## Sheet1
Manufacturer Name | Name of the medical device | Affected devices | Cause of the problem Ar | Cause of the problem En | Corrective action | Creation date
RICHARD WOLF | FLEX. GRASPING FORCEPS, FLEX. BIOPSY FORCEPS | 1 | - | as per attached FSN letter | as per attached FSN letter test | 2023-05-10 13:23
Richard wolf | FLEX. GRASPING FORCEPS, FLEX. BIOPSY FORCEPS | 1 | - | As per attached FSN letter | As per attached FSN letter | 2023-05-10 14:01
Spacelabs Healthcare, Inc. | Ultraview SL (UVSL) Command Module, Model 91496 | 2255 | - | Command modules will display the following when parameter processing is suspended depending on software version Older versions will have waveforms present with no alarms Newer versions have waveforms absent with no alarms This may cause clinicians to miss a potentially catastrophic event | Recommendations It is Spacelabs Healthcare’s recommendation that users must be adequately trained on all versions of the product that will be used. Please provide this notification and/or communicate correct and appropriate use to all users. Ensure that all users are aware that Suspend Processing status means that all parameter processing and alarms are stopped and will not function. The Suspended state must be manually re-engaged by pressing the “Resume Processing” button to return to active processing and return the alarms to the intended operation. Acknowledgment Requested Spacelabs requests that all customers acknowledge receipt and understanding of this notification by either mailing the included PDF copy of the response form or by completing the online form which can be accessed by using the URL below: https://spacelabshealthcare.com/support/recall-reply-form/2023-00001-c/ | 2023-05-14 09:18
richard | test | 1 | - | - | q233 | 2023-05-14 12:55
RX | TEST | 1 | - | RX | RX | 2023-05-14 13:03
Medtronic Inc. | Implantable Cardioverter Defibrillator | 6827 | - | There is a rare potential for reduced or no energy output during high voltage HV therapy typically J in implantable cardioverter defibrillators ICDs and cardiac resynchronization therapy defibrillators CRT Ds manufactured with a specific glassed feedthrough including currently available ICDs and CRT Ds x D x D There is an increased potential for a reduced or no energy HV therapy in the AX B configuration when all the following conditions are met x D The device has a glassed feedthrough manufactured after July x D There is significant separation of the layers of insulation materials in the feedthrough components of the device header x D An unintended current pathway forms within the void created by the insulation separation capable of conducting high levels of current during HV therapy x D x D When an unintended current pathway is detected during HV therapy the Short Circuit Protection SCP feature may trigger This behavior can be intermittent both full energy and reduced energy HV therapies within the same episode have been observed SCP events may also be lead related for both lead related and device related unintended current pathways the defibrillation waveform is truncated early in the energy delivery sequence resulting in reduced or no energy being delivered J x D | PATIENT MANAGEMENT RECOMMENDATIONS: Medtronic recognizes that each patient requires unique clinical considerations. Based on internal investigation and external consultation with our Independent Physician Quality Panel (IPQP), Medtronic provides the following guidance: • Prophylactic device replacement is NOT recommended. o The risk of mortality for patients after reprogramming is 0.001% at 9 years and is less than the risk of patient mortality due to complications associated with device replacement (0.032% - 0.043%1,2,3). • Program all HV therapy pathways B>AX in all therapy zones to minimize the risk for this issue. o Note: Using “Get Medtronic Nominals” will require manual reprogramming of Rx5 and Rx6 to B>AX for all ventricular therapies. o For patients remotely followed via CareLink, your Medtronic representative will provide a report to assist with identifying patients who may have one or more HV therapy pathways programmed AX>B. You may contact your local representative to obtain an updated copy of the report at any time. • Prioritize reprogramming patients who have both a history of HV therapy and Rx1 programmed AX>B. o Rx1 provides the greatest statistical likelihood to resolve an arrhythmia, and therefore it is important to minimize risk of a reduced- or no-energy HV therapy in the first sequence. • For remaining patients with AX>B programming in any HV therapy sequence, schedule (with appropriate discretion) the next follow-up for in-clinic reprogramming, to minimize potential for reduced- or no-energy HV therapies to occur. • Per standard practice, check tachyarrhythmia episodes to determine effectiveness of therapies that have been delivered. o Instruct patients to contact the clinic if they receive HV therapy or hear an audible tone coming from their device. o Verify delivered energy is consistent with programmed energy in the Episode Summary. Note: The image below highlights an episode experiencing intermittent reduced HV therapies (red boxes). • Contact your Medtronic representative if one of the following is observed as these may be an indication of either a device or lead-related issue: o Reduced- or no-energy HV therapy is displayed in Episode Text (regardless of programmed pathway) o A persistent drop of approximately 50% in RA, RV and LV pacing lead impedance measurements as this may be an indication of increased potential for a future reduced- or no-energy therapy. ACTIONS: • Acknowledge receipt of this notice. • This notice needs to be passed on to all those who need to be aware within your organization or to any organization where the potentially affected devices have been transferred. | 2023-05-14 19:36
Medtronic Inc. | Implantable Cardioverter Defibrillator Single Implantable Defibrillator Cardiac Implantable Defibrillator | 6827 | - | There is a rare potential for reduced or no energy output during high voltage HV therapy typically J in implantable cardioverter defibrillators ICDs and cardiac resynchronization therapy defibrillators CRT Ds manufactured with a specific glassed feedthrough including currently available ICDs and CRT Ds x D x D There is an increased potential for a reduced or no energy HV therapy in the AX B configuration when all the following conditions are met x D The device has a glassed feedthrough manufactured after July x D There is significant separation of the layers of insulation materials in the feedthrough components of the device header x D An unintended current pathway forms within the void created by the insulation separation capable of conducting high levels of current during HV therapy x D x D When an unintended current pathway is detected during HV therapy the Short Circuit Protection SCP feature may trigger This behavior can be intermittent both full energy and reduced energy HV therapies within the same episode have been observed SCP events may also be lead related for both lead related and device related unintended current pathways the defibrillation waveform is truncated early in the energy delivery sequence resulting in reduced or no energy being delivered J x D | PATIENT MANAGEMENT RECOMMENDATIONS: Medtronic recognizes that each patient requires unique clinical considerations. Based on internal investigation and external consultation with our Independent Physician Quality Panel (IPQP), Medtronic provides the following guidance: • Prophylactic device replacement is NOT recommended. o The risk of mortality for patients after reprogramming is 0.001% at 9 years and is less than the risk of patient mortality due to complications associated with device replacement (0.032% - 0.043%1,2,3). • Program all HV therapy pathways B>AX in all therapy zones to minimize the risk for this issue. o Note: Using “Get Medtronic Nominals” will require manual reprogramming of Rx5 and Rx6 to B>AX for all ventricular therapies. o For patients remotely followed via CareLink, your Medtronic representative will provide a report to assist with identifying patients who may have one or more HV therapy pathways programmed AX>B. You may contact your local representative to obtain an updated copy of the report at any time. • Prioritize reprogramming patients who have both a history of HV therapy and Rx1 programmed AX>B. o Rx1 provides the greatest statistical likelihood to resolve an arrhythmia, and therefore it is important to minimize risk of a reduced- or no-energy HV therapy in the first sequence. • For remaining patients with AX>B programming in any HV therapy sequence, schedule (with appropriate discretion) the next follow-up for in-clinic reprogramming, to minimize potential for reduced- or no-energy HV therapies to occur. • Per standard practice, check tachyarrhythmia episodes to determine effectiveness of therapies that have been delivered. o Instruct patients to contact the clinic if they receive HV therapy or hear an audible tone coming from their device. o Verify delivered energy is consistent with programmed energy in the Episode Summary. Note: The image below highlights an episode experiencing intermittent reduced HV therapies (red boxes). • Contact your Medtronic representative if one of the following is observed as these may be an indication of either a device or lead-related issue: o Reduced- or no-energy HV therapy is displayed in Episode Text (regardless of programmed pathway) o A persistent drop of approximately 50% in RA, RV and LV pacing lead impedance measurements as this may be an indication of increased potential for a future reduced- or no-energy therapy. ACTIONS: • Acknowledge receipt of this notice. • This notice needs to be passed on to all those who need to be aware within your organization or to any organization where the potentially affected devices have been transferred. | 2023-05-14 19:36
Medtronic Inc. | Implantable Cardioverter Implantable Cardioverter Defibrillator | 6827 | - | There is a rare potential for reduced or no energy output during high voltage HV therapy typically J in implantable cardioverter defibrillators ICDs and cardiac resynchronization therapy defibrillators CRT Ds manufactured with a specific glassed feedthrough including currently available ICDs and CRT Ds x D x D There is an increased potential for a reduced or no energy HV therapy in the AX B configuration when all the following conditions are met x D The device has a glassed feedthrough manufactured after July x D There is significant separation of the layers of insulation materials in the feedthrough components of the device header x D An unintended current pathway forms within the void created by the insulation separation capable of conducting high levels of current during HV therapy x D x D When an unintended current pathway is detected during HV therapy the Short Circuit Protection SCP feature may trigger This behavior can be intermittent both full energy and reduced energy HV therapies within the same episode have been observed SCP events may also be lead related for both lead related and device related unintended current pathways the defibrillation waveform is truncated early in the energy delivery sequence resulting in reduced or no energy being delivered J x D | PATIENT MANAGEMENT RECOMMENDATIONS: Medtronic recognizes that each patient requires unique clinical considerations. Based on internal investigation and external consultation with our Independent Physician Quality Panel (IPQP), Medtronic provides the following guidance: • Prophylactic device replacement is NOT recommended. o The risk of mortality for patients after reprogramming is 0.001% at 9 years and is less than the risk of patient mortality due to complications associated with device replacement (0.032% - 0.043%1,2,3). • Program all HV therapy pathways B>AX in all therapy zones to minimize the risk for this issue. o Note: Using “Get Medtronic Nominals” will require manual reprogramming of Rx5 and Rx6 to B>AX for all ventricular therapies. o For patients remotely followed via CareLink, your Medtronic representative will provide a report to assist with identifying patients who may have one or more HV therapy pathways programmed AX>B. You may contact your local representative to obtain an updated copy of the report at any time. • Prioritize reprogramming patients who have both a history of HV therapy and Rx1 programmed AX>B. o Rx1 provides the greatest statistical likelihood to resolve an arrhythmia, and therefore it is important to minimize risk of a reduced- or no-energy HV therapy in the first sequence. • For remaining patients with AX>B programming in any HV therapy sequence, schedule (with appropriate discretion) the next follow-up for in-clinic reprogramming, to minimize potential for reduced- or no-energy HV therapies to occur. • Per standard practice, check tachyarrhythmia episodes to determine effectiveness of therapies that have been delivered. o Instruct patients to contact the clinic if they receive HV therapy or hear an audible tone coming from their device. o Verify delivered energy is consistent with programmed energy in the Episode Summary. Note: The image below highlights an episode experiencing intermittent reduced HV therapies (red boxes). • Contact your Medtronic representative if one of the following is observed as these may be an indication of either a device or lead-related issue: o Reduced- or no-energy HV therapy is displayed in Episode Text (regardless of programmed pathway) o A persistent drop of approximately 50% in RA, RV and LV pacing lead impedance measurements as this may be an indication of increased potential for a future reduced- or no-energy therapy. ACTIONS: • Acknowledge receipt of this notice. • This notice needs to be passed on to all those who need to be aware within your organization or to any organization where the potentially affected devices have been transferred. | 2023-05-14 19:36
Medtronic Inc. | ICD-VR DVEX3E4 EV ICD EV4 PIVOTAL | 6827 | - | There is a rare potential for reduced or no energy output during high voltage HV therapy typically J in implantable cardioverter defibrillators ICDs and cardiac resynchronization therapy defibrillators CRT Ds manufactured with a specific glassed feedthrough including currently available ICDs and CRT Ds x D x D There is an increased potential for a reduced or no energy HV therapy in the AX B configuration when all the following conditions are met x D The device has a glassed feedthrough manufactured after July x D There is significant separation of the layers of insulation materials in the feedthrough components of the device header x D An unintended current pathway forms within the void created by the insulation separation capable of conducting high levels of current during HV therapy x D x D When an unintended current pathway is detected during HV therapy the Short Circuit Protection SCP feature may trigger This behavior can be intermittent both full energy and reduced energy HV therapies within the same episode have been observed SCP events may also be lead related for both lead related and device related unintended current pathways the defibrillation waveform is truncated early in the energy delivery sequence resulting in reduced or no energy being delivered J x D | PATIENT MANAGEMENT RECOMMENDATIONS: Medtronic recognizes that each patient requires unique clinical considerations. Based on internal investigation and external consultation with our Independent Physician Quality Panel (IPQP), Medtronic provides the following guidance: • Prophylactic device replacement is NOT recommended. o The risk of mortality for patients after reprogramming is 0.001% at 9 years and is less than the risk of patient mortality due to complications associated with device replacement (0.032% - 0.043%1,2,3). • Program all HV therapy pathways B>AX in all therapy zones to minimize the risk for this issue. o Note: Using “Get Medtronic Nominals” will require manual reprogramming of Rx5 and Rx6 to B>AX for all ventricular therapies. o For patients remotely followed via CareLink, your Medtronic representative will provide a report to assist with identifying patients who may have one or more HV therapy pathways programmed AX>B. You may contact your local representative to obtain an updated copy of the report at any time. • Prioritize reprogramming patients who have both a history of HV therapy and Rx1 programmed AX>B. o Rx1 provides the greatest statistical likelihood to resolve an arrhythmia, and therefore it is important to minimize risk of a reduced- or no-energy HV therapy in the first sequence. • For remaining patients with AX>B programming in any HV therapy sequence, schedule (with appropriate discretion) the next follow-up for in-clinic reprogramming, to minimize potential for reduced- or no-energy HV therapies to occur. • Per standard practice, check tachyarrhythmia episodes to determine effectiveness of therapies that have been delivered. o Instruct patients to contact the clinic if they receive HV therapy or hear an audible tone coming from their device. o Verify delivered energy is consistent with programmed energy in the Episode Summary. Note: The image below highlights an episode experiencing intermittent reduced HV therapies (red boxes). • Contact your Medtronic representative if one of the following is observed as these may be an indication of either a device or lead-related issue: o Reduced- or no-energy HV therapy is displayed in Episode Text (regardless of programmed pathway) o A persistent drop of approximately 50% in RA, RV and LV pacing lead impedance measurements as this may be an indication of increased potential for a future reduced- or no-energy therapy. ACTIONS: • Acknowledge receipt of this notice. • This notice needs to be passed on to all those who need to be aware within your organization or to any organization where the potentially affected devices have been transferred. | 2023-05-14 19:36
- | - | 1 | - | - | - | 2023-05-31 07:41
- | - | - | - | - | - | 2023-05-31 08:02
Abbott Ireland Diagnostics Division | Alinity i HBsAg Reagent Kit | 1 | - | - | - | 2023-05-31 08:51
Natus Neurology Incorporated | Disposable Ring Electrodes w/Lead | 212 | - | - | - | 2023-05-31 08:55
Abbott Diabetes Care Limited | FreeStyle Optium Blood Glucose and Ketone Monitoring System | 22 | - | - | - | 2023-05-31 08:58
Zimmer Inc. | Zimmer Periarticular Locking Plate System Distal | 7 | - | possible thread form issue of the locking holes that could result in locking screws that would not properly mate with the plate | Quarantine all affected product. Zimmer Biomet representative will remove the affected product from affected facilities. | 2023-07-16 11:37
Cook Medical Europe Limited | Fusion Lithotripsy Extraction Basket | 2 | - | Cook conducted a complaint investigation that confirmed on Apr that nonconforming Fusion Lithotripsy Extraction Baskets were manufactured and distributed This manufacturing nonconformance can appear in the field as inability to advance or retract the basket or the drive wire breaking or separating from handle with or without mechanical lithotripsy Cook has received five complaints worldwide for this nonconformance | Please complete the enclosed Customer Reply Form. Since devices are to be returned, our Customer Services department will contact you to organize the return and issue the relevant Returns Authorization number. Please include contact details on the Customer Reply Form so they can contact you. Returned Devices should be addressed to: Cook Medical EUDC Robert-Koch-Straße, 2 52499 Baesweiler Germany Credit will be provided for the returned affected devices where applicable. | 2023-07-18 14:35
Getinge AB | Aperlan Poka-Yoke agent A | 12 | - | The root cause of this issue has not yet been fully identified and we would like to inform you about certain measures that must be taken to ensure a safe product | Perform a control of the red and blue caps on the canisters for both Aperlan Poka-Yoke agent A and Poka-Yoke agent B for any deviation (loose or damaged caps) just before the canisters are placed in the ED-Flow AER (Automated Endoscope Re processor) for use. If the cap can be removed without effort or if it is not tightened at all: Do not use the product. Discard the canister according to your local regulations in terms of dangerous goods.  If any of the soldering points of the tampering ring is broken: Do not use the product. Discard the canister according to your local regulations in terms of dangerous goods. | 2023-07-20 17:11
Getinge AB | Aperlan Poka-Yoke agent B | 12 | - | The root cause of this issue has not yet been fully identified and we would like to inform you about certain measures that must be taken to ensure a safe product | Perform a control of the red and blue caps on the canisters for both Aperlan Poka-Yoke agent A and Poka-Yoke agent B for any deviation (loose or damaged caps) just before the canisters are placed in the ED-Flow AER (Automated Endoscope Re processor) for use. If the cap can be removed without effort or if it is not tightened at all: Do not use the product. Discard the canister according to your local regulations in terms of dangerous goods.  If any of the soldering points of the tampering ring is broken: Do not use the product. Discard the canister according to your local regulations in terms of dangerous goods. | 2023-07-20 17:11
Abbott | Proclaim XR SCS/Elite SCS/DRG Neurostimulation systems and Infinity DBS System | - | - | During standard use a Bluetooth x D x D paired Patient Controller is used to place a patient s implantable pulse x D generator in MRI mode which in turn disables the delivery of therapy Upon completion of the MRI the paired x D Patient Controller is then used to exit MRI mode and allow therapy to resume | Abbott recommends physicians to do the following to reduce the risk of loss of therapy and the need for IPG replacement surgery to restore therapy: • For Patient Controller, advise the patient not to delete the paired Bluetooth ® connection between their IPG and the Patient Controller and to not alter, damage or lose their Patient Controller while the IPG is in MRI mode. • For Clinician Programmer, maintain the paired Bluetooth ® connection between the IPG and the Clinician Programmer by avoiding deletion of the pairing and disabling automatic iOS ‡ software upgrades from the iOS ‡ Settings. • For questions about this issue or to report patients who are unable to exit MRI mode, please contact your local Abbott representative for assistance | 2023-07-26 09:36
Philips | Big Bore RT | 1 | - | Philips has identified multiple software issues affecting Philips Big Bore RT and Upgrade Brilliance CT Big Bore systems Some of these issues may cause the system to shut down unexpectedly or to generate unsatisfactory images which may result in the need to perform a rescan see Appendix A issues Detailed descriptions and advice to customers pertaining to all identified issues are provided in Appendix A in FSN letter | Philips will contact you to schedule a time for a Philips Field Service Engineer (FSE) to visit your site and install the software update (FCO 72800789) to resolve the issues listed in Appendix A. | 2023-07-27 10:16
Abbott GmbH. & Co. KG | Alinity s Anti-HBc Reagent Kit | 19 | - | As per attached letter | As per attached letter | 2023-07-27 13:51
Smith & Nephew plc. | TRIGEN TROCHANTERIC ANTEGRADE NAIL | 39 | - | A complaint was received indicating that a package contained a TRIGEN TROCHANTERIC ANTEGRADE NAIL LIME MM X CM DEGREE LEFT implant instead of a TRIGEN TROCHANTERIC ANTEGRADE NAIL LIME MM X CM DEGREE LEFT as described on the product label | ◼ Identify and quarantine the products in your warehouse. ◼ Identify and inform all users, which have received affected products with the attached Field Safety Notice for Recall. ◼ Collect return slips of the Field Safety Notice for Recall (to confirm awareness of all affected users) and forward them to the specified contact below. If acknowledgement is not provided by any of the affected users provide evidence of three attempts to notify these users. ◼ Complete the acknowledgment section on page 3 and return it to the specified contact below. ◼ Return collected products as instructed below. Please maintain awareness on this notice and resulting action until the Field Action is terminated to ensure effectiveness of the action. This notice needs to be passed on to all those who need to be aware within your organization or to any organization where the potentially affected products have been transferred. | 2023-07-27 16:38
Diversatek Healthcare | Zvu® Functional GI Software | - | - | Diversatek Healthcare has identified an issue with Zvu Functional GI Software Part Number ZVU Version shipped between and The product label contains the incorrect software revision It is labelled as revision but it contains revision | 1-identify those devices with version 3.3.2109.6 2-Notify Diversatek Healthcare customer service of the discarded quantity to receive a replacement. | 2023-07-29 20:04
Siemens Healthcare Diagnostics Inc | Atellica® CH 930 Analyzer | - | - | Reason for Correction x D The purpose of this communication is to inform you of an issue with the products indicated in Table above and to provide instructions on actions that your laboratory must take x D If you do not run any of the assays listed in Table there are no actions for your laboratory to take at this time x D Through a proactive internal screening Siemens Healthcare Diagnostics Inc has identified the potential for reagent carryover on the Atellica CH resulting in a positive bias that could impact quality control QC patient samples and calibrator results with specific assays as listed below in Tables The addition of Reagent Probe Cleaner RPC wash mitigates this issue x D The resolution for the reagent carryover issues is implemented in Atellica Solution Software SW v SP and v which will be released soon In the interim please follow the instructions in the Actions to be Taken by the Customer section until all Atellica CH Analyzers in your laboratory are updated to either SW v SP or v and higher x D Kindly review the attached UFSN | Kindly review the attached UFSN | 2023-07-31 09:01
TEST | TEST | - | - | TEST | TEST | 2023-08-01 14:05
Ral Diagnostics | pH= 7.0 BUFFER SOLUTION | 275 | - | The user observes the presence of bacilli on blood smears customer complaints have been | Product Removal / On-site device modification/inspection | 2023-08-07 14:33
Getinge | Datascope Cardiosave Hybrid and Rescue Intra-Aortic Balloon Pumps (IABP) | 100 | - | Issue Docking Power Battery Failure x D Issue Poor or No ECG Signal x D Issue Autofill Alarms x D Issue Gas Loss Gas Gain Alarms x D Issue System Over Temperature x D Issue Fiber Optic Damage x D | please refer to both letters | 2023-08-07 14:38
Agilent Technologies Singapore (International) Pte Ltd. | EnVision FLEX, High pH (Dako Omnis), EnVision FLEX Mini Kit, High pH (Dako Omnis) | 9 | - | we have identified an issue with increased level of artefacts high background staining on certain GV EnVision FLEX High pH Dako Omnis Kits and kits of GV EnVision FLEX Mini Kit High pH Dako Omnis | Actions to be taken by the customer 1. In case increased background staining/ artefacts are observed on the slides, inspect the stainer gap lids on the Dako Omnis. 2. If debris (an opaque brown layer) is visible on the lids, perform the process for manual cleaning of the stainer gap lids as described in the Dako Omnis Basic User Guide section 24. | 2023-08-07 14:46
Haemonetics Corporation | Leukotrap RCPL | - | - | Haemonetics has identified an increase in reports of particulate matter found inside platelet x D satellite bags for certain lots of finished devices similar to reference number RPT CCL The x D particles have been determined to be polyester originating from the ATS LPL platelet filter which x D is also present on product reference RPT CCL circled on schematics below The particles x D vary in size and may not be visible to the user during processing Some particles may be small x D enough to pass through additional pre transfusion filters x D Frequency of failures and complaints Haemonetics has received reports of polyester x D particles in products out of approximately units distributed globally In some x D instances particles may not be visually detectable None of the reports were on product x D reference RPT CCL x D Adverse events No adverse events have been reported in relation to this issue x D Product Distribution Information x D The actions defined in this FSN are applicable to RPT CCL lots manufactured after May st x D with lot numbers after and including | Actions to be taken by Distributor: Our records indicate that you have distributed the devices affected by this notice. We therefore ask you to perform the following activities:  Immediately inform all your impacted customers of this FSN, ask them to stop using the affected product and return any unused quantities. For your convenience, we have attached a sample customer letter in annex 1. If required, translate this letter into local language  If required by your local regulations, inform the appropriate authorities. Note: Haemonetics will inform the appropriate European authorities (including EEA and Turkey).  Complete the attached acknowledgement form in its entirety whether or not you have distributed affected product. Once complete, return the form to Haemonetics following the instructions on the form. Your response is vital to our monitoring of the effectiveness of this action. Haemonetics has put all inventory of potentially affected products on hold while our investigation is ongoing to identify and resolve the root cause of the issue. We apologize for any disruption this situation may cause your organization and we thank you for your business and continued support. Haemonetics is committed to continually improving its products and services, with safety and quality as our top priority. This action is being performed by Haemonetics in full transparency with regulatory authorities. Please contact your Haemonetics representative or customer support if you have any questions. | 2023-08-08 12:28
OCULUS Optikgerate GmbH | Pentacam® AXL | - | - | Problem description x D One of our suppliers did not meet the specifications for an optical component inside the x D devices An insufficient anti reflective coating was found which under certain circumstances can x D lead to an additional axial length signal x D If the correct axial length signal is weaker e g due to an off centre measurement or a cataract x D an incorrect axial length value can be displayed x D OCULUS identified the problem within the production process We have reviewed our records x D and identified previously shipped units where this issue is not likely but not completely ruled x D out to occur x D However there have been no reports of physical injury or serious negative consequences | Kindly review the attached UFSN corrective action | 2023-08-09 08:23
OCULUS Optikgerate GmbH | Pentacam® AXL Wave | - | - | Problem description x D One of our suppliers did not meet the specifications for an optical component inside the x D devices An insufficient anti reflective coating was found which under certain circumstances can x D lead to an additional axial length signal x D If the correct axial length signal is weaker e g due to an off centre measurement or a cataract x D an incorrect axial length value can be displayed x D OCULUS identified the problem within the production process We have reviewed our records x D and identified previously shipped units where this issue is not likely but not completely ruled x D out to occur x D However there have been no reports of physical injury or serious negative consequences | Kindly review the attached UFSN corrective action | 2023-08-09 08:23
St. Jude Medical | Trifecta Valve | 18 | - | Abbott has decided to voluntarily recall the Trifecta x D and Trifecta with Glide Technology GT valves x D On February Abbott communicated the potential for early Structural Valve Deterioration SVD and provided patient management considerations for those patients implanted with the Trifecta and Trifecta GT valves A review of published literature and clinical trial data has shown that the reported rates of SVD for the Trifecta valve do not consistently align with the prospective clinical trial data and demonstrate greater variation across medical centers relative to comparator valves Abbott s assessment of incidence and risks has not changed since the February communication x D Reasons for the Field Safety Corrective Action July x D Abbott had previously decided to discontinue its Trifecta family of valves to focus on tissue heart valve solutions that maximize possibilities for lifetime management of valvular heart disease and has requested withdrawal of its CE certification In consultation with the Competent Authority Abbott is now requesting removal of unused Trifecta and Trifecta GT valves | Patient Management Considerations: Abbott does not recommend that patients already implanted with the valve undergo a prophylactic device explant. Patient management considerations previously described in our Trifecta family of valves communication from 27 February 20231 remain in effect. Steps Abbott is Requesting You to Take: 1. Return any remaining unused product to Abbott. Your Abbott representative can assist you in returning these devices. 2. Complete and return the accompanying Acknowledgment Form to Abbott. | 2023-08-09 11:10
Philips Medical Systems Nederland B.V. | Allura Xper FD20 | 167 | - | The wired and wireless foot switch are used to control fluoroscopy exposure and other functions such as single shot light control and toggle between X ray planes for bi plane systems x D A foot switch pedal may get stuck in the active position when the user releases the pedal resulting in unintended radiation because of x D Build up of dense or sticky fluids on the foot switch e g blood or contrast fluid if not properly cleaned x D Use of protective covers that are either the wrong size or incorrectly placed on the foot switch x D Dislodgment of a screw holding the pick up bar see Figure of the foot switch subsequently becoming lodged within the foot switch housing x D An additional screw inadvertently left inside the foot switch housing during manufacturing x D | Philips will contact customers to schedule a visit to inspect the foot switch and provide a copy of IFU addendum (reference FCO72200545). Based on available information, systems may safely continue to be used in accordance with the device instructions for use and the provided instructions in this Urgent Field Safety Notice letter. | 2023-08-09 15:46
Philips Medical Systems Nederland B.V. | Allura Xper FD10 | 167 | - | The wired and wireless foot switch are used to control fluoroscopy exposure and other functions such as single shot light control and toggle between X ray planes for bi plane systems x D A foot switch pedal may get stuck in the active position when the user releases the pedal resulting in unintended radiation because of x D Build up of dense or sticky fluids on the foot switch e g blood or contrast fluid if not properly cleaned x D Use of protective covers that are either the wrong size or incorrectly placed on the foot switch x D Dislodgment of a screw holding the pick up bar see Figure of the foot switch subsequently becoming lodged within the foot switch housing x D An additional screw inadvertently left inside the foot switch housing during manufacturing x D | Philips will contact customers to schedule a visit to inspect the foot switch and provide a copy of IFU addendum (reference FCO72200545). Based on available information, systems may safely continue to be used in accordance with the device instructions for use and the provided instructions in this Urgent Field Safety Notice letter. | 2023-08-09 15:46
Sentec AG | Membrane Changer | 66 | - | During the manufacturing process there were components that required rework and during this step but before assembly an additional piece of material may have been left in the device itself This material is less than mm in diameter and is inert and non toxic x D Thesource of this material has been explicitly identified and Sentec AG can with certainty list theaffected production lots | SenTec advises its distributors as well as all end-customers to take the following immediate actions: On distributor level: 1. Inform sub-distributors (if any) 2. Inform customers 3. Inventory and logistic flow: stop ship to customers any items with the affected lot number 4. Confirm the completion of action 1 and 2 and your commitment to action 3 On customer level: 1. Follow the Membrane Changer Instructions for Use (included as Appendix II of the FSN; RF-013992-b). 2. Visually inspect the sensor after membrane change. 3. If any integrity of the sensor is compromised or it appears as like the following example – please immediately change the membrane again. 4. If you experience a defective membrane, please contact Customer Service and report a Complaint to receive a replacement product at no charge. | 2023-08-10 13:24
Sentec AG | Membrane Changer Insert | 66 | - | During the manufacturing process there were components that required rework and during this step but before assembly an additional piece of material may have been left in the device itself This material is less than mm in diameter and is inert and non toxic x D Thesource of this material has been explicitly identified and Sentec AG can with certainty list theaffected production lots | SenTec advises its distributors as well as all end-customers to take the following immediate actions: On distributor level: 1. Inform sub-distributors (if any) 2. Inform customers 3. Inventory and logistic flow: stop ship to customers any items with the affected lot number 4. Confirm the completion of action 1 and 2 and your commitment to action 3 On customer level: 1. Follow the Membrane Changer Instructions for Use (included as Appendix II of the FSN; RF-013992-b). 2. Visually inspect the sensor after membrane change. 3. If any integrity of the sensor is compromised or it appears as like the following example – please immediately change the membrane again. 4. If you experience a defective membrane, please contact Customer Service and report a Complaint to receive a replacement product at no charge. | 2023-08-10 13:24
B. Braun Melsungen AG | SPACEPLUS INFUSOMAT | 1030 | - | As per FSN attached | As per FSN attached | 2023-08-15 12:07
B. Braun Melsungen AG | Infusomat® Space - Upstream Occlusion Sensor | 1241 | - | As attached FSN | As attached FSN | 2023-08-15 13:10
Abbott | HeartMate Touch™ Communication System (Model: HMT1150) | - | - | Abbott is notifying customers of a single reported event with the HeartMate Touch x D Communication System also referred to as HeartMate Touch This event was not associated x D with patient harm The HeartMate Touch Communication System is part of the HeartMate x D and HeartMate II Left Ventricular Assist System LVAS and is used in the medical facilities x D for patient monitoring and system programming x D During a clinic visit the HeartMate Touch Communication System in one treatment room x D connected to the HeartMate Touch Wireless Adapter plugged into the Power Module in an x D adjacent treatment room without a user initiated action or notification to the user The x D connection to an unintended LVAS without any user interaction led to the clinician adjusting x D settings of another patient without the clinician s awareness The HeartMate Touch System in x D the adjacent room did not have battery charge and its power cable was not plugged in prior to x D being used it is unknown if the lack of battery and power connection contributed to the reported x D event | Abbott is reinforcing the existing cautions in the Instructions For Use when utilizing the HeartMate Touch Communication System: • The HeartMate Touch Communication System should be fully charged, or its power cable should be plugged in before starting this procedure. • Verify that the HeartMate Touch Communication System is connected to the intended HeartMate LVAS when making changes to LVAS settings. Page 1 of 3 The HeartMate LVAS System connects to the HeartMate Touch Communication System via connection to the Power Module and the associated HeartMate Touch Wireless Adapter that uses Bluetooth Wireless Technology. Prior to making any changes to pump settings, user should confirm that they are connected to the intended LVAS System by locating the adapter as described in the ‘HeartMate Touch Communication System Interface Overview’ section of the Instructions for Use | 2023-08-15 13:29
Philips | Essenta DR Compact | 10 | - | There is a potential cybersecurity issue with Essenta DR Compact Systems which could result in unauthorized disclosure or modification of patient data This issue may occur if a firewall is not in place the file transfer protocol FTP is utilized and the following sequence of events occurs x D x D Unauthorized personnel gains physical access to your facility x D Your facility does not have a firewall and or network segregation to isolate medical devices and x D The unauthorized person connects their device to the Essenta DR Compact system network at the specific time data is transferred data is not continuously transferred x D x D Under these conditions unauthorized access or modification of patient data is possible x D | Philips is providing this Field Safety Notice which contains recommendations for network protection described in Appendix A. | 2023-08-16 11:21
Johnson & Johnson International c/o European Logistics Centre | PDS II™ | - | - | Ethicon received seven complaints for one lot of undyed PDS II polydioxanone x D Sutures regarding low tensile strength Internal testing on returned product from this lot x D confirmed that some PDS II polydioxanone Sutures from this lot did not meet x D Ethicon s tensile strength requirement An investigation identified additional lots across x D the PDS Plus Antibacterial polydioxanone and PDS II polydioxanone Suture x D families that have the potential to be affected by the same root cause as the lot that x D received the complaints | Examine your inventory immediately to determine if you have product subject to this recall on hand and quarantine such product(s). | 2023-08-16 12:14
Johnson & Johnson International c/o European Logistics Centre | PDS Plus™ | - | - | Ethicon received seven complaints for one lot of undyed PDS II polydioxanone x D Sutures regarding low tensile strength Internal testing on returned product from this lot x D confirmed that some PDS II polydioxanone Sutures from this lot did not meet x D Ethicon s tensile strength requirement An investigation identified additional lots across x D the PDS Plus Antibacterial polydioxanone and PDS II polydioxanone Suture x D families that have the potential to be affected by the same root cause as the lot that x D received the complaints | Examine your inventory immediately to determine if you have product subject to this recall on hand and quarantine such product(s). | 2023-08-16 12:14
Bio-Rad | Liquichek Pediatric Control | 27 | - | We have observed that Bilirubin Total and Direct may not meet the Refrigerated Off Board x D or the Instrument On Board storage stability claims as printed in the package inserts x D Therefore the stability claims for these storage conditions for all analytes have been revised x D from days to days as indicated below by the highlighted text and bold font | Please obtain a revised copy of Liquichek Pediatric Control package insert through the Internet at http://myeinserts.qcnet.com. Discard all previous versions of the package insert for these product lots that you may have at your location and substitute with the revised package insert. • Contact Bio-Rad Laboratories Technical Support Department at 1-800-854-6737 if you have technical questions regarding this notice. Alternate methods for receiving the package insert or Instructions for Use (IFU) are available by contacting your local BioRad Laboratories Office. Contact Bio-Rad Laboratories Technical Support Department (USA) at 1-800-854-6737, option #2, or by email at qsd.techservice@bio-rad.com if you have technical questions regarding this notice. | 2023-08-16 15:10
Carl Zeiss Suzhou Co.,Ltd | OPMI LUMERA 300 | 5 | - | In certain devices there is a probability of specific screw missing during assembly This screw is used to prevent one internal thread cap rotating which ensure the connection between suspension arm and the suspended components The identified defect may lead to a potential risk x D In consequence we Carl Zeiss Suzhou Co Ltd have decided to take field safety corrective action on all OPMI LUMERA that may potentially have the problem to inform customers and prevent harm to patients | Urgent and important mandatory action to be taken by user: The gap shown is Fig 1 indicates the fasten status of thread cap (not visible) which carries all load of suspension arm. If this gap is not in its nominal value, there is a risk of OPMI falls from suspension arm. Before each use of OPMI LUMERA 300, please perform the following mandatory inspection: • Measure the gap shown in Fig. 1 If the gap is greater than 5mm, please Immediately stop using the device! In this case the OPMI LUMERA 300 can no longer be used. We are planning to inspect (repair if needed) of the suspension arm immediately. Our service staff will try to complete the action in AR's warehouse. | 2023-08-16 15:48
Accuray Incorporated | Radixact™ Treatment Delivery System | 1 | - | Accuray Incorporated Accuray has discovered an issue related to the use of a jib crane by service personnel while servicing the Radixact System It is possible that the upper roller bracket in the Radixact System gantry enclosure may not provide enough clearance for service personnel to use a jib crane for moving components while servicing the system | Accuray will inspect the upper roller bracket of each impacted Radixact systems to confirm there is adequate clearance for a jib crane. If there is sufficient clearance for the operation of a jib crane, then no action is necessary, and the inspection will be documented. If there is insufficient clearance, then the old bracket will be replaced with the new bracket following the released updated service procedures. | 2023-08-16 17:33
Olympus Medical System Corp | Bronchofiberscope, Bronchovideoscope, Ultrasonic Bronchofibervideoscope, Tracheal Intubation Fiberscope, Tracheal Intubation Videoscope, Airway Mobilescope | - | - | Olympus has received three complaints including one that was associated with an injury that upon trying to use an Olympus pulmonary endoscope model with an endotracheal tube the tip became lodged entrapped inside the endotracheal tube connector As a result of the complaint investigation it was determined that the scope bending section was too large for the endotracheal tube connector | Actions to be taken by the end user: Olympus requests you to take the following actions: 1. Inspect your inventory for the referenced devices and identify any device with the model names specified above. Please check all areas of the hospital to determine if any of these devices remain in inventory. 2. Carefully read the content of this Field Safety Notice as well as the attached “Addendum”. The Addendum provides information on the maximum insertion tube outer diameter for the affected bronchoscopes/tracheal intubation scopes/airway mobilescopes. 3. Ensure all personnel are completely knowledgeable and thoroughly aware of the contents. 4. Send the completed Reply Form back to your local Olympus representative at ra@olympus-mea.com latest by 21 August 2023. 5. If you have further distributed this product, identify your customers, forward them this notification, and appropriately document your notification process. | 2023-08-17 13:45
Philips | MultiDiagnost Eleva with Flat Detector | 6 | - | The wired foot switch is used to control fluoroscopy exposure and other functions such as single shot light control and toggle between X ray planes for bi plane systems x D x D Philips has identified instances where no or only intermittent X ray radiation initiation is possible through use of the wired foot switch as a result of x D x D A damaged foot switch cable or cable connector caused by large external force on the cable such as x D o The cable being run over by other medical equipment x D Supplier manufacturing issues associated with the production of certain components of the foot switch x D x D | Philips will be inspecting all affected systems to check the foot switch cable. Philips will contact to schedule a visit to perform this inspection (reference FCO72200534). | 2023-08-17 13:49
Philips | Allura Xper FD10/10, Allura Xper FD20, Allura Xper FD20 Biplane, Allura Xper FD10, Allura Xper FD20 OR Table, Allura Xper FD20/10, Allura CV20, Allura Xper FD10 OR Table, Allura Xper FD20/15, Azurion 3 M12, Azurion 3 M15, Azurion 7 B12, Azurion 7 B20, Azurion 7 M12, Azurion 7 M20, Azurion 5 M20, Cardio Vascular-Allura Centron | 159 | - | The wired foot switch is used to control fluoroscopy exposure and other functions such as single shot light control and toggle between X ray planes for bi plane systems x D x D Philips has identified instances where no or only intermittent X ray radiation initiation is possible through use of the wired foot switch as a result of x D x D A damaged foot switch cable or cable connector caused by large external force on the cable such as x D o A stuck cable being pulled along during table pivot swivel x D o The cable being caught inadvertently during moving transferring of equipment x D o The cable being run over by other medical equipment x D x D The strain relief connector not being properly secured during installation servicing or a broken strain relief connector caused by external forces applied on the foot switch cable With loss of the strain relief connector the foot switch cable can become disconnected from the system when force is applied to the cable x D A damaged foot switch cable caused by the cable getting stuck between the patient table and table cover because the table plinth is not adjusted correctly to prevent a gap between the table cover and floor x D Supplier manufacturing issues associated with the production of certain components of the foot switch x D | Philips has updated the Repair Instructions to ensure that the patient table and table cover is correctly adjusted to prevent a gap between the table cover and the floor. Philips has also updated its Preventative Maintenance Manual to include additional activities to help ensure the proper operation of the foot switch. Philips will be inspecting all affected systems to check the foot switch cable, to ensure that the table plinth is adjusted correctly, and to ensure that the strain relief connector is properly secured. Philips will contact to schedule a visit to perform this inspection (reference FCO72200534). | 2023-08-17 14:00
Creganna Medical | GORE TIPS SET | 3 | - | mislabeled products may indicate a longer expiration date | remove and replace affected product before reaching their approved expiration | 2023-08-17 15:09
GE Healthcare | Enterprise Archive | 42 | - | Issue has been identified where images from two different patients may be contained in a single study when stored in Enterprise Archive EA as outlined in the following events x B The user incorrectly selects Patient A order at the acquisition device such as a DX scanner when performing a study on Patient B and storing in EA x B The user changes the Patient Demographics on the study stored in EA to Patient B without changing the Study Instance UID x B When Patient A arrives and the original order for Patient A is selected again at the acquisition device the images of Patient A get added to the study used by Patient B when stored in EA This is due to the same study instance UID used for both the studies The Patient Demographics of the study are changed back to Patient A In the rare event that this issue is not identified it is possible to use the images from Patient B to diagnose Patient A which can result in misdiagnosis for Patient A x B In addition this issue can also occur if an outside study being ingested in to the EA DICOM archive does not maintain an unique study instance UID x B As per DICOM standard the same study instance UID should not be used for two different studies x D | Corrective Actions: [CA1] EA FMEA/CMT to address different patient Info present in a single Study (Study UID Reuse) [Process RC #2] [ EA v4x/v8x] [CA2] Update EA Requirements to cover handling Study UID misuse (Reuse/Duplicate Study UID case ) [Process RC #1] [EA v4x/v8x] [CA3] Release EA v8x SW basic consistency checker to address different patient Info present in a single Study (Study UID) [Technical RC] [CA4] Release EA v4x SW basic consistency checker to address different patient Info present in a single Study (Study UID) [Technical RC] [CA5]: Deploy Field Action FMI 85469 to correct existing Centricity Enterprise Archive (EA) and/or Enterprise Archive (EA) Install Base (IB). [Technical RC] Preventive Actions: [PA1] Complete EA FMEA/CMT to address similar hazard covering scenario such as duplicate UIDs for all UIDs [Process RC #2] [ EA v4x/v8x] [PA2] Address EA handling of Study UID/UID Misuse in EA forward production requirements [Process RC #1] [EA v8SP1+] [PA3] Usability Analysis, Testing and Fix identified Usability gaps for FMI [Process RC #1] | 2023-08-17 15:25
Philips Medical Systems Nederland B.V. | INGENIA AMBITION S | 2 | - | The magnet within the MR System is equipped with lifting blocks which are positioned on top of the magnet These blocks are used to allow the magnets to be moved in or out by a lifting crane e g during decommissioning or moving activities x D x D A mechanical issue was detected on these hoisting interface blocks on top of the magnet In the initial design of the hoisting blocks the diameter to meet the required bolt on interface surface for the swivel eye bolts used as lifting hardware did not meet the required specification x D x D An unacceptable hazardous situation could occur only when the magnet is lifted directly by crane during decommissioning or moving activities x D x D To solve this issue a spacer has been designed to fit around the lifting block boss and provide the required interface surface x D x D Philips has not received any reports of adverse events associated with this issue x D | A Philips representative will contact to schedule an on-site visit to resolve the issue by implementing a technical solution (FCO78100524). | 2023-08-21 10:00
Philips Medical Systems Nederland B.V. | INGENIA AMBITION X | 2 | - | The magnet within the MR System is equipped with lifting blocks which are positioned on top of the magnet These blocks are used to allow the magnets to be moved in or out by a lifting crane e g during decommissioning or moving activities x D x D A mechanical issue was detected on these hoisting interface blocks on top of the magnet In the initial design of the hoisting blocks the diameter to meet the required bolt on interface surface for the swivel eye bolts used as lifting hardware did not meet the required specification x D x D An unacceptable hazardous situation could occur only when the magnet is lifted directly by crane during decommissioning or moving activities x D x D To solve this issue a spacer has been designed to fit around the lifting block boss and provide the required interface surface x D x D Philips has not received any reports of adverse events associated with this issue x D | A Philips representative will contact to schedule an on-site visit to resolve the issue by implementing a technical solution (FCO78100524). | 2023-08-21 10:00
Great glove xinghua co LTD. | Nahdi examination gloves | 1 | - | Not conforming to the technical specification ISO Annex A Water tightness testing for Medical | dispose current stock and implement Water tightness testing for Medical gloves for future orders | 2023-08-21 14:59
Siemens Healthcare Diagnostics Inc | Atellica® CH LITH_2 | 14560 | - | Reason for Communication x D In September Siemens Healthcare Diagnostics Inc issued an Urgent Field Safety Notice x D UFSN ACHC A OUS to inform customers with in date reagent lots of the potential for x D falsely elevated lithium results x D Investigation findings have since identified the reagent chimney within the reagent packs x D contributed to foam formation x D Siemens has implemented changes to the lithium reagent pack and test definition Starting with x D reagent lot the lithium reagents will be packaged using reagent packs without the chimney x D Atellica Solution Software SW v or higher must be installed prior to running LITH x D This SW version includes x D Change in Onboard Stability from days to days x D Change in Pack Calibration Interval from days to days x D All other performance characteristics remain the same x D The Instructions for Use for the LITH assay Rev has been updated with the x D above information x D Siemens is actively working towards replenishing inventory levels for this product In order to x D ensure that all customers have an adequate supply full order quantities may not be immediately x D available | Actions to be Taken by the Customer Customers may now order Lithium_2 reagents for use on the Atellica® CH 930 Analyzer. Please be advised that new orders will have to be placed. The SMN has not changed. Ensure your system has been upgraded to Atellica Solution Software (SW) v1.27 or higher prior to running LITH_2. Download the latest revision of the IFU. Please retain this letter with your laboratory records and forward this letter to those who may have received this product. If you have any questions, please contact your Siemens Healthineers Customer Care Center or your local Siemens Healthineers technical support representative. | 2023-08-23 08:43
Siemens Healthcare GmbH | MAGNETOM Aera | 148 | - | In some cases when lubricating grease is present within the headband of the over ear headphones it may cause visible dot or line shaped fat isointense artifacts during head examinations only with the MAGNETOM MRI systems listed in the attachment x D | Please refer to the attachment for steps that the user can take to avoid the possible risks associated with the mentioned issue. | 2023-08-23 11:39
Siemens Healthcare GmbH | MAGNETOM Prisma fit | 148 | - | In some cases when lubricating grease is present within the headband of the over ear headphones it may cause visible dot or line shaped fat isointense artifacts during head examinations only with the MAGNETOM MRI systems listed in the attachment x D | Please refer to the attachment for steps that the user can take to avoid the possible risks associated with the mentioned issue. | 2023-08-23 11:39
Siemens Healthcare GmbH | MAGNETOM Altea | 148 | - | In some cases when lubricating grease is present within the headband of the over ear headphones it may cause visible dot or line shaped fat isointense artifacts during head examinations only with the MAGNETOM MRI systems listed in the attachment x D | Please refer to the attachment for steps that the user can take to avoid the possible risks associated with the mentioned issue. | 2023-08-23 11:39
Siemens Healthcare GmbH | MAGNETOM Skyra | 148 | - | In some cases when lubricating grease is present within the headband of the over ear headphones it may cause visible dot or line shaped fat isointense artifacts during head examinations only with the MAGNETOM MRI systems listed in the attachment x D | Please refer to the attachment for steps that the user can take to avoid the possible risks associated with the mentioned issue. | 2023-08-23 11:39
Siemens Healthcare GmbH | MAGNETOM Avanto | 148 | - | In some cases when lubricating grease is present within the headband of the over ear headphones it may cause visible dot or line shaped fat isointense artifacts during head examinations only with the MAGNETOM MRI systems listed in the attachment x D | Please refer to the attachment for steps that the user can take to avoid the possible risks associated with the mentioned issue. | 2023-08-23 11:39
Siemens Healthcare GmbH | MAGNETOM Skyra fit | 148 | - | In some cases when lubricating grease is present within the headband of the over ear headphones it may cause visible dot or line shaped fat isointense artifacts during head examinations only with the MAGNETOM MRI systems listed in the attachment x D | Please refer to the attachment for steps that the user can take to avoid the possible risks associated with the mentioned issue. | 2023-08-23 11:39
Siemens Healthcare GmbH | MAGNETOM Vida | 148 | - | In some cases when lubricating grease is present within the headband of the over ear headphones it may cause visible dot or line shaped fat isointense artifacts during head examinations only with the MAGNETOM MRI systems listed in the attachment x D | Please refer to the attachment for steps that the user can take to avoid the possible risks associated with the mentioned issue. | 2023-08-23 11:39
Siemens Healthcare GmbH | MAGNETOM Avanto Fit | 148 | - | In some cases when lubricating grease is present within the headband of the over ear headphones it may cause visible dot or line shaped fat isointense artifacts during head examinations only with the MAGNETOM MRI systems listed in the attachment x D | Please refer to the attachment for steps that the user can take to avoid the possible risks associated with the mentioned issue. | 2023-08-23 11:39
Siemens Healthcare GmbH | MAGNETOM Vida Fit | 148 | - | In some cases when lubricating grease is present within the headband of the over ear headphones it may cause visible dot or line shaped fat isointense artifacts during head examinations only with the MAGNETOM MRI systems listed in the attachment x D | Please refer to the attachment for steps that the user can take to avoid the possible risks associated with the mentioned issue. | 2023-08-23 11:39
Siemens Healthcare GmbH | MAGNETOM Lumina | 148 | - | In some cases when lubricating grease is present within the headband of the over ear headphones it may cause visible dot or line shaped fat isointense artifacts during head examinations only with the MAGNETOM MRI systems listed in the attachment x D | Please refer to the attachment for steps that the user can take to avoid the possible risks associated with the mentioned issue. | 2023-08-23 11:39
Siemens Healthcare GmbH | MAGNETOM Sola | 148 | - | In some cases when lubricating grease is present within the headband of the over ear headphones it may cause visible dot or line shaped fat isointense artifacts during head examinations only with the MAGNETOM MRI systems listed in the attachment x D | Please refer to the attachment for steps that the user can take to avoid the possible risks associated with the mentioned issue. | 2023-08-23 11:39
Siemens Healthcare GmbH | MAGNETOM Prisma | 148 | - | In some cases when lubricating grease is present within the headband of the over ear headphones it may cause visible dot or line shaped fat isointense artifacts during head examinations only with the MAGNETOM MRI systems listed in the attachment x D | Please refer to the attachment for steps that the user can take to avoid the possible risks associated with the mentioned issue. | 2023-08-23 11:39
Siemens Healthcare GmbH | MAGNETOM Sola Fit | 148 | - | In some cases when lubricating grease is present within the headband of the over ear headphones it may cause visible dot or line shaped fat isointense artifacts during head examinations only with the MAGNETOM MRI systems listed in the attachment x D | Please refer to the attachment for steps that the user can take to avoid the possible risks associated with the mentioned issue. | 2023-08-23 11:39
Siemens Shenzhen Magnetic Resonance Ltd.(SSMR) | MAGNETOM Amira | 148 | - | In some cases when lubricating grease is present within the headband of the over ear headphones it may cause visible dot or line shaped fat isointense artifacts during head examinations only with the MAGNETOM MRI systems listed in the attachment x D | Please refer to the attachment for steps that the user can take to avoid the possible risks associated with the mentioned issue. | 2023-08-23 11:39
Arthrex INC | Knee Scorpion | 6 | - | wrong labelling | return the product with the wrong label or sterilize the product before use. | 2023-08-23 12:41
Hamilton Medical | Hamilton -C1/T1/MR1 | -14 | - | A degrading capacitor on the control board might leak electrolyte onto the control x D board causing a short circuit on the board and or the capacitor might lose its x D function The primary function of this capacitor is to provide energy to store x D Realtime Clock information in case of total loss of power external AC supply and x D battery for example during storage x D As a result of the short circuit affected HAMILTON C T MR ventilators might x D switch to the Ambient State x D | 1. Provide alternative ventilation immediately. 2. You must switch off ventilator power to exit the “Ambient State”. 3. Once patient safety is ensured, it is required to have the ventilator serviced. 4-For affected devices listed in “FSCA 2023-07-03 List of Affected Devices” identify the end customer, exchange the control board . | 2023-08-23 14:05
Hamilton Medical | Hamilton C1/T1/MR1 | 14 | - | A degrading capacitor on the control board might leak electrolyte onto the control x D board causing a short circuit on the board and or the capacitor might lose its x D function The primary function of this capacitor is to provide energy to store x D Realtime Clock information in case of total loss of power external AC supply and x D battery for example during storage | 1. Provide alternative ventilation immediately. 2. You must switch off ventilator power to exit the “Ambient State”. 3. Once patient safety is ensured, it is required to have the ventilator serviced. 4- For affected devices listed in “FSCA 2023-07-03 List of Affected Devices” identify the end customer, exchange the control board | 2023-08-23 14:17
Boston Scientific Corporation | Guider Softip XF 5F | 663 | - | Stryker Neurovascular has observed that certain lots of Guider SoftipTM XF Guide x D Catheters appear to have the incorrect tip curve shape The impacted products were x D distributed with a tip curve shape which differs from the labeled tip curve shape | 1. Immediately check your internal inventory for affected devices. 2. Segregate the affected units in a secure location for return to Stryker. 3. Circulate this Recall-Removal notice internally to all interested/affected parties. 4. Maintain awareness of this communication internally until all required actions have been completed within your facility. 5. Inform Stryker if any of the subject devices have been distributed to other organizations. If yes, provide contact details so that Stryker can inform the recipients appropriately. 6. Please inform Stryker of any adverse events concerning the use of the subject devices. 7. Complete the attached customer response form. It may be that you no longer have any physical inventory on site. Completing this form will allow us to update our records and will also negate the need for us to send any further unnecessary communications on this matter. 8. Email the completed form to your local Stryker contact. 9. Product Return Information will be provided to you by your designated Sales Representative. | 2023-08-24 09:58
Boston Scientific Corporation | Guider Softip XF 6F | 663 | - | Stryker Neurovascular has observed that certain lots of Guider SoftipTM XF Guide x D Catheters appear to have the incorrect tip curve shape The impacted products were x D distributed with a tip curve shape which differs from the labeled tip curve shape | 1. Immediately check your internal inventory for affected devices. 2. Segregate the affected units in a secure location for return to Stryker. 3. Circulate this Recall-Removal notice internally to all interested/affected parties. 4. Maintain awareness of this communication internally until all required actions have been completed within your facility. 5. Inform Stryker if any of the subject devices have been distributed to other organizations. If yes, provide contact details so that Stryker can inform the recipients appropriately. 6. Please inform Stryker of any adverse events concerning the use of the subject devices. 7. Complete the attached customer response form. It may be that you no longer have any physical inventory on site. Completing this form will allow us to update our records and will also negate the need for us to send any further unnecessary communications on this matter. 8. Email the completed form to your local Stryker contact. 9. Product Return Information will be provided to you by your designated Sales Representative. | 2023-08-24 09:58
Boston Scientific Corporation | Guider Softip XF 7F | 663 | - | Stryker Neurovascular has observed that certain lots of Guider SoftipTM XF Guide x D Catheters appear to have the incorrect tip curve shape The impacted products were x D distributed with a tip curve shape which differs from the labeled tip curve shape | 1. Immediately check your internal inventory for affected devices. 2. Segregate the affected units in a secure location for return to Stryker. 3. Circulate this Recall-Removal notice internally to all interested/affected parties. 4. Maintain awareness of this communication internally until all required actions have been completed within your facility. 5. Inform Stryker if any of the subject devices have been distributed to other organizations. If yes, provide contact details so that Stryker can inform the recipients appropriately. 6. Please inform Stryker of any adverse events concerning the use of the subject devices. 7. Complete the attached customer response form. It may be that you no longer have any physical inventory on site. Completing this form will allow us to update our records and will also negate the need for us to send any further unnecessary communications on this matter. 8. Email the completed form to your local Stryker contact. 9. Product Return Information will be provided to you by your designated Sales Representative. | 2023-08-24 09:58
Boston Scientific Corporation | Guider Softip XF 8F | 663 | - | Stryker Neurovascular has observed that certain lots of Guider SoftipTM XF Guide x D Catheters appear to have the incorrect tip curve shape The impacted products were x D distributed with a tip curve shape which differs from the labeled tip curve shape | 1. Immediately check your internal inventory for affected devices. 2. Segregate the affected units in a secure location for return to Stryker. 3. Circulate this Recall-Removal notice internally to all interested/affected parties. 4. Maintain awareness of this communication internally until all required actions have been completed within your facility. 5. Inform Stryker if any of the subject devices have been distributed to other organizations. If yes, provide contact details so that Stryker can inform the recipients appropriately. 6. Please inform Stryker of any adverse events concerning the use of the subject devices. 7. Complete the attached customer response form. It may be that you no longer have any physical inventory on site. Completing this form will allow us to update our records and will also negate the need for us to send any further unnecessary communications on this matter. 8. Email the completed form to your local Stryker contact. 9. Product Return Information will be provided to you by your designated Sales Representative. | 2023-08-24 09:58
Berlin Heart GmbH | Apex Cannula | 54 | - | Berlin Heart urges you to pay close attention to the integrity of the cannula and to situations in which incorrect use or improper handling of EXCOR cannulae may result in partial or complete breaches on the surface of the cannula x D x D It is imperative that if a breach occurs the cannula should be trimmed per the directions provided in the Instructions for Use EN x Edition chapter Identifying and Rectifying Errors Tab Possible error scenarios Cannula damaged by trained medical professionals x D x D While root cause investigations are still ongoing it is important to recognize possible hazardous scenarios described in attached appendix and what actions should be completed immediately x D | What Actions Berlin Heart is Asking You to Take: 1. EXCOR cannulae should be inspected every four hours by a healthcare provider. (As defined in our Instructions for Use in Chapter 14). In particular, care must be taken to observe whether any sign of a cannula breach is present. Special attention should be given to the titanium connector -- cannula junction as this location is the most common site of a breach. Signs of cannula breaches are as follows: a. Cuts, tears, or any compromise to the integrity of cannula b. Discoloration 2. Special care should be taken to ensure proper placement of cable ties at the titanium• connector/ cannula junction, that the velour wrap leaves at least 5 cm (2 inches) of cannula exposed at the titanium-connector/ cannula junction and that the cannula does not come into contact with sharp objects. Please refer to the Appendix for images of proper placement of cable ties and velour wraps. 3. The correction does not require removal of the device from use. Please note, that an exchange of your EXCOR components is not necessary unless the cannula is showing signs of breaching. 4. Please complete the included response form acknowledging receipt and understanding of recall communication. Please return the response form to vigilance@berlinheart.de within 7 days of receipt. What to Do if the Issue Occurs: In the event of a cannula breach, • Immediately stop the support by disconnecting the driving tube from driving unit, • clamp the cannula and • trim the cannula proximal to the breached or damaged area. • reconnect the cannula to the blood pump. | 2023-08-24 10:21
Philips | Xper Flex Cardio Physiomonitoring system | 81 | - | Following exposure to defibrillation voltage the system returns to normal operating mode within seconds However if the Xper Flex Cardio Device has been exposed to greater than the specified energy related to the defibrillator and or ESU use it might fail to return to normal operating mode within seconds and it may require a restart in order to continue ECG monitoring While this could result in a delay in treatment the risk of harm is unlilkely x D The specifications of the Xper Flex Cardio Deviceregarding the absorption of energy into the device are noted in the Instructions For Use IFU Additionally the IFU contains warnings and best practices when used in conjunction with defibrillators and ESU This design limitation could result in an inability of the Xper Flex Cardio Deviceto recover and may need the user to restart the device to continue ECG monitoring x D | Philips is distributing this Field Safety Notice to the affected customers / users to highlight the instructions specified in the IFU. | 2023-08-28 11:11
Philips | Philips Hemo system | 81 | - | Following exposure to defibrillation voltage the system returns to normal operating mode within seconds However if the Xper Flex Cardio Device has been exposed to greater than the specified energy related to the defibrillator and or ESU use it might fail to return to normal operating mode within seconds and it may require a restart in order to continue ECG monitoring While this could result in a delay in treatment the risk of harm is unlilkely x D The specifications of the Xper Flex Cardio Deviceregarding the absorption of energy into the device are noted in the Instructions For Use IFU Additionally the IFU contains warnings and best practices when used in conjunction with defibrillators and ESU This design limitation could result in an inability of the Xper Flex Cardio Deviceto recover and may need the user to restart the device to continue ECG monitoring x D | Philips is distributing this Field Safety Notice to the affected customers / users to highlight the instructions specified in the IFU. | 2023-08-28 11:11
Alba Bioscience Limited | ORTHO Sera Anti-Lea Murine Monoclonal IgM. | 206 | - | A Customer Complaint investigation has confirmed that the lots of ORTHOTM Sera Anti Lea x D stated above show unexpected negative results with Lea antigen positive red cells This x D phenomenon was noted during testing of these lots against reagent red blood cells and x D EDTA red cells positively expressing the Lea antigen At time of product release all lots x D met release criteria however during in house testing we have noted deterioration in x D performance over shelf life leading to the potential for failure to detect red cells with x D reduced expression of the Lea antigen | The reagents listed above should be removed from use and discarded. Quotient is recommending that a review of previous donor and patient testing performed with these lots of antisera should be conducted. • Repeat testing should be conducted using an alternative lot of antisera if previously obtained results appear anomalous. • All required regulatory notifications are being completed. We request that you help us meet our compliance requirements. Please perform the required actions: complete and sign the attached form and return it to Quotient at your earliest opportunity by email: Vigilance.Notifications@quotientbd.com. • Refer to the enclosed Confirmation of Receipt form for product replacement and credit. | 2023-08-29 09:51
Ohmeda Medical, a Division of Datex-Ohmeda, Inc. | Giraffe OmniBed Carestation CS1 | 193 | - | On January th Border Operations BOP during the inspection of wall latches identified the west sidewalls of Giraffe OmniBed s GOB had incorrect latches OmniBed Sidewalls P N West Acrylic Sidewall Assembly OmniBed identified with north latch P N corresponding to GINC Sidewalls north latch instead of P N corresponding to GOB Sidewalls north latch as required per drawing ADW Rev x D | Customer letter will be shipped to all affected consignees to inspect all their Omnibed units and if incorrect latch is observed need to quarantine the unit Replace incorrect latch on affected products with a new latch. | 2023-08-30 17:13
Abbott | Apical Coring Knife and HeartMate 3 LVAD Kits | - | - | Abbott is notifying customers of a finding with the HeartMate Apical Coring Knife provided with x D the HeartMate Left Ventricular Assist System LVAS kits Model INT and the Apical x D Coring Knife Model separately distributed for use during HeartMate implant x D procedures Note that distributed HeartMate II Implant Kits are not affected by this notice x D Starting in April Abbott has received complaints regarding inability of the Apical Coring x D Knife to start and or complete resecting heart tissue when the physician was preparing the x D patient s heart for implantation of the Left Ventricular Assist Device LVAD As of August x D a total of eight complaints had been reported about this issue corresponding to an x D overall complaint rate of about two percent There has been no patient harm reported x D beyond extended surgical time procedure time with these complaints In all these events the x D physician was able to complete the procedure by using an Apical Coring Knife from the backup x D LVAS implant kit or an alternate surgical tool | At this time, this product is not being removed from the field and does not need to be returned. While Abbott continues to investigate the root cause of the issue and implements appropriate corrective actions, clinicians can continue to use the Apical Coring Knife provided with the HeartMate 3 LVAS kits and the separately distributed Apical Coring Knives listed above at their discretion until further notice. Abbott is reinforcing the following warning in the Instructions for Use when utilizing the Apical Coring Knife: • During the implant process, a complete backup system (implant kit and external components) must be available on-site and in close proximity for use in the event of an emergency. Additionally, if unusual resistance is encountered during the coring procedure, immediately stop coring and complete using the backup Apical Coring Knife (from the LVAS kits or a separately distributed Apical Coring Knife, if available). Please continue to report any difficulties you encounter with the Apical Coring Knife | 2023-09-04 19:24
Fresenius Kabi AG | AmiCORE Apheresis Kit – Single Needle with Two Platelet Containers and PAS Connector | 350 | - | The saline and return lines were swapped at the connection to the AmiCORE Cassette | The root cause is related to the operator inadvertently not performing the assembly operation as described in the work instructions. Due to this reason the mis-assembly occurred between the return line and the saline line. The assembled products are checked at the end of the assembly line, but the issue was not detected at the inspection station. | 2023-09-05 12:49
Abbott | GALLANT™, NEUTRINO™ NxT, AND ENTRANT™ ICDs AND CRT-Ds | 1488 | - | Abbott is informing customers of a rare potential for a Bluetooth BLE circuit component issue on a subset of Gallant x D Neutrino and Entrant Implantable Cardioverter Defibrillators ICDs and Cardiac Resynchronization Therapy x D Defibrillators CRT Ds manufactured prior to April This issue has been associated with a risk of loss of Bluetooth x D communication and therefore a loss of remote monitoring higher than normal current consumption and reduced device x D longevity | Abbott has developed upgraded device firmware version pr00.10.87.04, which eliminates the potential for devices affected by this issue from entering the high current consumption mode should the Bluetooth (BLE) circuit component issue occur | 2023-09-07 11:03
Abbott | EnSite™ X EP System | - | - | The issue is isolated to EnSite X EP Systems using version software when used in x D conjunction with a TactiFlex Ablation Catheter The issue does not occur with EnSite X x D EP Systems using version software or lower or when using other catheters | Abbott is working to correct the software causing this issue. Abbott will notify you when a software upgrade is available and schedule the installation of the software upgrade. | 2023-09-10 11:23
Werfen Australia Pty Ltd | ACL TOP SW 6.4.0. | 31 | - | There is the potential for sample misidentification under certain unusual circumstances identified on the ACL TOP Family Series software versions or later due to an unexpected behaviour that has occurred in an ACL TOP SW SN when a rack ID was inserted in position S the last tube was not recognised and the position was shown in grey x D A certain number of conditions must be true for this rare occurrence to occur and the occurrence would not alert the user to the misidentification All the following conditions must be true x D User introduces a sample rack in any of the sample track positions and all samples are properly identified x D The barcode reader has not moved to another position x D The user decides to remove the rack and swap one of the samples with a sample that has a different ID x D The user reintroduces the sample rack in the same sample track | Based on the above please take one of the following immediate workflows to mitigate this issue. To avoid the potential sample misidentification, users should not change any samples in the sample rack that have been properly identified. If the user decides they must remove a sample from a sample rack that has been properly identified by the ACL TOP Family 50 Series/ACL TOP 970 CL Instrument, they may do so only after the bar code reader has returned to the home position or the samples have been completely resulted. Alternatively, a user may remove a properly identified sample from a sample rack and must leave that sample rack position empty. • Share this information with your laboratory staff and update your internal procedures, as needed. • Forward this notification to all affected locations in your facility. • Post this notification on all affected systems in your facility. • Retain a copy of this notification for your records. • Complete and return the Response Tracking Form to acknowledge the above actions. Based on the above please take one of the following immediate workflows to mitigate this issue. To avoid the potential sample misidentification, users should not change any samples in the sample rack that have been properly identified. If the user decides they must remove a sample from a sample rack that has been properly identified by the ACL TOP Family 50 Series/ACL TOP 970 CL Instrument, they may do so only after the bar code reader has returned to the home position or the samples have been completely resulted. Alternatively, a user may remove a properly identified sample from a sample rack and must leave that sample rack position empty. • Share this information with your laboratory staff and update your internal procedures, as needed. • Forward this notification to all affected locations in your facility. • Post this notification on all affected systems in your facility. • Retain a copy of this notification for your records. • Complete and return the Response Tracking Form to acknowledge the above actions. | 2023-09-10 15:43
NOUVAG AG | Control unit Dispenser DP 30 | 1 | - | On July Medtronic received a Field Safety Notification FSN from Nouvag AG PVH OEM related to the DP Infiltration Pumps distributed by Medtronic Specifically Nouvag received a finding from their Notified Body that they did not have clinical evidence to support pediatric indications Nouvag is conducting a Field Corrective Action to x D Inform customers and partners about the FSN x D Forward a revised IFU to customers removing pediatric patient population indications x D | • Please distribute this notice to all those who need to be aware within your organization or to any organization where the potentially affected product may have been transferred. • Complete the Customer Acknowledgement Form for the affected Dispenser DP 30 (REF 4180 & 4187) product that your account has received. Return the completed form to Medtronic. • Please maintain a copy of the Nouvag Field Safety Notification letter for your records. | 2023-09-10 16:39
Siemens Medical Solution USA, Inc | ACUSON Redwood | 9 | - | If a user generated preset for an L transducer created on an ACUSON Redwood system is used with an ACUSON Redwood system the ACUSON Redwood system will display underestimated measurement results when using an L transducer and viewing in the Dual format side by side visualization function x D This issue does NOT affect the ACUSON Redwood system and also does NOT affect the ACUSON Redwood system as long as the user generated presets for L transducers are NOT carried over from an ACUSON Redwood system | This issue will be solved with a free-of-charge software update to your ACUSON Redwood ultrasound system. The software update will correct the compatibility of pre-existing user-generated presets. | 2023-09-11 11:28
Dr. Schumacher GmbH | DESCOSEPT SPEZIAL | 12 | - | found contamination at a low level in an individual case | There is no increased risk for products already used, as this concerns contamination at a very low level and the germ does not pose a risk of ubiquitous occurrence. The returned samples will be examined. We will inform you promptly about the results. | 2023-09-11 13:25
Shenzhen Mindray Bio-medical Electronics Co., Ltd. | Ventilator | 90 | - | It has been brought to our attention that in the situation of continuous long period x D closed system suction some specific uncommon operations may cause severe high x D negative airway pressure that affect the ventilator status and alarms will be triggered to x D notify the medical staff to intervene After the suction is stopped and tube disconnected x D the ventilation can restore automatically | To prevent above issue, Mindray has released an enhancement software version, which allows the ventilator to reduce the high negative pressure by pressure release mechanisms and to maintain ventilation. | 2023-09-11 15:54
SAUDI MAIS CO FOR MEDICAL PRODUCTS | Sodium Bicarbonate | 31440 | - | leakage from neck side of the bi bag | It is advised to stop the use of affected product immediately and quarantine. Inform to our customer services as our customer services will collect the affected products. We will be replacing the remaining effected quantities with the product with a new Lot. | 2023-09-12 09:25
GE Healthcare Finland Oy, Kuortaneenkatu 2, Helsinki, FI-00510, Finland | F2-01 | 3 | - | If the F frame has not been powered down completely for more than days it can cause the data communication interruption between CARESCAPE Canvas CARESCAPE ONE and E modules that are connected through F frame If this situation were to occur it could cause a partial loss of monitoring potentially causing a delay in treatment x D | CA1 to RC1: Design Change to update F2-01 software version to EMBC_PRO_1.16.0 to fix the software coding error CA2 to RC2: Update F2-01 EMBC Pro Software Design Specification to add the special interrupt source design requirement CA3 to RC1: FMI# 36160 initiated for this F2-01 run time issue to update the affected units | 2023-09-12 13:31
GE Medical Systems Ultrasound & Primary Care Diagnostics LLC, 9900 Innovation Drive, Wauwatosa WI, 53226, USA | Venue Go | 9 | - | GE HealthCare has become aware that some Venue Go Standard Carts can develop an internal failure of the vertical horizontal tilt adjustment mechanism which can result in the Venue Go system detaching from the cart and falling | Venue Go Cart tilt replacement for all reported complaints | 2023-09-12 14:25
Becton Dickinson and Company | BD BACTEC™ Plus Aerobic/F Culture Vials | - | - | BD has confirmed through two customer complaints that certain vials of the BD BACTEC Plus x D Aerobic F Culture Vials noted above have a labeling issue where duplicate barcode sequence x D numbers have been identified on more than one vial x D An internal collision analysis revealed that it is unlikely that a customer site will receive vials with x D duplicate sequence numbers This is based on the distribution of sequence number label batches for x D the affected lots and the regional distribution patterns Although unlikely this issue could cause a x D customer to receive more than one BD BACTEC Plus Aerobic F Culture Vials with duplicate x D sequence numbers | BD Actions: BD is investigating and will implement appropriate measures to prevent recurrence of this product issue. Customer Actions:  Review the information in Table 1 to determine if BD BACTEC™ Plus Aerobic/F Culture Vials in your possession are impacted.  Users should verify that the accession number on the BD BACTEC™ Instrument loading screen matches the vial's accession number while scanning and loading onto the instrument. If an error message or other system action occurs when loading vials into the BD BACTEC™ Instrument that indicates a potential issue with duplicate sequence numbers, it is recommended to follow the instructions within the message and to view the table of “System Alerts” located in “7 – Troubleshooting” of the BD BACTEC™ FX Instrument User’s Manual before taking any further action in vial processing. The link to the BD BACTEC™ FX Instrument User’s Manual is as follows: https://eifu.bd.com/en/search/searchresults? term=441385&cats=9444_,9441_9443,9441_9442,1185_,1185 There are no additional follow-up activities recommended for processed patient samples.  Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 9th October 2023.  Circulate this notice to all those who need to be aware within your organisation or to any organisation where the potentially affected product has been transferred.  If you experience any issues, please report as a complaint as per your normal process. Distributor Actions:  Review the information in Table 1 to determine if the BD BACTEC™ Plus Aerobic/F Culture Vials in your possession are impacted.  Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 9th October 2023.  Complete and return the Customer Response Form following completion of your reconciliation activities.  If you experience any issues, please report as a complaint as per your normal process. | 2023-09-14 01:43
Siemens Healthcare GmbH | Ysio Max | 41 | - | Following a complaint it was identified that during system installation the mounting of the BABIX holder was possibly performed incorrectly | The field corrective action (Update XP047/23/S) will initiate an onsite visit of a service engineer to check the mounting for the BABIX holder on the Bucky wall stand and if necessary, mount it correctly. Update XP047/23/S is planned to be released in October 2023 and will be provided free of charge. | 2023-09-14 15:07
Abbott Laboratories | Multichem S Plus, Multichem IA Plus , Multichem P | 4412 | - | As per attached letter | As per attached letter | 2023-09-14 15:51
iLine Microsystems S.L. | MicroINR easy control | 9 | - | We have detected an error in the QR code on the label attached to the microlNR EasyControl instructions for use When this code is read with the microINR Expert barcode reader a message of expired lot is displayed | We attach to this communication an annex with the correct QR code for the affected microlNR EasyControl lot you have received during 2023. If any client is using this lot with the microINR Expert Meter, they should read the QR code provided now instead of the one on the instructions for use. The microINR Expert Meter has also the option of using the microINR EasyControl as an external quality control material and in this case a QR code is not required. Please, share with any final user using the microINR EasyControl with microlNR Expert meters the new QR code we provide in this communication. | 2023-09-17 11:01
Roche Diabetes Care GmbH Sandhofer Strasse 116 68305 Mannheim, Germany | Accu-Chek Solo | - | - | In the previous FSN was related to the Accu Chek Solo system diabetes manager In the updated one we have added information that the device may be affected if it is not used correctly according to the updated instructions of use attached just to highlight the correct way to use the items ONLY | We decided to replace the affected devices with new ones that include the updated software to facilitate the process for the patients instead of collecting the affected devices from the patients to update the software and return it to them again | 2023-09-18 10:13
BALT EXTRUSION SAS | HYBRID WIRE | 9 | - | The investigations showed that the root cause is related to label design issues Specifically a recopying error of the guidewire drawing in the label | Procedure to be applied by distributors: • Inform your customers and your local competent authority about this notice • Complete and return the “Notice Receipt form” Procedure to be applied by the hospital staff: • Communicate this information to staff within the hospital that may use HYBRID12D300 references and lots stents or any other person if deemed necessary. • Complete and return the “Notice Receipt form”. By returning the completed Notice Receipt form by e-mail or mail, you acknowledge that you have read and understood this Field Safety Notice. | 2023-09-18 17:36
Alcon Laboratories Inc. | Alcon Phaco Tips (see Attachment 1) | 55088 | - | Alcon has detected an increase in complaints of plastic particles associated with Phaco Tips Laboratory analysis of retrieved particles revealed that the composition of the particles is consistent with the plastic wrench material Globally complaints of plastic particles associated with Phaco Tips represent less than of Phaco Tips sales volume | In order to mitigate the potential for the wrench to be damaged by overtightening the Phaco Tip and associated generation of plastic particles, Alcon is advising customers of the following precautions: • To fasten the tip into the handpiece, grasp the plastic wrench by the cylindrical part of the wrench. Overtightening by leveraging the wings can increase the likelihood that the wrench may slip over the tip-to-wrench interface, causing damage to the wrench, and has the potential to generate plastic particles which may remain on the exterior of the Phaco Tip and/or distal end of the phaco handpiece. • If the wrench slips, or you suspect that the wrench may have slipped, during the process of threading the tip to the handpiece, do not apply the infusion sleeve over the Phaco Tip. • Thoroughly flush the tip using the console irrigation functionality. • Inspect the Phaco Tip and distal end of the phaco handpiece for particle presence under magnification and illumination. • If particles are present on the Phaco Tip and/or phaco handpiece, use the console irrigation functionality to flush the tip again. • After flushing, inspect again to verify no particles are present. Upon confirming that the observed particle(s) is no longer present, proceed with application of the infusion sleeve and instrument setup. | 2023-09-19 11:20
Masimo Corporation | Radical-7 | - | - | Masimo has identified that certain batteries used in Radical devices may experience excessive swelling All x D batteries can swell over time However certain batteries were found to have expanded to the point where the x D battery cover may be displaced and or the Radical cannot be fully docked | Assess batteries inside Radical-7 devices to determine if batteries are experiencing excessive swelling. Any Radical-7 with a battery cover being pushed out or displaced is considered to have a battery with excessive swelling. If excessive swelling is not present, there is no additional action required. If excessive swelling is detected, contact Masimo to request replacement batteries. The replacement batteries will be delivered in a kit with instructions on how to properly install the new batteries | 2023-09-19 11:41
ResMed SAS | Astral 100 or Astral 150 ventilator | - | - | ResMed s Astral ventilator is designed to deliver a Total Power Failure TPF audible alarm for at x D least minutes to alert the carer clinician when the ventilator experiences a total loss of power that x D results in a stop of ventilation The TPF audible alarm is powered by an electrical component known x D as a supercapacitor To monitor performance this component is tested at each two year x D preventative maintenance x D ResMed is bringing to your attention an issue with the performance of the supercapacitor component x D This component degrades over time which may cause the TPF alarm to sound for less than x D minutes or in some cases to not sound at all This component does not affect any other alarm x D functionality including low and critically low battery alarms This issue does not affect therapy x D delivery x D To date ResMed has not received any reports of serious adverse events related to this issue | Actions by ResMed This Field Safety Corrective Action will consist of two (2) phases. 1. The first phase includes this communication to customers as well as an update to the Astral service process to improve detection of this issue. ResMed will provide replacement circuit boards for ventilators that are identified in service to be affected by this issue. 2. A second phase will be communicated to customers when further ventilator correction(s), currently under development, are available, with specific actions to follow. Actions to be taken by healthcare providers and distributors • Complete and return the acknowledgment form. • Immediately provide a copy of the patient/carer letter to all patients and carers. • Immediately provide a copy of this notice and the patient/carer letter to relevant healthcare professionals. • Consistently follow the 2-year preventative maintenance schedule of the Astral ventilator to help detect and replace circuit boards with degraded supercapacitors. • Immediately prioritise the return of ventilators that are due for preventative maintenance, paying particular attention to those used on ventilator-dependent patients. Actions to be taken by healthcare providers/distributors that service Astral Ventilators • Forward this notice to the person(s) responsible for Service Centre operations and device maintenance. • Follow the updated Astral Service process detailed in Tech Note 1063673. Note, ResMed Service Centres will also follow this updated process. • As per the instructions in Tech Note 1063673, enter all required information regarding service of affected ventilators to ensure replacement circuit boards are provided. | 2023-09-19 11:50
Cardiac Pacemaker Inc. | EMBLEM S-ICD | - | - | During an EMBLEM S ICD system impedance measurement low energy pulses are automatically sent every three days When this test is performed sensing is momentarily disabled to prevent non cardiac artifacts from being over sensed by the device or displayed on the S ECG | software update for the EMBLEM S-ICD | 2023-09-19 11:54
BD Switzerland Sàrl | BD Alaris™ Syringe Module, BD Alaris™ PCU Module and Alaris™ PCA Module | 4 | - | BD has become aware that Monoject syringe manufacturer Cardinal Health has rebranded its x D syringes from Covidien Monoject to Cardinal Health Monoject and retained the original Covidien x D SKU numbers Cardinal Health issued a notification to customers in March indicating that they x D were changing the manufacturing location and process for the syringes On September x D Cardinal Health issued a quality hold for their Cardinal Health Monoject syringes due to compatibility x D issues with certain syringe infusion pumps We understand that Cardinal Health is undertaking an indepth x D technical investigation of the issue x D BD s preliminary investigation has determined that there are differences in syringe dimensions x D associated with these changes that may impact performance with the BD Alaris Syringe and PCA x D Modules Therefore BD is issuing this notification to advise that the BD Alaris Syringe and PCA x D Modules have been validated for use with Covidien branded Monoject syringes BD has not x D validated the BD Alaris Syringe and PCA Modules for use with Cardinal Health branded Monoject x D syringes with the exception of one Cardinal Health Monoject branded syringe the detachable x D plunger empty barrel SKU which is unchanged and may continue to be used with the x D Alaris PCA Module subject to country availability regulatory status x D Please note that the BD Alaris Infusion System User Manuals warns against the use of incompatible x D syringe sizes and models The use of incompatible syringes can impact pump operation resulting in x D inaccurate fluid delivery delayed generation of occlusion alarms and other potential problems x D Figure below shows representative examples of the Covidien Monoject syringes compatible on x D the left and the Cardinal Health Monoject syringes not compatible on the right | BD Actions: 1. In order to maintain supply to customers, BD has increased production of BD syringe SKUs. 2. BD has updated the BD Alaris™ Syringe and PCA Compatibility Lists with Attachments 1 and 2 of this letter. 3. BD will continue to evaluate syringe compatibility and make updates as appropriate. Customer Actions:  You should cease use of Cardinal Health branded Monoject™ syringes with the BD Alaris™ PCU, Syringe and PCA Modules; except for SKU 8881135609, which is not impacted by this notice. o Note: The Cardinal Health Monoject™ branded syringe, the detachable plunger, empty barrel SKU 8881135609, is validated and may continue to be used with the Alaris™ PCA Module (subject to country availability/regulatory status).  Refer to Attachments 1 and 2 of this notice for a full selection of syringes that are compatible with the BD Alaris™ Syringe and PCA Module (subject to country availability/regulatory status). o For ordering assistance, please contact your local BD office/representative.  Replace your current BD Alaris™ Syringe and PCA Module Compatibility Lists with Attachments 1 and 2 of this notice.  Complete and return the Customer Response Form even if you no longer have/use BD Alaris™ PCU, Syringe and PCA Modules by 20th October 2023.  Circulate this notice to all those who need to be aware within your organisation or to any organisation where the BD Alaris™ PCU, Syringe and PCA Modules have been transferred.  If you experience any issues please report as a complaint as per your normal process. Distributor Actions:  Identify the facilities where you have distributed BD Alaris™ PCU, Syringe and PCA Modules and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 20th October 2023.  Complete and return the Customer Response Form to BD following completion of your reconciliation activities.  If you experience any issues please report as a complaint as per your normal process. | 2023-09-19 14:49
Zimmer Surgical Inc. | ZIMMER DERMATOME BLADES ZIMMER DERMATOME BLADES | 420 | - | skin grafts being thin and non uniform when using the affected blades | locate and quarantine affected products return all affected products | 2023-09-20 09:23
Draeger Medical Systems, Inc. | Babyroo TN300 | 1 | - | The TN Babyroo mattress tray MU may tilt without engaging the tilt mechanism when the mattress assembly is pushed down The tilt mechanism may not lock release correctly due to either loose or missing screws x D Note There have been no reports of adverse patient impact relating to this issue in the field x D The issue would be obvious to users or service personnel because when pushing on the mattress tray with minimal pressure the tilt mechanism will shift and the mattress will rapidly tilt | The attached Safety Notice is being issued to all customers that instructs the user on how to check the device to see if the mattress tray tilt mechanism is operational. | 2023-09-20 13:07
Randox Laboratories. | Liquid Urine Control Level 2 | - | - | Randox Laboratories is conducting a Device Recall for Liquid Urine Controls Level UC x D lot UC due to the following x D There is vial to vial variation resulting in some vials recovering positive for hCG x D which should be negative and high and outside range for cortisol x D There has been a transcription error for Creatinine in the Instructions For Use IFU x D The target and ranges for Creatinine for the Roche Creatinine Plus method have x D been listed incorrectly | -Discontinue use of and discard any of the above immediately. Provide Randox with photographic evidence of the destruction of the kits. • Review your inventory of these products and assess your laboratories needs for reimbursement for discarded inventory. • Review previous results for this controls and ensure patient results were not reported if control results were not within range - Complete and return the response form 12187-QA to technical.services@randox.com within five working days | 2023-09-20 14:00
Leica Microsystems (Schweiz) AG | PROVEO 8 Stand F42 | - | - | Leica Microsystems is issuing this field safety notice as a corrective action for some M OHX PROVIDO and Proveo Surgical Microscopes that were shipped between December st and August th x D This field safety notice supersedes the previous field safety notice regardless of the upgrade status of your affected system x D At Leica Microsystems we are dedicated to provide the highest standards of quality to our customers Part of this dedication is proactively responding to quality defects when they occur x D You are receiving this letter because according to our data you are the holder of a product that is subject to this field safety corrective action x D Description of the Problem x D Leica Microsystems has become aware of process weaknesses of the sub supplier producing the affected Power Supply Unit x D These process weaknesses have caused that the Power Supply can become overheated due to internal component overstressing x D Consequently due to the overstressing of the internal components one specific component can fail which will cause the device to unexpectedly shutdown x D This shutdown is caused by the interruption of the energy supply caused by the component defect inside the Power Supply Unit x D Leica Microsystems has performed a thorough investigation which revealed that the identified unusual accumulation of defects reported are related to devices delivered from December st till August th x D To date the complaints received by Leica Microsystems have not caused any actual serious injury to the patient or users x D As a matter of precaution Leica Microsystems has chosen to initiate a field safety corrective action to replace the affected Power Supply Unit on all potentially affected M OHX PROvido and Proveo surgical operating microscopes delivered between December st and August th The replacement part will ensure that the affected systems will not actually develop the failure and the system will function according to specification x D A Leica Microsystems representative will be contacting your facility to arrange this non cost field update of your instrument x D The field update will be scheduled to occur at your convenience within the next months | As a matter of precaution, Leica Microsystems has chosen to initiate a field safety corrective action to replace the affected Power Supply Unit on all potentially affected M530 OHX, PROvido and Proveo 8 surgical operating microscopes delivered between December 1st, 2021 and August 25th, 2023. The replacement part will ensure that the affected systems will not actually develop the failure and the system will function according to specification. A Leica Microsystems representative will be contacting your facility to arrange this non-cost field update of your instrument. The field update will be scheduled to occur at your convenience within the next 6 months. Advice on action to be taken by the user: • You can continue to use the system pending this field update. Based on analysis of data available to us, the probability is remote that an unacceptable patient risk will occur due to the unexpected shutdown of the system caused by the affected Power Supply Unit. Nonetheless, please consider the type and duration of any intended surgical procedure and evaluate whether the system can be deemed fit for such purpose. • However, please ensure that all users of the affected microscope are aware of the “Safety notes” as described in section 3 of the User Manual to follow the recommendation to have a backup device available and will follow the checklist “Checklist before the operation” – see excerpts of latest User Manual below: o Section 3 “Safety notes”: “Like any other instrument in the operating theater, this system may fail. Leica Microsystem (Schweiz) AG therefore recommends that a backup system be kept available during the operation.” o “Checklist before the operation”: “Connect the power cable. + Switch on the microscope.” | 2023-09-20 14:54
Leica Microsystems (Schweiz) AG | OHX Stand for M530 | - | - | Leica Microsystems is issuing this field safety notice as a corrective action for some M OHX PROVIDO and Proveo Surgical Microscopes that were shipped between December st and August th x D This field safety notice supersedes the previous field safety notice regardless of the upgrade status of your affected system x D At Leica Microsystems we are dedicated to provide the highest standards of quality to our customers Part of this dedication is proactively responding to quality defects when they occur x D You are receiving this letter because according to our data you are the holder of a product that is subject to this field safety corrective action x D Description of the Problem x D Leica Microsystems has become aware of process weaknesses of the sub supplier producing the affected Power Supply Unit x D These process weaknesses have caused that the Power Supply can become overheated due to internal component overstressing x D Consequently due to the overstressing of the internal components one specific component can fail which will cause the device to unexpectedly shutdown x D This shutdown is caused by the interruption of the energy supply caused by the component defect inside the Power Supply Unit x D Leica Microsystems has performed a thorough investigation which revealed that the identified unusual accumulation of defects reported are related to devices delivered from December st till August th x D To date the complaints received by Leica Microsystems have not caused any actual serious injury to the patient or users x D As a matter of precaution Leica Microsystems has chosen to initiate a field safety corrective action to replace the affected Power Supply Unit on all potentially affected M OHX PROvido and Proveo surgical operating microscopes delivered between December st and August th The replacement part will ensure that the affected systems will not actually develop the failure and the system will function according to specification x D A Leica Microsystems representative will be contacting your facility to arrange this non cost field update of your instrument x D The field update will be scheduled to occur at your convenience within the next months | As a matter of precaution, Leica Microsystems has chosen to initiate a field safety corrective action to replace the affected Power Supply Unit on all potentially affected M530 OHX, PROvido and Proveo 8 surgical operating microscopes delivered between December 1st, 2021 and August 25th, 2023. The replacement part will ensure that the affected systems will not actually develop the failure and the system will function according to specification. A Leica Microsystems representative will be contacting your facility to arrange this non-cost field update of your instrument. The field update will be scheduled to occur at your convenience within the next 6 months. Advice on action to be taken by the user: • You can continue to use the system pending this field update. Based on analysis of data available to us, the probability is remote that an unacceptable patient risk will occur due to the unexpected shutdown of the system caused by the affected Power Supply Unit. Nonetheless, please consider the type and duration of any intended surgical procedure and evaluate whether the system can be deemed fit for such purpose. • However, please ensure that all users of the affected microscope are aware of the “Safety notes” as described in section 3 of the User Manual to follow the recommendation to have a backup device available and will follow the checklist “Checklist before the operation” – see excerpts of latest User Manual below: o Section 3 “Safety notes”: “Like any other instrument in the operating theater, this system may fail. Leica Microsystem (Schweiz) AG therefore recommends that a backup system be kept available during the operation.” o “Checklist before the operation”: “Connect the power cable. + Switch on the microscope.” | 2023-09-20 14:54
Leica Microsystems (Schweiz) AG | PROVIDO | - | - | Leica Microsystems is issuing this field safety notice as a corrective action for some M OHX PROVIDO and Proveo Surgical Microscopes that were shipped between December st and August th x D This field safety notice supersedes the previous field safety notice regardless of the upgrade status of your affected system x D At Leica Microsystems we are dedicated to provide the highest standards of quality to our customers Part of this dedication is proactively responding to quality defects when they occur x D You are receiving this letter because according to our data you are the holder of a product that is subject to this field safety corrective action x D Description of the Problem x D Leica Microsystems has become aware of process weaknesses of the sub supplier producing the affected Power Supply Unit x D These process weaknesses have caused that the Power Supply can become overheated due to internal component overstressing x D Consequently due to the overstressing of the internal components one specific component can fail which will cause the device to unexpectedly shutdown x D This shutdown is caused by the interruption of the energy supply caused by the component defect inside the Power Supply Unit x D Leica Microsystems has performed a thorough investigation which revealed that the identified unusual accumulation of defects reported are related to devices delivered from December st till August th x D To date the complaints received by Leica Microsystems have not caused any actual serious injury to the patient or users x D As a matter of precaution Leica Microsystems has chosen to initiate a field safety corrective action to replace the affected Power Supply Unit on all potentially affected M OHX PROvido and Proveo surgical operating microscopes delivered between December st and August th The replacement part will ensure that the affected systems will not actually develop the failure and the system will function according to specification x D A Leica Microsystems representative will be contacting your facility to arrange this non cost field update of your instrument x D The field update will be scheduled to occur at your convenience within the next months | As a matter of precaution, Leica Microsystems has chosen to initiate a field safety corrective action to replace the affected Power Supply Unit on all potentially affected M530 OHX, PROvido and Proveo 8 surgical operating microscopes delivered between December 1st, 2021 and August 25th, 2023. The replacement part will ensure that the affected systems will not actually develop the failure and the system will function according to specification. A Leica Microsystems representative will be contacting your facility to arrange this non-cost field update of your instrument. The field update will be scheduled to occur at your convenience within the next 6 months. Advice on action to be taken by the user: • You can continue to use the system pending this field update. Based on analysis of data available to us, the probability is remote that an unacceptable patient risk will occur due to the unexpected shutdown of the system caused by the affected Power Supply Unit. Nonetheless, please consider the type and duration of any intended surgical procedure and evaluate whether the system can be deemed fit for such purpose. • However, please ensure that all users of the affected microscope are aware of the “Safety notes” as described in section 3 of the User Manual to follow the recommendation to have a backup device available and will follow the checklist “Checklist before the operation” – see excerpts of latest User Manual below: o Section 3 “Safety notes”: “Like any other instrument in the operating theater, this system may fail. Leica Microsystem (Schweiz) AG therefore recommends that a backup system be kept available during the operation.” o “Checklist before the operation”: “Connect the power cable. + Switch on the microscope.” | 2023-09-20 14:54
Randox Laboratories | Liquid Protein Calibrators | - | - | As part of our ongoing quality monitoring Randox Laboratories have restandardised Ferritin in Liquid x D Protein Calibrators IT to reference material NISBC Current calibrators lots IT x D IT packed into batches and have been reassigned as part of the x D restandardisation x D Ferritin results for Quality Control material and patient samples is expected to recover higher by x D approximately across the assay range following this restandardisatjon Updated calibrator x D targets for the batches have been listed in Tablet below The updated Instructions For Use IFU is x D available on www randox com please discard the previous version of the IFU and download the latest x D version | Action to be taken: • Review results generated with the affected batches in line with the clinical profile of the patient. • Discuss the contents of this notice with your Medical Director. • Complete and return the response form 12187-QA to technical.services@randox.com within five working days. Transmission of Field Safety Notice: Send a copy of the FSN to all affected customers and to those who need to be aware within your organisation | 2023-09-20 15:22
Randox Laboratories | Human Assayed Multi-Sera Control Level 3 | - | - | Randox Laboratories can confirm that the stability for AST GOT has been reduced from x D days to days when stored at C for Human Assayed Multi Sera Control Level HE x D lot UE The Instructions For Use IFU have been updated and is now available on x D www randox com Please discard any previous versions of the IFU and download the updated x D IFU This change applies to this lot only | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Please accept our apologies for any inconvenience caused. Thank you for your patience and understanding. If you have any questions or concerns please contact Randox Technical Services. | 2023-09-20 16:35
werfen (IL) | HemosIL von Willebrand Factor Antigen | 50 | - | This is a follow up to the original notification dated April that advised your facility of an issue with the labeled open vial C stability claim for HemosIL von Willebrand Factor Antigen PN x D Our investigation into reported customer complaints identified that the product was not meeting its labeled claim for Open vial stability at C of months A new open vial stability claims of days at C was established and cleared by the FDA x D Note There is no impact on patient samples as the data flag CE Baseline Average High is generated and no results are reported | The open vial stability claim has been reduced to 14 days for all in-date and future product lots of HemosIL von Willebrand Factor Antigen (Part No. 0020002300). Based on the above, please take the following immediate actions for HemosIL von Willebrand Factor Antigen (PN 0020002300): • Use the following updated open vial stability claim for HemosIL von Willebrand Factor Antigen (PN 0020002300). Current IFU Open Vial Stability at 2-8°C Reduced Open Vial Stability at 2-8°C 3 Months 14 Days • Share this information with your staff, notifying them of the updated open vial stability claim for HemosIL von Willebrand Factor Antigen (Part No. 0020002300). • Retain a copy of this letter in your files as a record of the notification. • Complete and return the enclosed Mandatory Response Tracking Form to acknowledge the above actions. | 2023-09-21 09:21
Randox Laboratories | Liquid Protein Calibrators | - | - | Reason for Action x D As part of our ongoing quality monitoring Randox Laboratories have restandardised Ferritin in Liquid x D Protein Calibrators IT to reference material NISBC Current calibrators lots IT x D IT packed into batches and have been reassigned as part of the x D restandardisation x D Ferritin results for Quality Control material and patient samples is expected to recover higher by x D approximately across the assay range following this restandardisatjon Updated calibrator x D targets for the batches have been listed in Tablet below The updated Instructions For Use IFU is x D available on www randox com please discard the previous version of the IFU and download the latest x D version | * Review results generated with the affected batches in line with the clinical profile of the patient. • Discuss the contents of this notice with your Medical Director. • Complete and return the response form 12187-QA to technical.services@randox.com within five working days. Transmission of Field Safety Notice: Send a copy of the FSN to all affected customers and to those who need to be aware within your organisation | 2023-09-25 10:21
Randox Laboratories | Immunoassay Specialty I Level 1 | - | - | Randox Laboratories can confirm that the OH Vitamin D in Immunoassay Speciality I Level x D IAS lot EC was assigned for Roche using the Roche nd generation Elecsys Vitamin x D D Total II Roche have since released Vitamin D Total II rd Generation reagent and EC x D recovers close to the bottom of the assay range for this reagent Randox cannot therefore x D provide target and ranges for this method x D A Product Information has been provided and is available on the Instructions For Use IFU x D available on www randox com | Action to be taken: Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2023-09-25 10:31
Randox Laboratories | RX Daytona Plus (with ISE), RX Daytona Plus (without ISE), RX Imola, RX Modena, RX Monaco | - | - | Randox are currently investigating RX Parameter rerun settings Sample Volume Low High x D and Dilutions used for Auto reruns on the following analysers RX Daytona Plus with ISE RX x D Daytona Plus without ISE RX Imola RX Modena and Sample Volume Decrement x D Increment on the RX Monaco analyser x D Results outside the upper and lower limit technical range will have an evaluation flag or x D added to the measurement result If the Automatic rerun function has been enabled the x D analyser will apply the rerun settings and perform the rerun function and results will be x D flagged with r Indicating that the result is a rerun measurement result x D x D Randox recommendations x D Auto rerun settings are disabled see Appendix A for guidance for each Rx instrument x D Results with flag reported as the bottom of the technical range Sample Volume x D settings Low should not be applied x D Results with flag reported as the top of the technical range Sample Volume x D settings high and dilutions should not be applied A manual dilution should be x D performed and results multiplied by the dilution factor x D results with r should not be reported x D The RX Parameter rerun settings will be updated with a parameter release in due course the x D Instructions For Use will also be updated at this time | - Discuss the contents of this notice with your Medical Director. • Complete and return the response form 12187-QA to technical.services@randox.com within five working day | 2023-09-25 15:03
Atrium Medical Corporation | Atrium Ocean, Oasis, and Express Single Collection Chest | 1123 | - | Atrium Getinge is initiating a voluntary Medical Device Correction for the Atrium Ocean Oasis and Express chest drains The Instructions for Use IFU for the Atrium Ocean Oasis and Express chest drains do not provide sufficient precaution instruction for proper set up of catheter s and patient line connections with single collection chamber chest drains No Devices Need to Be Returned | -Your facility can continue use of the device. No devices need to be returned. -Please ensure that all Atrium Ocean, Oasis, or Express Single Collection Chest Drain users at your facility are aware of this Safety Notice and post a copy of the Notice on Page 4 in all inventory locations within your facility where the devices are stored. o Notification of the release of the updated IFU containing the new Precaution will be communicated to all customers upon release, including reminder the Notice on Page 4 may be removed upon the facility’s receipt of product with the updated IFU. | 2023-09-26 23:59
Ortho Clinical Diagnostics Inc. | VITROS 3600 Immunodiagnostic System | 0 | - | Ortho Clinical Diagnostics QuidelOrtho received complaints related to this issue as of Jul we have received no complaints QuidelOrtho confirmed the issue and determined that certain Proboscis Piston Assembly or Assemblies manufactured between Apr Date code and Dec Date code may not perform as intended due to an out of specification barrel component This may cause VITROS VersaTips and or VITROS MicroTips to be improperly seated on the barrel component of the Assembly which may result in an increased frequency of condition codes related to the actions of tip pickup aspirate dispense mix tip seal and discard Due to the small VITROS MicroTip and VersaTip dimension variability within manufacturing tolerances and the way these Tips interact with the Assembly Assemblies manufactured within the affected date range may not perform as intended based on the Tips currently in use x D To determine if a laboratory is using an Assembly that was manufactured within the affected date range customers are instructed to check the date code stamped into the barrel of the Assembly Refer to the customer letter for an example image of how to locate the date code on the barrel component of the Assembly If a laboratory has an Assembly with a date code between the Assembly is considered affected and must be replaced QuidelOrtho will provide a credit OR no cost replacement for each affected Assembly This includes all affected Assemblies in the laboratory including spare Assemblies not yet installed on a VITROS System specifically for biased results related to this issue nor have we experienced biased results during our investigation | Determine if any of the impacted VITROS Systems in their laboratory are using an affected Proboscis/Piston Assembly by checking the date code stamped into the barrel component. Once a replacement Proboscis/Piston Assembly is received, discard all affected Assemblies within the date range of 040722 - 121922. • Determine if their inventory of VITROS spare parts contains any affected Proboscis/Piston Assemblies by checking the date code stamped into the barrel component. Once a replacement Proboscis/Piston Assembly is received, discard all affected Assemblies within the date range of 040722 - 121922. • Ensure their laboratory is performing the Leak/Hysteresis test during daily maintenance and performing Quality Control to ensure a properly functioning Proboscis/Piston Assembly until a replacement Assembly is available. • Customers must complete the enclosed Confirmation of Receipt form. Upon receiving their completed Confirmation of Receipt form, QuidelOrtho will credit their account or send a replacement Proboscis/Piston Assembly (Part. No. J55375), if applicable. | 2023-09-27 09:34
Ortho Clinical Diagnostics Inc. | VITROS 4600 Chemistry System | 0 | - | Ortho Clinical Diagnostics QuidelOrtho received complaints related to this issue as of Jul we have received no complaints QuidelOrtho confirmed the issue and determined that certain Proboscis Piston Assembly or Assemblies manufactured between Apr Date code and Dec Date code may not perform as intended due to an out of specification barrel component This may cause VITROS VersaTips and or VITROS MicroTips to be improperly seated on the barrel component of the Assembly which may result in an increased frequency of condition codes related to the actions of tip pickup aspirate dispense mix tip seal and discard Due to the small VITROS MicroTip and VersaTip dimension variability within manufacturing tolerances and the way these Tips interact with the Assembly Assemblies manufactured within the affected date range may not perform as intended based on the Tips currently in use x D To determine if a laboratory is using an Assembly that was manufactured within the affected date range customers are instructed to check the date code stamped into the barrel of the Assembly Refer to the customer letter for an example image of how to locate the date code on the barrel component of the Assembly If a laboratory has an Assembly with a date code between the Assembly is considered affected and must be replaced QuidelOrtho will provide a credit OR no cost replacement for each affected Assembly This includes all affected Assemblies in the laboratory including spare Assemblies not yet installed on a VITROS System specifically for biased results related to this issue nor have we experienced biased results during our investigation | Determine if any of the impacted VITROS Systems in their laboratory are using an affected Proboscis/Piston Assembly by checking the date code stamped into the barrel component. Once a replacement Proboscis/Piston Assembly is received, discard all affected Assemblies within the date range of 040722 - 121922. • Determine if their inventory of VITROS spare parts contains any affected Proboscis/Piston Assemblies by checking the date code stamped into the barrel component. Once a replacement Proboscis/Piston Assembly is received, discard all affected Assemblies within the date range of 040722 - 121922. • Ensure their laboratory is performing the Leak/Hysteresis test during daily maintenance and performing Quality Control to ensure a properly functioning Proboscis/Piston Assembly until a replacement Assembly is available. • Customers must complete the enclosed Confirmation of Receipt form. Upon receiving their completed Confirmation of Receipt form, QuidelOrtho will credit their account or send a replacement Proboscis/Piston Assembly (Part. No. J55375), if applicable. | 2023-09-27 09:34
Ortho Clinical Diagnostics Inc. | VITROS 5600 Integrated System | 0 | - | Ortho Clinical Diagnostics QuidelOrtho received complaints related to this issue as of Jul we have received no complaints QuidelOrtho confirmed the issue and determined that certain Proboscis Piston Assembly or Assemblies manufactured between Apr Date code and Dec Date code may not perform as intended due to an out of specification barrel component This may cause VITROS VersaTips and or VITROS MicroTips to be improperly seated on the barrel component of the Assembly which may result in an increased frequency of condition codes related to the actions of tip pickup aspirate dispense mix tip seal and discard Due to the small VITROS MicroTip and VersaTip dimension variability within manufacturing tolerances and the way these Tips interact with the Assembly Assemblies manufactured within the affected date range may not perform as intended based on the Tips currently in use x D To determine if a laboratory is using an Assembly that was manufactured within the affected date range customers are instructed to check the date code stamped into the barrel of the Assembly Refer to the customer letter for an example image of how to locate the date code on the barrel component of the Assembly If a laboratory has an Assembly with a date code between the Assembly is considered affected and must be replaced QuidelOrtho will provide a credit OR no cost replacement for each affected Assembly This includes all affected Assemblies in the laboratory including spare Assemblies not yet installed on a VITROS System specifically for biased results related to this issue nor have we experienced biased results during our investigation | Determine if any of the impacted VITROS Systems in their laboratory are using an affected Proboscis/Piston Assembly by checking the date code stamped into the barrel component. Once a replacement Proboscis/Piston Assembly is received, discard all affected Assemblies within the date range of 040722 - 121922. • Determine if their inventory of VITROS spare parts contains any affected Proboscis/Piston Assemblies by checking the date code stamped into the barrel component. Once a replacement Proboscis/Piston Assembly is received, discard all affected Assemblies within the date range of 040722 - 121922. • Ensure their laboratory is performing the Leak/Hysteresis test during daily maintenance and performing Quality Control to ensure a properly functioning Proboscis/Piston Assembly until a replacement Assembly is available. • Customers must complete the enclosed Confirmation of Receipt form. Upon receiving their completed Confirmation of Receipt form, QuidelOrtho will credit their account or send a replacement Proboscis/Piston Assembly (Part. No. J55375), if applicable. | 2023-09-27 09:34
Ortho Clinical Diagnostics Inc. | VITROS XT 7600 Integrated System | 0 | - | Ortho Clinical Diagnostics QuidelOrtho received complaints related to this issue as of Jul we have received no complaints QuidelOrtho confirmed the issue and determined that certain Proboscis Piston Assembly or Assemblies manufactured between Apr Date code and Dec Date code may not perform as intended due to an out of specification barrel component This may cause VITROS VersaTips and or VITROS MicroTips to be improperly seated on the barrel component of the Assembly which may result in an increased frequency of condition codes related to the actions of tip pickup aspirate dispense mix tip seal and discard Due to the small VITROS MicroTip and VersaTip dimension variability within manufacturing tolerances and the way these Tips interact with the Assembly Assemblies manufactured within the affected date range may not perform as intended based on the Tips currently in use x D To determine if a laboratory is using an Assembly that was manufactured within the affected date range customers are instructed to check the date code stamped into the barrel of the Assembly Refer to the customer letter for an example image of how to locate the date code on the barrel component of the Assembly If a laboratory has an Assembly with a date code between the Assembly is considered affected and must be replaced QuidelOrtho will provide a credit OR no cost replacement for each affected Assembly This includes all affected Assemblies in the laboratory including spare Assemblies not yet installed on a VITROS System specifically for biased results related to this issue nor have we experienced biased results during our investigation | Determine if any of the impacted VITROS Systems in their laboratory are using an affected Proboscis/Piston Assembly by checking the date code stamped into the barrel component. Once a replacement Proboscis/Piston Assembly is received, discard all affected Assemblies within the date range of 040722 - 121922. • Determine if their inventory of VITROS spare parts contains any affected Proboscis/Piston Assemblies by checking the date code stamped into the barrel component. Once a replacement Proboscis/Piston Assembly is received, discard all affected Assemblies within the date range of 040722 - 121922. • Ensure their laboratory is performing the Leak/Hysteresis test during daily maintenance and performing Quality Control to ensure a properly functioning Proboscis/Piston Assembly until a replacement Assembly is available. • Customers must complete the enclosed Confirmation of Receipt form. Upon receiving their completed Confirmation of Receipt form, QuidelOrtho will credit their account or send a replacement Proboscis/Piston Assembly (Part. No. J55375), if applicable. | 2023-09-27 09:34
Draeger Medical, Inc. | Filter/HME TwinStar | 789816 | - | Feedback from the market has shown that condensate can form more frequently in our TwinStar Plus filter HME which can lead to increased resistance under certain circumstances | Sending the safety instructions. | 2023-09-27 14:09
Hamilton,Switzerland | HAMILTON MEDICAL VENTILATORS (MR1,T1) | 14 | - | the device may interrupt ventilation due to a faulty capacitor Hamilton Medical AG received nine x D complaints involving ventilation interruption while neither serious injuries nor deaths related to this issue were x D reported | Replace Control Board | 2023-09-27 14:13
Hamilton,Switzerland | HAMILTON T1 ,MR1 | 14 | - | t the device may interrupt ventilation due to a faulty capacitor Hamilton Medical AG received nine x D complaints involving ventilation interruption while neither serious injuries nor deaths related to this issue were x D reported | Replace Control Board | 2023-09-27 14:33
Roche | VENTANA anti-ALK (D5F3) Antibody | 1 | - | This is an active investigation and we are working to further identify and address the root cause In an effort to balance the actions taken for this product quality issue with the need to maintain market availability for customers and patients we are providing you with the following guidance | Immediately discontinue use of and discard any remaining inventory of the impacted VENTANA anti-ALK (D5F3) lots listed in Table A below. Further, a short date has been assigned for each at 14 months post-manufacture. Roche recommends that customers review patient samples tested after the short date listed in Table A below and retest any clinically diagnosed negative ALK (D5F3) results, especially where only appendix is being utilized as a positive tissue control. | 2023-09-27 22:03
List Attached | Tempus Pro Monitor | - | - | An issue has been identified with the Tempus Pro Monitor where an error may occur during video x D laryngoscope use with the Tempus Pro Video Laryngoscope or immediately after the Tempus Pro Video x D Laryngoscope has been unplugged from the Tempus Pro Monitor If this error occurs the user is x D presented with an unexpected full screen message informing the user an error has occurred requiring x D shutdown and restart of the Tempus Pro Monitor This full screen message prevents the user from x D viewing any graphical representation of patient vital signs however text and numerical values are still x D visible on the device s display screen | Philips has developed a software update to resolve this issue that is expected to be available in Q3 2023. When it is available, Philips will provide the updated software and installation instructions (via a downloadable link) for customer installation. Upon customer request, Philips can also provide a USB Flash Drive with the updated version of software. If you need any further information or support concerning this issue, please contact your local Philips representative. met.quality@philips.com This notice has been reported to the appropriate Regulatory Agencies. Be sure to report any occurrence of this issue to Philips, your Philips representative, or to your local Regulatory authority. | 2023-10-01 16:24
Ortho Clinical Diagnostics Inc. | ALKP Slides | - | - | QuidelOrtho received complaints related to under recovery when testing an elevated patient sample with VITROS ALKP Slides The VITROS Systems have algorithms designed to perform a check for substrate depletion when processing VITROS MicroSlides This check is intended to prevent samples with high concentrations of ALKP i e above the measuring range from being reported as within the ALKP measuring range when tested undiluted Samples containing high levels of ALKP or an interfering substance could deplete the available substrate in the reagent layer of a VITROS ALKP Slide and not post the DP code which may result in the System reporting an incorrect value lower than the actual value To mitigate this issue QuidelOrtho recommends performing a x dilution on sample results reported between U L when tested undiluted If the x diluted result is below U L after correcting for the dilution factor then the original undiluted result should be reported If the x diluted result is above U L after correcting for the dilution factor then the diluted result should be reported Undiluted sample results above U L should be diluted per the Instructions for Use IFU and your laboratories protocol Diluting results within the measuring range is typically not recommended by Ortho however due to this issue a dilution within the measuring range is recommended to confirm the result does not exceed the upper end of the measuring range of U L on highly elevated samples | When reviewing sample results generated using VITROS ALKP Slides or VITROS XT TBIL-ALKP Slides, any result between 600-1500 U/L should be diluted using a 10x dilution and retested. | 2023-10-02 09:27
Ortho Clinical Diagnostics Inc. | VITROS Chemistry Products TBIL- ALKP Slides | - | - | QuidelOrtho received complaints related to under recovery when testing an elevated patient sample with VITROS ALKP Slides The VITROS Systems have algorithms designed to perform a check for substrate depletion when processing VITROS MicroSlides This check is intended to prevent samples with high concentrations of ALKP i e above the measuring range from being reported as within the ALKP measuring range when tested undiluted Samples containing high levels of ALKP or an interfering substance could deplete the available substrate in the reagent layer of a VITROS ALKP Slide and not post the DP code which may result in the System reporting an incorrect value lower than the actual value To mitigate this issue QuidelOrtho recommends performing a x dilution on sample results reported between U L when tested undiluted If the x diluted result is below U L after correcting for the dilution factor then the original undiluted result should be reported If the x diluted result is above U L after correcting for the dilution factor then the diluted result should be reported Undiluted sample results above U L should be diluted per the Instructions for Use IFU and your laboratories protocol Diluting results within the measuring range is typically not recommended by Ortho however due to this issue a dilution within the measuring range is recommended to confirm the result does not exceed the upper end of the measuring range of U L on highly elevated samples | When reviewing sample results generated using VITROS ALKP Slides or VITROS XT TBIL-ALKP Slides, any result between 600-1500 U/L should be diluted using a 10x dilution and retested. | 2023-10-02 09:27
Olympus Medical Systems Corp. | Olympus Flexible Endoscopes | 2 | - | Following repair activities Olympus uses a water flush of endoscope channels as part of our final endoscope channel inspection and we subsequently dry the endoscope channels with compressed filtered air Olympus discovered that our channel air drying process was not validated and that a small percentage of scopes returned to customers after repair had a wet channel Olympus is in the process of implementing a validated drying process at our repair centers | As a precautionary measure, Olympus requires that you inspect endoscopes received from our repair facilities for the presence of water prior to reprocessing. If you observe any moisture or water in your returned endoscope, please contact your local Olympus representative to coordinate a return. Until we have implemented the validated process, you will receive a reminder with each returned repaired device. See attached FSN | 2023-10-02 17:12
DiaMed GmbH | ID-DiaPanel Plus 6 | - | - | With this letter we would like to inform you about a typing error affecting the Antigen table of the set of ID DiaPanel Plus x D The phenotype mentioned in the antigen table of cell of the above mentioned lots of ID Diapanel Plus is incorrect It is indicated CCDwD ee while the correct one should have been CCCwD ee x D The product itself and the remaining content of the antigen table are correct The product can be used without any restriction x D Please find enclosed with this field action letter the correct antigen table Also the correct antigen table can also be downloaded on the following website x D https ih area bio rad com system files B X V pdf | Please find enclosed with this field action letter the correct antigen table. Also, the correct antigen table can also be downloaded on the following website: https://ih-area.bio-rad.com/system/files/B05300.35X_V.02.pdf | 2023-10-03 14:56
PEROUSE MEDICAL | dolphin | 477 | - | PEROUSE MEDICAL became aware of complaints concerning the defective packaging of the DOLPHIN inflation device x D The blister may have a visible crack in the side | x Identify the Device x Quarantine the Device x Return the Device | 2023-10-03 15:46
Abbott | CardioMEMS Patient Electronics System | 86 | - | If the PES power connector plug becomes frayed and or has exposed wiring the following may occur x D PES does not power on This problem is routinely reported quickly and the PES will be replaced by x D Abbott Technical Support Therefore it is not likely that this will interrupt ongoing hemodynamicguided x D heart x D failure x D management x D long x D enough x D to x D result x D in x D unanticipated x D worsening x D or x D acute x D x D decompensation x D Power connector plug sparking could cause fire potentially resulting in burns The potential for x D medically reversible or transient adverse health consequence is remote x D Discomfort from mild shock or sparking resulting from exposed wiring The potential for medically x D reversible or transient adverse health consequences is remote x D Should a patient s power connector plug present these issues a replacement PES will be sent to minimise x D any interruption in daily readings | Abbott will update the IFU in the appropriate language by geography to add clarity on how to connect the power adapter cable to the power connector plug. Updated IFUs will be available to patients on the Abbott website under Manuals & Technical Resources: (https://www.cardiovascular.abbott/int/en/hcp/products/heartfailure/pulmonary-pressure-monitors/cardiomems/manuals-and-resources.html). Until these updates are finalised, Abbott is reinforcing the warning in the IFU and providing supplemental guidance. See Enclosure: Patient Letter | 2023-10-07 18:56
Carestream Health,Inc. | DRX-Revolution Mobile X-Ray System | 112 | - | There is a potential for unexpected failure of the electrical components within the Carestream Health Inc x D generator installed in DRX Revolution Mobile X ray System s | Inform all personnel that utilize the system of the potential issue. • Should an event of this nature occur: o Move the system away from patients/staff. o The system will be inoperable. o Call the Carestream Customer Care Center in the U.S. at 1-800-328-2910, your local Carestream Service support number or Carestream Health Authorized Representative to place a service call. • If an Adverse Health Event is encountered, notify FDA MedWatch o MedWatch: Use the MedWatch Online Reporting Form to report adverse events or call 800-332-1088 . o If outside the US, notify the local competent authority in your region. | 2023-10-08 09:26
Beckman Coulter, Inc. | DxI 9000 Access Immunoassay Analyzer | 1 | - | Beckman Coulter has determined that the DxI Access x D Immunoassay Analyzer may incorrectly identify samples x D The issue only occurs under the following conditions x D The DxI analyzer is connected to a host system LIS or x D middleware x D Normal sample processing occurs through the host system x D a A sample Sample A with test orders is submitted from the x D host system and loaded onto the DxI analyzer using a x D DxLAB sample rack x D b The analyzer completes the tests and transmits test results x D back to the host system for Sample A x D The host system sends the analyzer new information or test order x D updates for Sample A This can happen immediately or days later x D This action assigns Sample A to the previous rack and position it x D occupied during step b x D A new sample Sample B with tests orders is submitted from the x D host system and placed into the same rack and position to which x D Sample A was previously assigned and the rack is loaded onto the x D DxI analyzer x D The analyzer reads the tube barcode but disregards the sample ID x D of Sample B and reverts to information for Sample A x D a Any pending tests for Sample A will be run using Sample B x D instead resulting in an erroneous result for Sample A In x D addition the tests ordered for Sample B will not run and no x D results will be reported for Sample B Sample B will be x D offloaded with no notification to the user x D b If Sample A received a new information update such as x D patient demographic information from the host system but no x D test order updates the tests ordered for Sample B will not run x D and no results will be reported for Sample B Sample B will be x D offloaded with no notification to the user x D DocuSign Envelope ID DD D E C F DE AE E AD F C A C x D S Kraemer Blvd Brea CA www beckmancoulter com x D FA Page x D Confidential Company Proprietary x D Subsequent update messages from the host system for Sample A will x D initiate the sequence starting at Step The sequence will occur each time x D that the host system submits additional test requests for Sample A while any x D other sample not just Sample B occupies the same rack and position that x D was initially occupied by Sample A x D The issue only affects samples loaded using DxLAB sample racks x D loaded directly onto the sample handler of a DxI analyzer x D connected to a host system x D This issue does not affect samples that are sampled from an x D automation line |  Perform a retrospective review of all add on tests or rerun tests that were generated on a DxI 9000 analyzer that is connected to a host system.  Retesting is recommended on Sample A.  Perform one of the following: Option 1: Disconnect the DxI 9000 Access Immunoassay Analyzer from the host system. You may continue ordering tests through the analyzer console. Option 2: Follow the instructions provided in Appendix A. | 2023-10-09 16:21
Lowenstein Medical Innovation GmbH + Co. | Affected are all intensive care ventilators elisa 300, elisa 500, elisa 600, elisa 800 and elisa 800VIT with software version < 1.11.3, 2.04.7, 2.09.14, 2.10.6 or 2.13.2. | - | - | During market monitoring we have become aware of cases in which delayed recognition of the patient may occur after x D disconnection when automatic patient detection APD and tube compensation are activated which may take up to min x D The following events must coincide for this to occur x D automatic patient detection APD is activate x D tube compensation x D patient is disconnected e g for suctioning x D due to the delayed recognition the device still remains in recognition mode thus delivering only l min byflow x D spontaneously breathing patient inhales forcefully and continuously x D As a result of this the following scenario may occur x D due to the fact that only l min flow is supplied a negative pressure is created and after a certain time the alarm x D Airway pressure negative is triggered x D the alarm causes the safety valve to open to allow the patient to breathe x D the opened safety valve prevents the PEEP or an expiratory flow from being reached to exit the detection mode x D after a delay time of s the safety valve closes patient detection is possible again x D in the worst case the patient s inhalation effort occurs before the device has detected the patient which again leads to x D the opening of the safety valve x D The reason for this is the missing update of the calculation of the pressure drop in the tube x D x D Problem x D The software bug described above also has an impact during open suctioning where automatic patient detection APD is always x D activated regardless of the device configuration If the ventilation tube is disconnected during preoxygenation without waiting for x D it to end the device automatically switches to suction mode If tube compensation is switched on the problem described above x D occurs x D The reason for this is the missing update of the calculation of the pressure drop in the tube x D FSCA | Problem 1: • Deactivation of automatic patient detection (APD) Problem 2: • Open suction must be performed in such a way that the end of preoxygenation (2 min) is waited for or manually switched to the actual suction, but the ventilation tube is not disconnected during preoxygenation. | 2023-10-09 16:56
BD Switzerland Sàrl, Eysins, Switzerland | BD Alaris™ Pump Sets, Gravity, and Extension infusion sets | - | - | BD has identified internally that the infusion sets listed in Appendix contain di ethylhexyl phthalate DEHP and have not been labelled accordingly The infusion sets contain DEHP by design and the content remains unchanged since product launch | BD Actions: BD will be updating the labelling for these devices to state “Contains or presence of phthalate: DEHP.” Clinical User Actions: • BD recommends using DEHP-free devices when specified by drug manufacturers. • For devices in-situ with patients requiring a DEHP-free treatment, BD instructs to cease use and an alternate treatment or device be sourced. • For devices that have already been used, there are no further actions that need to be taken. Customer Actions: • Review the information in Appendix 1 and determine if the devices in your possession are impacted. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 7th November 2023. • Circulate this notice to all those who need to be aware within your organisation or to any organisation where the potentially affected product has been transferred. • If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: • Review the information in Appendix 1 and determine if the devices in your possession are impacted. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 7th November 2023. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. | 2023-10-11 00:21
Randox Laboratories | Glucose-6-Phosphate Dehydrogenase (G-6-PDH) | - | - | Randox Laboratories is conducting a Device Recall for G PDFI reagent catalogue number x D PD batch The R NADP and R Substrate are showing low rate of absorption x D change resulting in cases of low or no reaction taking place leading to low Q C recovery to x D target and patient results reported below the normal range in erythrocytes mU x D U g Hb C | -Discontinue use of and discard any of the above immediately. Provide Randox with photographic evidence of the destruction of the kits. • Review your reagent inventory of these products and assess your laboratories needs for reimbursement for discarded inventory. • Review results generated with the affected batches in line with the clinical profile of the patient. • Discuss the contents of this notice with your Medical Director. • Complete and return the response form 12187-QA to technical.services(5>randox.com within five working days. | 2023-10-11 13:29
Fresenius Medical Care AG & Co. KGaA | stay•safe® Disinfection Cap | 18200 | - | An increased exposure to povidone iodine could lead to iodine induced hypothyroidism and associated clinical sequelae x D This is especially a concern for infants and children with smaller peritoneal fill volumes in which iodine overload could result in secondary hypothyroidism | Fresenius Medical Care recommends close monitoring of thyroid function in patients with small peritoneal fill volumes, typically infants and children, when using iodine containing products. Fresenius Medical Care is in the process of updating the labeling to include these medical warnings and contraindications. | 2023-10-11 15:20
Fresenius Medical Care AG & Co. KGaA | stay•safe® Disinfection Cap | - | - | please check case ISA | please check case ISA-28-09-23-237 | 2023-10-11 15:25
Siemens Healthcare GmbH | Cios Alpha | 18 | - | What is the issue and when does it occur x D Potentially affected system Cios Spin VA x D x D With respect to DIN x D So far in D mode information about image orientation regarding laterality i e L left R right cannot be x D defined and saved with the clinical image by the user x D x D Potentially affected systems Cios Alpha VA Cios Spin VA Cios Flow VA x D x D With respect to IEC x D So far in radiography modes SUB DR and DCM the autostore function can be deactivated by the user however x D in D mode the autostore function is activated by default and cannot be deactivated by the user | The corrective action will be implemented with the upcoming software version VA30K as measure AX036/23/S. | 2023-10-12 10:33
Siemens Healthcare GmbH | Cios Flow | 18 | - | What is the issue and when does it occur x D Potentially affected system Cios Spin VA x D x D With respect to DIN x D So far in D mode information about image orientation regarding laterality i e L left R right cannot be x D defined and saved with the clinical image by the user x D x D Potentially affected systems Cios Alpha VA Cios Spin VA Cios Flow VA x D x D With respect to IEC x D So far in radiography modes SUB DR and DCM the autostore function can be deactivated by the user however x D in D mode the autostore function is activated by default and cannot be deactivated by the user | The corrective action will be implemented with the upcoming software version VA30K as measure AX036/23/S. | 2023-10-12 10:33
Siemens Healthcare GmbH | Cios Spin | 18 | - | What is the issue and when does it occur x D Potentially affected system Cios Spin VA x D x D With respect to DIN x D So far in D mode information about image orientation regarding laterality i e L left R right cannot be x D defined and saved with the clinical image by the user x D x D Potentially affected systems Cios Alpha VA Cios Spin VA Cios Flow VA x D x D With respect to IEC x D So far in radiography modes SUB DR and DCM the autostore function can be deactivated by the user however x D in D mode the autostore function is activated by default and cannot be deactivated by the user | The corrective action will be implemented with the upcoming software version VA30K as measure AX036/23/S. | 2023-10-12 10:33
Abbott GmbH. & Co. KG | Alinity s System | 18 | - | Please find attached letter | Please find attached letter | 2023-10-15 10:58
Edwards LLC | Acumen | 20 | - | The Acumen IQ Sensors Flotrac Sensors and TruWave Disposable Pressure Transducers DPT are being used for various procedures in the monitoring of physiological parameters The Acumen IQ sensor provides pressure values used in the measurement of systolic diastolic and mean arterial pressure values which provides blood pressure values cardiac output stroke volume stroke volume variation and systemic vascular resistance values to allow the correct differential diagnosis of patients x D x D Through post market surveillance Edwards Lifesciences has recently received customer complaints noting flushing difficulties during setup of the devices listed above The investigation revealed a manufacturing issue causing a potential occlusion within certain devices Some devices were fully occluded and could not be flushed while other devices were partially occluded causing slower flushing Our data indicates the issue potentially occurs in of affected lots x D x D Edwards has confirmed a manufacturing defect leading to possible flushing difficulties No patient harm was reported in any of these complaints The manufacturing issue was resolved and newly manufactured devices are not impacted x D x D The risk to health is limited as these products are used by highly trained clinicians experienced in identifying and mitigating any hazards that arise during the case There is no anticipated patient harm due to the issue as it is detected prior to use when flushing the device per the Instructions For Use IFU | Edwards Lifesciences is voluntarily notifying customers of an issue regarding certain lots of TruWave Disposable Pressure Transducers (DPTs), FloTrac Sensors, and Acumen IQ Sensors. No product return is necessary. A letter will be sent to all affected customers notifying them about the issue. Customers are instructed to acknowledge this notification, and return the acknowledgment form to Edwards Lifesciences. | 2023-10-15 15:06
Instrumentation Laboratory Co. | ACL TOP Family / ACL TOP Family 50 Series | 82 | - | Mandatory Parameters P for the ACL TOP Family and P and P for the x D ACL TOP Family Series contain the final mitigations to address the following field actions x D Reduction of on board stability claim from days to days for HemosIL Liquid Anti Xa x D Reduction of on board stability claim from days to hours for HemosIL Homocysteine x D Controls x D Updated test definitions to mitigate carryover for HemosIL Liquid Anti Xa x D NOTE This upgrade is mandatory for all ACL TOP Family ACL TOP Family Series regardless x D of whether the above reagents are currently in use x D The above changes to the on board stability claims apply to all currently released lots and future x D lots of HemosIL Liquid Anti Xa and HemosIL Homocysteine Controls | Please take the following actions:  Share this information with your laboratory staff and update your internal procedures, as needed.  Forward this notification to all affected locations in your facility.  Post this notification on all affected systems in your facility. NOTE: All customers with ACL TOP Family 50 Series systems with software (SW) versions 6.1.0, 6.2.0, 6.3.0 (Windows 7 OS) must be upgraded to SW 6.4.1 prior to installation of these parameters. For ACL TOP Family 50 Series systems with Windows 10 OS, you must be at SW version 6.5.2 prior to installation of these parameters. All ACL TOP Family customers need to be at a minimum of SW version 5.2.0 to install the Mandatory Test Parameters. | 2023-10-16 11:07
Philips | Defirbrillator/ Monitor HeartStart Intrepid | 65 | - | The current HeartStart Intrepid Service Manual does not detail IEC Class I electrical safety test verification methods The lack of this testing does not introduce a failure but it may prevent a failure from being detected This could occur when an IEC Class I device is tested as a Class II device after a servicing event x D x D The issue was identified as the result of a service provider s inquiry There have been no reports of patient harm x D x D The HeartStart Intrepid is a monitor defibrillator used in an emergency medical services or hospital environment by qualified medical personnel trained in its operation to provide pacing defibrillation and synchronized cardioversion therapies It is intended to measure heart rate and rhythm blood oxygen saturation exhaled CO systolic diastolic and mean blood pressure and temperature and deliver external cardiac pacing | Philips has released an addendum to the HeartStart Intrepid Service Manual for devices categorized as IEC Class I equipment describing how to conduct electrical safety tests in accordance with IEC 60601 and IEC 62353. It is included with this notification and may be used immediately. Philips will contactthe customer to arrange complimentary retesting to Class I as necessary. | 2023-10-16 11:07
Maquet Cardiopulmonary GmbH | CARDIOHELP-i | 2 | - | It was found that the fixation weld employed by the supplier of the nut bolt to the slide rail plate was insufficient x D The deficient weld allows the nut bolt to break off from the slide rail plate with slight manual force | Replacement of the slide rail holders | 2023-10-16 18:43
Carestream Health,Inc. | DRX-Revolution Mobile X-Ray System | - | - | a potential for unexpected failure of the electrical components within the Carestream Health Inc generator | Inform all personnel that utilize the system of the potential issue. • Should an event of this nature occur: o Move the system away from patients/staff. o The system will be inoperable. o Call the Carestream Customer Care Center in the U.S. at 1-800-328-2910, your local Carestream Service support number or Carestream Health Authorized Representative to place a service call. • If an Adverse Health Event is encountered, notify FDA MedWatch o MedWatch: Use the MedWatch Online Reporting Form to report adverse events or call 800-332-1088. o If outside the US, notify the local competent authority in your region. • Review and complete the “Consignee Notification Acknowledgement” form as instructed. | 2023-10-18 09:50
Agilent Technologies Singapore Pte Ltd. | GA08561-2, Monoclonal Mouse Anti-Human MSH2, Clone FE11 (Dako Omnis), -GA08661-2, Monoclonal Rabbit Anti-Human MSH6, Clone EP49 (Dako Omnis), and -GA08761-2, Monoclonal Rabbit Anti-Human PMS2, Clone EP51 (Dako Omnis). | 8 | - | In the period between February and January IFUs incorrectly specified tumors of the intestinal tract in the Intended Use Since the error was detected in January the indication was corrected to colorectal cancer and the eIFUs were published on Agilent com The product quality is not impacted and the product is safe for use | notify all customer about the changes. | 2023-10-22 11:00
Abbott Laboratories | Alinity s System | 18 | - | As per attached FSN letter | As per attached FSN letter | 2023-10-23 11:24
Intuitive Surgical,Inc. | da Vinci Surgical System | 2 | - | The da Vinci X and Xi system instrument arms USMs have Instrument Carriages that may separate from the Insertion Axis Linear Rail during shipping handling or repositioning of the arms when the system is powered down | schedule a site visit to inspect the insertion of linear rail and replace the USM if necessary. | 2023-10-24 16:35
Philips Healthcare (Suzhou) Co., Ltd. | Multiva 1.5T | 2 | - | Philips has identified an issue where a specific component failure in the Gradient Coil of the affected Multiva T MR Systems listed in this letter may act as a heat source with a potential to produce smoke and or fire x D x D If the component failure occurs the user may observe the following x D Repeated scan abort error messages due to gradient amplifier fault detection described in Section x D Unusual noise from the system in the examination or technical room x D Smoke and or fire from within the system x D x D Philips has received one complaint of smoke fire caused by this issue which is associated with the Gradient Coil type used in the MR Systems identified in this letter There was no report of injury or serious harm to hospital staff or patients however damage to property has occurred | Philips is providing this customer letter containing guidance until the investigation into this issue is complete. Philips is committed to addressing the issue and continues to investigate its root cause. We anticipate providing an update to you on the development of our plan to address the issue by in early 2024. | 2023-10-25 10:30
Philips Medical Systems | Achieva 1.5T | 16 | - | Philips has identified an issue where a specific component failure in the Gradient Coil of the affected T MR Systems listed in this letter may act as a heat source with a potential to produce smoke and or fire x D x D If the component failure occurs the user may observe the following x D Repeated scan abort error messages due to gradient amplifier fault detection described in Section x D Unusual noise from the system in the examination or technical room x D Smoke and or fire from within the system x D x D Philips has received one complaint of smoke fire caused by this issue which is associated with the Gradient Coil type used in the MR Systems identified in this letter There was no report of injury or serious harm to hospital staff or patients however damage to property has occurred x D | Philips is providing this customer letter containing guidance until the investigation into this issue is complete. Philips is committed to addressing the issue and continues to investigate its root cause. We anticipate providing an update to you on the development of our plan to address the issue in early 2024. | 2023-10-25 10:38
Philips Medical Systems | Ingenia 1.5T CX | 16 | - | Philips has identified an issue where a specific component failure in the Gradient Coil of the affected T MR Systems listed in this letter may act as a heat source with a potential to produce smoke and or fire x D x D If the component failure occurs the user may observe the following x D Repeated scan abort error messages due to gradient amplifier fault detection described in Section x D Unusual noise from the system in the examination or technical room x D Smoke and or fire from within the system x D x D Philips has received one complaint of smoke fire caused by this issue which is associated with the Gradient Coil type used in the MR Systems identified in this letter There was no report of injury or serious harm to hospital staff or patients however damage to property has occurred x D | Philips is providing this customer letter containing guidance until the investigation into this issue is complete. Philips is committed to addressing the issue and continues to investigate its root cause. We anticipate providing an update to you on the development of our plan to address the issue in early 2024. | 2023-10-25 10:38
Philips Medical Systems | SmartPath to dStream for 1.5T | 16 | - | Philips has identified an issue where a specific component failure in the Gradient Coil of the affected T MR Systems listed in this letter may act as a heat source with a potential to produce smoke and or fire x D x D If the component failure occurs the user may observe the following x D Repeated scan abort error messages due to gradient amplifier fault detection described in Section x D Unusual noise from the system in the examination or technical room x D Smoke and or fire from within the system x D x D Philips has received one complaint of smoke fire caused by this issue which is associated with the Gradient Coil type used in the MR Systems identified in this letter There was no report of injury or serious harm to hospital staff or patients however damage to property has occurred x D | Philips is providing this customer letter containing guidance until the investigation into this issue is complete. Philips is committed to addressing the issue and continues to investigate its root cause. We anticipate providing an update to you on the development of our plan to address the issue in early 2024. | 2023-10-25 10:38
ELEKTA LIMITED | Elekta Synergy Platform | 7 | - | There is a potential for an uncontrolled extension of iViewGT XVI detector arms if the microswitch within the middle arm is incorrectly setup x D | Users need to be aware of attachment content which contains recommendations to the customer of workarounds while waiting for the permanent correction. It is important that the user does not open or close the detector arm while it is positioned above the patient, as specified in the Instructions for Use. Elekta Service Representative will be fitting a bracket to the middle arm assembly to ensure that the micro-switch will activate only when the detector-locked solenoid is fully engaged in the interface bracket. Target release date for the correction is March 2024 | 2023-10-25 16:20
ELEKTA LIMITED | Elekta Synergy | 7 | - | There is a potential for an uncontrolled extension of iViewGT XVI detector arms if the microswitch within the middle arm is incorrectly setup x D | Users need to be aware of attachment content which contains recommendations to the customer of workarounds while waiting for the permanent correction. It is important that the user does not open or close the detector arm while it is positioned above the patient, as specified in the Instructions for Use. Elekta Service Representative will be fitting a bracket to the middle arm assembly to ensure that the micro-switch will activate only when the detector-locked solenoid is fully engaged in the interface bracket. Target release date for the correction is March 2024 | 2023-10-25 16:20
ELEKTA LIMITED | Versa HD | 7 | - | There is a potential for an uncontrolled extension of iViewGT XVI detector arms if the microswitch within the middle arm is incorrectly setup x D | Users need to be aware of attachment content which contains recommendations to the customer of workarounds while waiting for the permanent correction. It is important that the user does not open or close the detector arm while it is positioned above the patient, as specified in the Instructions for Use. Elekta Service Representative will be fitting a bracket to the middle arm assembly to ensure that the micro-switch will activate only when the detector-locked solenoid is fully engaged in the interface bracket. Target release date for the correction is March 2024 | 2023-10-25 16:20
Becton, Dickinson and Company, BD Biosciences | CD11c (S-HCL-3) APC | 8 | - | BD has confirmed through eight customer complaint investigations that the above affected lot of BD CD b APC CE has low or dim fluorescence signal reported x D x D | BD Actions: BD has implemented the appropriate corrective actions to prevent recurrence. Clinical/End User Actions: • Clinical laboratory staff should cease use of the affected lot of BD™ CD11b APC, CE. • Identify if the affected lot is available at the laboratory, ensure the new reagent lot validation internal procedure is followed and check the results. • Quarantine, and replace all units of the affected lot with an unaffected lot accordingly. • If the affected lot was used in testing of patient samples, review the results of the lot validation testing to identify if the intensity of staining was appropriate, and review the patient results for accuracy and consider other diagnostic tests, if necessary. Customer Actions: • Inspect your inventory, identify, and quarantine any units of the affected lot of BD™ CD11b APC, CE. • Destroy all units of the affected lot. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 30th November 2023. • Circulate this notice to all those who need to be aware within your organisation or to any organisation where the affected product has been transferred. • If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: • Cease distribution of the affected lot of BD™CD11b APC, CE. • Identify, quarantine, and destroy all units of the affected lot of BD™CD11b APC, CE. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 30th November 2023. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. | 2023-10-26 01:54
Olympus Medical Systems Corp. | EVIS X1 VIDEO SYSTEM CENTER OLYMPUS CV-1500 | 45 | - | Olympus has received one complaint that the water delivery tube came in contact with the touch panel and the Air flow button was unintentionally turned on This resulted in an adverse event and the patient had a pneumothorax due to the air insufflation during the Peroral Endoscopic Mytomy POEM procedure | 1. Ensure all personnel are completely knowledgeable and thoroughly aware of the recommendation to utilize the Lock touch panel function during procedures and the warning to keep objects away from the touch screen in order to prevent unintentional actuation of the touch panel. 2. Please indicate on the enclosed reply form that you have received this notification and taken the required actions by filling out and returning the completed enclosed Reply Form back to your local Olympus representative latest by 08.11.2023. | 2023-10-26 11:17
DiaSorin Italia S.p.A | LIAISON® Chlamydia trachomatis IgG | 3 | - | Our records confirm that you may have received the lot of the LIAISON Chlamydia trachomatis IgG product x D x D We have identified a potential issue on the calibration setting which might lead to an unexpected shift of the negative samples doses x D If the issue occurs it may result in x D Negative control out of CoA range thus leading to an Invalid Run x D False Positive results x D x D As reported in the Instructions For Use IFU Chlamydia IgG testing should be done in combination with Chlamydia IgA and the final interpretation of the single specimen should come from the combined results of both tests Furthermore diagnosis of infectious diseases should not be established on the basis of a single test result but should be determined in conjunction with clinical findings and other diagnostic procedures as well as in association with medical judgement x D Therefore a false positive result is not likely to cause adverse health consequences x D | • If you have Invalid Runs due to the Negative Control out of range, please contact your DiaSorin Representative for guidance on an alternative method. • If you have valid analytic sessions, please confirm all the Positive Results with an alternative method. • According to the risk assessment above, it is not deemed necessary to re-evaluate your whole previous analytical sessions. If there are some positive samples not concordant with clinical findings and other diagnostic procedures or with medical judgement, please evaluate to confirm them with an alternative method. Please, contact your DiaSorin Representative for further information. | 2023-10-26 15:23
Change Healthcare Canada | - | - | - | Circumstances under which the issue occurs x D Currently Change Healthcare systems leveraging the PACS event notification service for x D messaging include Change Healthcare Workflow Intelligence CHWI and Change Healthcare x D Image Repository CHIR Other applications such as secondary PACS RIS systems other overlay x D products and third party VNA s may be impacted if they are configured to use the PACS event x D notification service These systems will only be impacted if study information from the CHRS x D database includes certain special characters and is used for creating outgoing x D messages sent using the CHRS PACS event notification service x D This failure can cause data quality issues in any system that receives PACS Event Notifications x D from CHRS This failure has the potential to result in data values in interfaced applications that x D are not in sync with actions performed on CHRS data For example a patient is merged in CHRS x D and the non surviving patient record deleted patient last name value has a last name of x D O Brian The special character in the deleted patient name causes the merge message to fail to x D send to the interfaced system Therefore the intended non surviving patient record would be x D deleted in the PACS database but would persist in the interfaced database This creates a x D situation where the interfaced database contains a patient record that is missing studies | Immediate actions to take to reduce patient impact: Users are advised to check the CHRS Exception Tracker on a regular basis, as failed updates will be listed with details on the failure. Product updates that will address the issue: Change Healthcare has developed a software update that will address this issue for all impacted versions. Recommended actions to take on potentially affected data: Impacted data can be identified by using the CHRS Exception Tracker. Customers are advised to review records of recent updates from CHRS to external systems and manually update any changes to study records that may have been missed. This notice must be distributed to all personnel within your organization who need to be aware of this Urgent Field Safety Notice. Customers should alert other affiliated parties that may be affected by this Urgent Field Safety Notice. To ensure effectiveness of any required corrective actions, please maintain awareness of this Urgent Field Safety Notice until the issue has been fully resolved. To acknowledge your understanding of this urgent field safety notice please complete and return the acknowledgement letter. To request the installation of the product update on your system, please contact our Customer Support department. You can find the toll-free number for your region at https://iwcrm.changehealthcare.com. If you have any questions or need further clarification about this notice, please do not hesitate to reach out to our Customer Support department using the toll-free number provided in the link above. Once again, until the product update is installed at your site, Change Healthcare recommends that the immediate actions to avoid the issue by reviewing updates from CHRS to downstream systems be taken. | 2023-10-29 13:27
FEATHER SAFETY RAZOR CO., LTD | MICRO FEATHER DISPOSABLE OPHTHALMIC SCALPEL WITH PLASTIC HANDLE | - | - | The manufacturer found out that the plastic handle of the product changes over time resulting in weakening of the fixing strength In several circumstances the blade fell out of the handle | Actions to be taken by the user Immediately stop using and contact your distributor to return the product. FSCA taken by the manufacturer The manufacturer will recall the product, and remove the suspected Lot. from the Said Arabian Market. | 2023-10-29 13:35
Maquet SAS | Maquet SAS’ OR Light Systems | - | - | Maquet SAS Getinge is initiating a voluntary Urgent Medical Device Correction for Maquet SASs OR lights due to the potential for a light system to fall in the operating room x D x D Maquet SAS Getinge has received a customer complaint regarding a fall of a light system in an operating room in Hong Kong Investigation into this event demonstrated that the recommendation of maintenance and service was not performed as instructed in the current user and maintenance manuals No serious injury has been reported as a result of this occurrence x D | Please distribute this Field Safety notice and forward the link to access manuals to anyone who may need to use the manuals, and especially to your service or service provider who performs your maintenance. Refer to the enclosed customer letter for details. | 2023-10-29 15:32
Olympus Medical Systems Corp. | BRONCHOVIDEOSCOPE OLYMPUS | 158 | - | Olympus has received four adverse event complaints of endobronchial combustion during therapeutic procedures with the Olympus bronchoscope model BF XT of which one involved High frequency therapy equipment The other three adverse events involved unknown energy therapy equipment | 1.Ensure all personnel are completely knowledgeable and thoroughly aware of the Warnings in affected bronchoscope’s Operation Manual for use with high-frequency devices and that Olympus high-frequency compatible bronchoscopes are compatible only with Combination equipment listed in the operation manual. 2.Please indicate on the enclosed reply form that you have received this notification and taken the required actions by filling out and returning the completed enclosed Reply Form back to your local Olympus representative latest by 10.11.2023. | 2023-10-29 17:57
UTAK Laboratories, Inc. | Validity Control 2 ADULTERANT TOXICOLOGY CONTROL | 242 | - | While doing containment for a related issue it was discovered by UTAK that particulate matter x D may be present in the product | Our team has opened a corrective action and is actively looking into the root cause of this matter. | 2023-10-30 11:31
UTAK Laboratories, Inc. | Validity Control 3 ADULTERANT TOXICOLOGY CONTROL | 242 | - | While doing containment for a related issue it was discovered by UTAK that particulate matter x D may be present in the product | Our team has opened a corrective action and is actively looking into the root cause of this matter. | 2023-10-30 11:31
Nova Biomedical Corporation | Stat Profile Prime Plus | 3 | - | Reporting enormous Creatinine patient results associated with the initial start up of a new BUN Creatinine Sensor Cartridge | 1. Software upgrades have been performed on ALL Prime Plus analyzers installed at customer sites as well as all demonstration systems. 2. Affected Customers were notified and given International Field Correction Notice Response Form 1 and 2 from Manufacturer. 3. Software for affected units upgraded as per manufacturer's instructions. 4. Affected 2 Customers signed that the corrective action (Software Upgrade) was completed. | 2023-10-30 14:52
Maquet Cardiopulmonary GmbH | Cardiohelp-i | 95 | - | The background of this Field Safety Corrective Action FSCA pertains to a nonconformance recognized in the production of the CARDIOHELP system A production tool viz a cable used to assess leakage current at the sensor panel connection hub Figure and Figure only had connection with out of contacts within that particular connection hub The root cause of the improper contact originates in an incorrect drawing of the production tool in question resulting in an incorrectly manufactured production tool Therefore testing with this cable is insufficient The insufficiency subsequently runs afoul of the EN Medical Electrical Equipment standard x D x D An additional error was identified in the process of the investigation When the CARDIOHELP Service Manual was created the measurement of the patient leakage current was not taken into account This error is unrelated and incidental to the finding explained in the previous paragraph x D x D x D | Affected CARDIOHELP-i are not requested to be returned and can be used as usual. A local Getinge representative will contact you to arrange the Electrical Safety Test according to IEC 62353 of the Cardiohelp-I and subsequent repair, ifnecessary. Please ensure that the affected device will be made available for the necessary verification as per scheduled date. | 2023-10-30 20:10
Maquet Cardiopulmonary GmbH | Cardiohelo-i | 95 | - | The IFU of the CARDIOHELP system states incorrect factory settings The device itself is working as intended and the error only refers to the IFU The following factory settings are incorrect for further information please refer to Annex I x D x D x D x D Mismatching information in IFU regarding PVen PAux Venous Bubble Intervention x D x D False statement in IFU regarding deactivated automatic locking in MECC Thapp x D x D | Replacement of the incorrect Instruction to Use | 2023-10-30 20:22
Ortho Clinical Diagnostics | ORTHO VISION Analyzer | - | - | During an internal review QuidelOrtho identified that Test ID included a calculated x D Rh Anti D or RhD Interpretation Result when in fact no Anti D column was used for the x D test x D Since no Rh Anti D column is used for Test ID the Interpretation Result is x D erroneously determined only via the Control column in the cassette If the Control column x D interpretation is Negative then the Rh result is Rh Pos positive If the Control column x D interpretation is Positive then the Rh result is a | Until the ADD with the fix is released, discontinue the use of test ID 10023 (4 ABO(FWD)-44 + (RVS)-A1,A2,B) for Blood Grouping and Rh Typing. • As an alternative, Test ID 10021 or 10022 must be used when performing Blood Grouping and Rh Typing. • QuidelOrtho is recommending a one-time retrospective review of stored donor blood using the BioVue test ID 10023 that may be erroneously identified as RhD positive. • Acknowledge your understanding of this notification by completing the enclosed Confirmation of Receipt form no later than November 10, 2023. • If your laboratory has experienced the issue described in this notification and you have not already done so, please report the occurrence to your local Global Services Organization (Previously Ortho Care). | 2023-10-31 13:47
Ortho Clinical Diagnostics | ORTHO VISION Max Analyzer | - | - | During an internal review QuidelOrtho identified that Test ID included a calculated x D Rh Anti D or RhD Interpretation Result when in fact no Anti D column was used for the x D test x D Since no Rh Anti D column is used for Test ID the Interpretation Result is x D erroneously determined only via the Control column in the cassette If the Control column x D interpretation is Negative then the Rh result is Rh Pos positive If the Control column x D interpretation is Positive then the Rh result is a | Until the ADD with the fix is released, discontinue the use of test ID 10023 (4 ABO(FWD)-44 + (RVS)-A1,A2,B) for Blood Grouping and Rh Typing. • As an alternative, Test ID 10021 or 10022 must be used when performing Blood Grouping and Rh Typing. • QuidelOrtho is recommending a one-time retrospective review of stored donor blood using the BioVue test ID 10023 that may be erroneously identified as RhD positive. • Acknowledge your understanding of this notification by completing the enclosed Confirmation of Receipt form no later than November 10, 2023. • If your laboratory has experienced the issue described in this notification and you have not already done so, please report the occurrence to your local Global Services Organization (Previously Ortho Care). | 2023-10-31 13:47
Ortho Clinical Diagnostics | ORTHO Optix™ Reader | - | - | During an internal review QuidelOrtho identified that Test ID included a calculated x D Rh Anti D or RhD Interpretation Result when in fact no Anti D column was used for the x D test x D Since no Rh Anti D column is used for Test ID the Interpretation Result is x D erroneously determined only via the Control column in the cassette If the Control column x D interpretation is Negative then the Rh result is Rh Pos positive If the Control column x D interpretation is Positive then the Rh result is a | Until the ADD with the fix is released, discontinue the use of test ID 10023 (4 ABO(FWD)-44 + (RVS)-A1,A2,B) for Blood Grouping and Rh Typing. • As an alternative, Test ID 10021 or 10022 must be used when performing Blood Grouping and Rh Typing. • QuidelOrtho is recommending a one-time retrospective review of stored donor blood using the BioVue test ID 10023 that may be erroneously identified as RhD positive. • Acknowledge your understanding of this notification by completing the enclosed Confirmation of Receipt form no later than November 10, 2023. • If your laboratory has experienced the issue described in this notification and you have not already done so, please report the occurrence to your local Global Services Organization (Previously Ortho Care). | 2023-10-31 13:47
Olympus Medical Systems Corp. | EVIS EXERA III GASTROINTESTINAL VIDEOSCOPE | 9 | - | The GIF TH Reprocessing Manual has mandatory instructions for cleaning and disinfection with an optional final Ethylene Oxide Gas EOG sterilization x D As a result of testing Olympus discovered EOG sterilization failures in which the GIF TH failed to achieve sterilization following the GIF TH s Reprocessing Manual As a result of the subsequent investigation Olympus determined that changes are required to the endoscope s channel drying steps prior to EOG sterilization in order to ensure effective endoscope sterilization Accordingly Olympus has updated the instructions for the endoscope channel s drying process and validated that the new updated channel drying steps will result in GIF TH endoscope EOG sterilization The revised instructions Dry the endoscope in Chapter of the GIF TH Reprocessing Manual are provided in the attached Addendum x D | Deliver the FSN+the Addendum attached to the end user and ensure all reprocessing personnel are completely knowledgeable and thoroughly aware of the contents. | 2023-10-31 14:00
Randox Laboratories | Total Bilirubin Vanadate Oxidation Method | - | - | Randox have had reports of elevated patient results using Plasma lithium heparin samples x D with Total Bilirubin reagent TBil Vanadate Oxidation Method catalogue numbers BR x D BR and BR Randox are currently investigating Plasma lithium heparin sample x D interference with the sample collection and preparation steps For this reason the only x D sample type that is recommended at this time is Serum Following our investigation Randox x D will complete a further customer notice to update our Instructions for Use IFU | - Total Bilirubin Vanadate Oxidation Method can only to be used on Serum samples. • Review results generated with Plasma (lithium heparin) samples in line with the clinical profile of the patient. • Discuss the contents of this notice with your Medical Director. • Complete and return the response form 12187-QA to technical.services(5)randox.com within five working days | 2023-10-31 14:58
Medline Industries Inc. | Tracheostomy Clean and Care Tray | 6600 | - | the tracheostomy brush a component within the Tracheostomy Care and Cleaning Kits has a sharp edge at the tip that can potentially puncture tubing during use and or cause user injury There have been complaints of tracheostomy brush bristles detaching prior to use in the packaging and or during use for which medical intervention may be required if bristles were to reach the patient s airway Inventory of affected Tracheostomy Care and Cleaning Kits can be disposed of for credit | 1. Immediately check your stock for item number and lot number, listed on the destruction form, quarantine all affected product. Destroy all product belonging to the affected lot numbers listed on the destruction form. 2. Please return the completed enclosed destruction form listing the quantity of affected product destroyed. Even if you do not have any affected product please complete and return the form, as indicated on the destruction form. When we receive your completed destruction form your account will be issued a credit, if applicable. 3. If you have transferred possession of this product to another individual, you are required by US FDA regulations to notify them of this recall communication. | 2023-10-31 17:08
Beckman Coulter Inc. | UniCel DxI 800 Access Immunoassay System | 17 | - | Beckman Coulter has determined that Access UniCel DxI and UniCel DxI luminometer lock nuts fitted on multiple instrument serial numbers may contain Cadmium levels that are above the limit allowable by the Restriction of Hazardous Substances RoHS x D Beckman Coulter has determined that the UniCel DxI and UniCel DxI bulk feeder II chute block sensor fitted on multiple instrument serial numbers may contain Cadmium levels that are above the limit allowable by the Restriction of Hazardous Substances RoHS x D | • No action is necessary by end users. • Due to the environmental hazard of Cadmium, Beckman Coulter will dispose of all affected luminometer lock nuts, bulk feeder II chute sensor blocks, and/or the instruments that contain these parts, according to the applicable regulations. • During the next semi-annual preventative maintenance or repair visit, a Beckman Coulter service representative will: o Replace any affected UniCel DxI 600 and UniCel DxI 800 bulk feeder II chute sensor block with a RoHS-compliant block. o Due to the extended instrument downtime (loss of diagnostic capacity for up to a day) caused by extensive disassembly required to replace the luminometer lock nuts and extremely limited exposure risk, Beckman Coulter will affix an alert label to the back of any Access 2, UniCel DxI 600, and UniCel DxI 800 instruments that contain the affected luminometer lock nuts. This label notifies laboratory personnel that the luminometer lock nut Cadmium levels are above the allowable RoHS limit, and that the part will not be replaced. o Upon decommissioning of the instrument, Beckman Coulter will dispose of the affected part per applicable regulations. | 2023-10-31 17:09
Beckman Coulter Inc. | UniCel DxI 600 Access Immunoassay System | 17 | - | Beckman Coulter has determined that Access UniCel DxI and UniCel DxI luminometer lock nuts fitted on multiple instrument serial numbers may contain Cadmium levels that are above the limit allowable by the Restriction of Hazardous Substances RoHS x D Beckman Coulter has determined that the UniCel DxI and UniCel DxI bulk feeder II chute block sensor fitted on multiple instrument serial numbers may contain Cadmium levels that are above the limit allowable by the Restriction of Hazardous Substances RoHS x D | • No action is necessary by end users. • Due to the environmental hazard of Cadmium, Beckman Coulter will dispose of all affected luminometer lock nuts, bulk feeder II chute sensor blocks, and/or the instruments that contain these parts, according to the applicable regulations. • During the next semi-annual preventative maintenance or repair visit, a Beckman Coulter service representative will: o Replace any affected UniCel DxI 600 and UniCel DxI 800 bulk feeder II chute sensor block with a RoHS-compliant block. o Due to the extended instrument downtime (loss of diagnostic capacity for up to a day) caused by extensive disassembly required to replace the luminometer lock nuts and extremely limited exposure risk, Beckman Coulter will affix an alert label to the back of any Access 2, UniCel DxI 600, and UniCel DxI 800 instruments that contain the affected luminometer lock nuts. This label notifies laboratory personnel that the luminometer lock nut Cadmium levels are above the allowable RoHS limit, and that the part will not be replaced. o Upon decommissioning of the instrument, Beckman Coulter will dispose of the affected part per applicable regulations. | 2023-10-31 17:09
Beckman Coulter Inc. | Access 2 Immunoassay System | 17 | - | Beckman Coulter has determined that Access UniCel DxI and UniCel DxI luminometer lock nuts fitted on multiple instrument serial numbers may contain Cadmium levels that are above the limit allowable by the Restriction of Hazardous Substances RoHS x D Beckman Coulter has determined that the UniCel DxI and UniCel DxI bulk feeder II chute block sensor fitted on multiple instrument serial numbers may contain Cadmium levels that are above the limit allowable by the Restriction of Hazardous Substances RoHS x D | • No action is necessary by end users. • Due to the environmental hazard of Cadmium, Beckman Coulter will dispose of all affected luminometer lock nuts, bulk feeder II chute sensor blocks, and/or the instruments that contain these parts, according to the applicable regulations. • During the next semi-annual preventative maintenance or repair visit, a Beckman Coulter service representative will: o Replace any affected UniCel DxI 600 and UniCel DxI 800 bulk feeder II chute sensor block with a RoHS-compliant block. o Due to the extended instrument downtime (loss of diagnostic capacity for up to a day) caused by extensive disassembly required to replace the luminometer lock nuts and extremely limited exposure risk, Beckman Coulter will affix an alert label to the back of any Access 2, UniCel DxI 600, and UniCel DxI 800 instruments that contain the affected luminometer lock nuts. This label notifies laboratory personnel that the luminometer lock nut Cadmium levels are above the allowable RoHS limit, and that the part will not be replaced. o Upon decommissioning of the instrument, Beckman Coulter will dispose of the affected part per applicable regulations. | 2023-10-31 17:09
Beckman Coulter, Inc. | Electrode-K | 100 | - | Beckman Coulter has determined that K electrodes with the lot number and serial numbers listed in the table above were filled with an incorrect inner liquid |  Check the serial number (S/N) on the label of any K electrodes currently installed on your instruments: o If the S/N of the K electrodes is listed above:  Stop using the K electrode.  Dispose of the K electrode according to your laboratory protocol.  Replace the K electrode with one S/N that is not listed above.  Refer to the “Replace the Na, K, or CL Electrode” procedure in the instrument Instructions for Use (IFU).  Perform ISE calibration and QC according to the IFU, if you remove or replace an electrode.  Check the S/N on the box label of any K electrodes in your inventory: o If the S/N of the K electrodes is listed above, dispose of the K electrode according to your laboratory protocol.  No retrospective review of results is necessary. | 2023-10-31 23:31
Ortho Clinical Diagnostics | VITROS Chemistry Products CKMB Slides (60 Slides) | 0 | - | Per the VITROS CKMB Slides Instructions for Use IFU Principals of the Procedure section the spread layer of the CKMB slide contains goat antihuman CK MM antibodies which inhibit CK MM muscle activity and of the CK MB heart activity The remaining CK activity represents of the total CK MB isoenzyme activity plus any CK BB which is relatively rare x D Per the current VITROS CKMB Slides IFU Known Interferences Section a Total CK activity greater than U L may result in falsely elevated CK MB results It is limited at U L by design which is due to the level of goat antihuman CK MM antibodies added during manufacturing Samples with Total CK U L should be diluted prior to analysis Therefore CK will not typically impact VITROS CKMB results below the CK result of U L x D Ortho Clinical Diagnostics QuidelOrtho has confirmed the issue and determined that lots manufactured from Coating of VITROS CKMB Slides do not adequately inhibit CK MM up to a total CK U L as stated in the IFU x D | • Per the VITROS CKMB Slides IFU, evaluate the Total CK of the sample. Elevated Total CK activity (>800 U/L) may cause a significantly elevated VITROS CKMB Slides result. If the CK activity exceeds 800 U/L, consult with your Medical Director. • Consider an alternate method of testing for samples with CK activity greater than 800 U/L. • Complete the enclosed Confirmation of Receipt form no later than 10-NOV-2023. • If your laboratory chooses not to use lots from the affected coating, QuidelOrtho will credit your account. • Save this notification with your User Documentation or post this notification by each VITROS 250/350/5,1 FS/4600/5600/XT 3400/XT 7600 System until the issue is resolved. • Please forward this notification if the product is distributed outside of your facility. | 2023-11-02 09:37
Ortho Clinical Diagnostics | VITROS Chemistry Products CKMB Slides (18 Slides) | 0 | - | Per the VITROS CKMB Slides Instructions for Use IFU Principals of the Procedure section the spread layer of the CKMB slide contains goat antihuman CK MM antibodies which inhibit CK MM muscle activity and of the CK MB heart activity The remaining CK activity represents of the total CK MB isoenzyme activity plus any CK BB which is relatively rare x D Per the current VITROS CKMB Slides IFU Known Interferences Section a Total CK activity greater than U L may result in falsely elevated CK MB results It is limited at U L by design which is due to the level of goat antihuman CK MM antibodies added during manufacturing Samples with Total CK U L should be diluted prior to analysis Therefore CK will not typically impact VITROS CKMB results below the CK result of U L x D Ortho Clinical Diagnostics QuidelOrtho has confirmed the issue and determined that lots manufactured from Coating of VITROS CKMB Slides do not adequately inhibit CK MM up to a total CK U L as stated in the IFU x D | • Per the VITROS CKMB Slides IFU, evaluate the Total CK of the sample. Elevated Total CK activity (>800 U/L) may cause a significantly elevated VITROS CKMB Slides result. If the CK activity exceeds 800 U/L, consult with your Medical Director. • Consider an alternate method of testing for samples with CK activity greater than 800 U/L. • Complete the enclosed Confirmation of Receipt form no later than 10-NOV-2023. • If your laboratory chooses not to use lots from the affected coating, QuidelOrtho will credit your account. • Save this notification with your User Documentation or post this notification by each VITROS 250/350/5,1 FS/4600/5600/XT 3400/XT 7600 System until the issue is resolved. • Please forward this notification if the product is distributed outside of your facility. | 2023-11-02 09:37
Philips Medical Systems Nederland B.V. | Allura Xper FD20 | 65 | - | Philips has identified a potential safety issue with the Philips Allura and Azurion Product Families having monoplane fixed ceiling mounted systems The ceiling mounted L arm contains a rotation cover that may potentially be susceptible to falling if a collision between the L arm and other hospital equipment i e an operating light were to occur Although the cover is retained by a safety chain if a collision were to occur the chain may also detach resulting in the cover falling on the patient user or bystander There have been incidents where minor injuries such as bruises or scratches were reported To date there have been no reported serious injuries or deaths related to this issue | As a remedy, Philips is replacing the existing covers of all the affected Philips Allura and Philips Azurion with the monoplane fixed ceiling mounted systems with a redesigned L-arm rotation cover (bolted instead of spring clips) to make sure it is resistant to external collisions. Philips will contact all affected customers to arrange for a Field Service Engineer visit to replace the L-arm rotation cover free of charge (reference FCO72200510 for the Philips Allura systems, including the Philips Poly G H5000, FCO72200511 for the Philips Azurion systems). | 2023-11-02 11:54
Philips Medical Systems Nederland B.V. | Allura Xper FD10 | 65 | - | Philips has identified a potential safety issue with the Philips Allura and Azurion Product Families having monoplane fixed ceiling mounted systems The ceiling mounted L arm contains a rotation cover that may potentially be susceptible to falling if a collision between the L arm and other hospital equipment i e an operating light were to occur Although the cover is retained by a safety chain if a collision were to occur the chain may also detach resulting in the cover falling on the patient user or bystander There have been incidents where minor injuries such as bruises or scratches were reported To date there have been no reported serious injuries or deaths related to this issue | As a remedy, Philips is replacing the existing covers of all the affected Philips Allura and Philips Azurion with the monoplane fixed ceiling mounted systems with a redesigned L-arm rotation cover (bolted instead of spring clips) to make sure it is resistant to external collisions. Philips will contact all affected customers to arrange for a Field Service Engineer visit to replace the L-arm rotation cover free of charge (reference FCO72200510 for the Philips Allura systems, including the Philips Poly G H5000, FCO72200511 for the Philips Azurion systems). | 2023-11-02 11:54
Philips Medical Systems Nederland B.V. | Allura Xper FD10 OR Table | 65 | - | Philips has identified a potential safety issue with the Philips Allura and Azurion Product Families having monoplane fixed ceiling mounted systems The ceiling mounted L arm contains a rotation cover that may potentially be susceptible to falling if a collision between the L arm and other hospital equipment i e an operating light were to occur Although the cover is retained by a safety chain if a collision were to occur the chain may also detach resulting in the cover falling on the patient user or bystander There have been incidents where minor injuries such as bruises or scratches were reported To date there have been no reported serious injuries or deaths related to this issue | As a remedy, Philips is replacing the existing covers of all the affected Philips Allura and Philips Azurion with the monoplane fixed ceiling mounted systems with a redesigned L-arm rotation cover (bolted instead of spring clips) to make sure it is resistant to external collisions. Philips will contact all affected customers to arrange for a Field Service Engineer visit to replace the L-arm rotation cover free of charge (reference FCO72200510 for the Philips Allura systems, including the Philips Poly G H5000, FCO72200511 for the Philips Azurion systems). | 2023-11-02 11:54
Philips Medical Systems Nederland B.V. | Allura Xper FD20 OR Table | 65 | - | Philips has identified a potential safety issue with the Philips Allura and Azurion Product Families having monoplane fixed ceiling mounted systems The ceiling mounted L arm contains a rotation cover that may potentially be susceptible to falling if a collision between the L arm and other hospital equipment i e an operating light were to occur Although the cover is retained by a safety chain if a collision were to occur the chain may also detach resulting in the cover falling on the patient user or bystander There have been incidents where minor injuries such as bruises or scratches were reported To date there have been no reported serious injuries or deaths related to this issue | As a remedy, Philips is replacing the existing covers of all the affected Philips Allura and Philips Azurion with the monoplane fixed ceiling mounted systems with a redesigned L-arm rotation cover (bolted instead of spring clips) to make sure it is resistant to external collisions. Philips will contact all affected customers to arrange for a Field Service Engineer visit to replace the L-arm rotation cover free of charge (reference FCO72200510 for the Philips Allura systems, including the Philips Poly G H5000, FCO72200511 for the Philips Azurion systems). | 2023-11-02 11:54
Philips Medical Systems Nederland B.V. | Azurion 5 M20 | 65 | - | Philips has identified a potential safety issue with the Philips Allura and Azurion Product Families having monoplane fixed ceiling mounted systems The ceiling mounted L arm contains a rotation cover that may potentially be susceptible to falling if a collision between the L arm and other hospital equipment i e an operating light were to occur Although the cover is retained by a safety chain if a collision were to occur the chain may also detach resulting in the cover falling on the patient user or bystander There have been incidents where minor injuries such as bruises or scratches were reported To date there have been no reported serious injuries or deaths related to this issue | As a remedy, Philips is replacing the existing covers of all the affected Philips Allura and Philips Azurion with the monoplane fixed ceiling mounted systems with a redesigned L-arm rotation cover (bolted instead of spring clips) to make sure it is resistant to external collisions. Philips will contact all affected customers to arrange for a Field Service Engineer visit to replace the L-arm rotation cover free of charge (reference FCO72200510 for the Philips Allura systems, including the Philips Poly G H5000, FCO72200511 for the Philips Azurion systems). | 2023-11-02 11:54
Philips Medical Systems Nederland B.V. | AZURION 7 M12 | 65 | - | Philips has identified a potential safety issue with the Philips Allura and Azurion Product Families having monoplane fixed ceiling mounted systems The ceiling mounted L arm contains a rotation cover that may potentially be susceptible to falling if a collision between the L arm and other hospital equipment i e an operating light were to occur Although the cover is retained by a safety chain if a collision were to occur the chain may also detach resulting in the cover falling on the patient user or bystander There have been incidents where minor injuries such as bruises or scratches were reported To date there have been no reported serious injuries or deaths related to this issue | As a remedy, Philips is replacing the existing covers of all the affected Philips Allura and Philips Azurion with the monoplane fixed ceiling mounted systems with a redesigned L-arm rotation cover (bolted instead of spring clips) to make sure it is resistant to external collisions. Philips will contact all affected customers to arrange for a Field Service Engineer visit to replace the L-arm rotation cover free of charge (reference FCO72200510 for the Philips Allura systems, including the Philips Poly G H5000, FCO72200511 for the Philips Azurion systems). | 2023-11-02 11:54
Philips Medical Systems Nederland B.V. | AZURION 7 M20 | 65 | - | Philips has identified a potential safety issue with the Philips Allura and Azurion Product Families having monoplane fixed ceiling mounted systems The ceiling mounted L arm contains a rotation cover that may potentially be susceptible to falling if a collision between the L arm and other hospital equipment i e an operating light were to occur Although the cover is retained by a safety chain if a collision were to occur the chain may also detach resulting in the cover falling on the patient user or bystander There have been incidents where minor injuries such as bruises or scratches were reported To date there have been no reported serious injuries or deaths related to this issue | As a remedy, Philips is replacing the existing covers of all the affected Philips Allura and Philips Azurion with the monoplane fixed ceiling mounted systems with a redesigned L-arm rotation cover (bolted instead of spring clips) to make sure it is resistant to external collisions. Philips will contact all affected customers to arrange for a Field Service Engineer visit to replace the L-arm rotation cover free of charge (reference FCO72200510 for the Philips Allura systems, including the Philips Poly G H5000, FCO72200511 for the Philips Azurion systems). | 2023-11-02 11:54
INOVA Diagnostics, Inc. | Quanta Flash Calprotectin Reagents | 19 | - | As a result of an investigation into the QUANTA Flash Calprotectin Reagents lot number x D we have confirmed an error with the expiration date listed on the Special Wash x D bottle Part Number lot number within the kit The expiration date on x D the Special Wash bottle is incorrectly listed as instead of The x D expiration dates on the outer packaging of the kit and on the Reagent Cartridge are x D correctly listed as x D Inova Diagnostics Inc a Werfen Company is initiating this advisory notice This advisory x D notice only affects the QUANTA Flash Calprotectin Reagents lot number | • Discard any remaining stock of the QUANTA Flash Calprotectin Reagents lot number 233854 after 2023-11-01. Follow local regulatory requirements for proper disposal. • Please distribute this information to all appropriate personnel in your laboratory and forward this information to all parties that may use this product, including others to whom you may have transferred this product. • Retain a copy of this letter in your laboratory records. | 2023-11-02 13:45
Medtronic, Inc. | LINQ™ II LNQ22 | 24 | - | In April Medtronic received an initial report of an ICM implanted October that experienced oversensing and undersensing resulting in clinician notifications of false tachycardia episodes It was stated the episodes began in March and reprogramming was attempted In May oversensing and interference noise was again reported Further troubleshooting was attempted but the noise was still present Therefore the ICM was replaced in May seven months after implant Through destructive analysis completed at the end of August it was found the reported oversensing noise was caused by moisture ingress under the titanium Ti foil of one of the tissue contacting electrodes that led to corrosion of the device s E electrode x D x D Investigation using acoustic microscopy identified a dark band in the laser bond from the edge of the Ti Foil visual signature on the electrode visible only via acoustic microscopy that indicated delamination of Ti foil from the Ti sputter used to bond the Ti foil to the device s sapphire wafer Penetrating dye testing confirmed a moisture pathway between the Ti foil and the Ti sputter see Figure red box Analysis of bond patterns determined that the laser bonder the oldest laser bonder used in the LINQ II manufacturing process which was decommissioned on Sep has a unique bond pattern that is different from the other laser bonders currently in use The bond pattern and laser characteristics e g power focus sequence led to a bond area near the electrode that can have a marginal bond potentially leading to a moisture pathway However a visible dark band on the electrode alone does not necessarily indicate a device will fail as there could be sufficient bonds within the dark line potential moisture pathway to prevent the forward progress of moisture x D x D In the absence of any manufacturing data to identify devices produced by the bonder that are at higher risk of being susceptible to moisture ingress Medtronic investigated CareLink data to better understand and characterize the issue Through a longitudinal investigation of the seconds of daily ECG that devices can send to CareLink provided the patient s device is enrolled in CareLink the patient s home monitor Is connected to CareLink and the patient is regularly in wireless transmission range of their home monitor including evaluation of internal device metrics correlated with noise Medtronic identified devices produced by the bonder with a marked transition from normal operation which includes some noise for an estimated of transmissions to potentially problematically noisy device operation with increased major noise and potentially leading to overall signal attenuation The devices identified by this evaluation included all returned devices confirmed to have electrode corrosion as of the time the data for the investigation was pulled May x D Medtronic also evaluated the CareLink records for devices produced by other laser bonders and found zero devices with the same performance transition x D x D Medtronic concludes that the devices identified in our assessment are at a higher risk of electrode corrosion and conversely that the broad population of devices that do not show this performance transition are at lower risk of having experienced moisture ingress though moisture ingress may still occur at some point in the future x D x D Based on the devices identified in the CareLink assessment the projected failure rate for devices due to moisture ingress leading to electrode corrosion is over the year LINQ II device lifetime across the population of the laser bonded devices x D | • Please review your inventory for devices with serial numbers listed in Table 1 of the FSN. Identify, quarantine, and return non-implanted devices. Your local Medtronic Representative can assist you as necessary>. • Encourage enrollment in and regular transmissions to CareLink. o Medtronic will apply recurring algorithmic searches on CareLink for the noise pattern and notify the clinician if present. The information used to identify this pattern is not visible to the clinician through CareLink. No further action is required for patients regularly transmitting to CareLink. • For patients not followed in CareLink: o Consider whether enrolling in CareLink is an option, per HRS/EHRA/APHRS/LAHRS guidance. Ongoing CareLink monitoring will reduce the potential for episodes caused by amplified noise to overwrite true episodes before they are reviewed. o If noise interferes with the ability to assess the patient’s rhythm or reason for monitoring, contact your Medtronic Representative. • If the ICM is no longer in use, no further action is necessary. | 2023-11-02 15:39
Medtronic, Inc. | CareLink SmartSync | 64 | - | There is a rare potential for reduced or no energy output during high voltage HV therapy typically J in implantable cardioverter defibrillators ICDs and cardiac resynchronization therapy defibrillators CRT Ds manufactured with a specific glassed feedthrough including currently available ICDs and CRT Ds x D x D There is an increased potential for a reduced or no energy HV therapy in the AX B configuration when all the following conditions are met x D The device has a glassed feedthrough manufactured after July x D There is significant separation of the layers of insulation materials in the feedthrough components of the device header x D An unintended current pathway forms within the void created by the insulation separation capable of conducting high levels of current during HV therapy x D x D When an unintended current pathway is detected during HV therapy the Short Circuit Protection SCP feature may trigger This behavior can be intermittent both full energy and reduced energy HV therapies within the same episode have been observed SCP events may also be lead related for both lead related and device related unintended current pathways the defibrillation waveform is truncated early in the energy delivery sequence resulting in reduced or no energy being delivered J x D | November 2023 Communication Install SmartSync software version 3.13.3. | 2023-11-04 18:46
Abbott Molecular Division Inc. | Alinity m HCV AMP Kit | 2 | - | Please find attached FSN letter | Please find attached FSN letter | 2023-11-06 11:36
Medtronic XOMED, Inc | NIM TriVantage EMG Endotracheal Tube | 1136 | - | A total of complaints were received by Medtronic reporting events in which the user has experienced noise from the NIM System lead off or high impedance issues leading to a Lead Off error message or intermittent connection issues or loss of functionality after the patient has been intubated with a TriVantage EMG Endotracheal Tube A shift within the supplier of raw material occurred during the November timeframe which caused a significant drop in reliability The root cause analysis of the issue is still under investigation | • Identify, segregate, and quarantine affected products within your inventory. The list of affected lot numbers is included in Attachment A. Return affected product in your inventory to Medtronic. Instructions on how to return this form and any impacted products to Medtronic can be found on the Customer Acknowledgment Form. Your local Medtronic Representative can assist you with the initiation of the return. • Please complete and return the Customer Acknowledgment Form even if you do not have the affected product | 2023-11-07 14:04
Respironics, Inc. | BiPAP A40 Pro International | 33 | - | There is an update to the Instructions for Use IFU for the labeling of this device Keredusy has been removed as a disinfection method due to unknown long term effects of this disinfection method on the device components | Philips has updated the BiPAP A40 Pro, BiPAP A40 EFL, BiPAP A30 EFL, and System One humidifier IFU (attached to this letter). The next revisions to the A40/A30 IFUs will incorporate the information provided in the attached addendums. | 2023-11-07 15:45
Philips | System One Heated Humidifier INTL | 33 | - | There is an update to the Instructions for Use IFU for the labeling of this device Keredusy has been removed as a disinfection method due to unknown long term effects of this disinfection method on the device components | Philips has updated the BiPAP A40 Pro, BiPAP A40 EFL, BiPAP A30 EFL, and System One humidifier IFU (attached to this letter). The next revisions to the A40/A30 IFUs will incorporate the information provided in the attached addendums. | 2023-11-07 15:45
Philips | A Series, Sys One, Heated Hum, INTL | 33 | - | There is an update to the Instructions for Use IFU for the labeling of this device Keredusy has been removed as a disinfection method due to unknown long term effects of this disinfection method on the device components | Philips has updated the BiPAP A40 Pro, BiPAP A40 EFL, BiPAP A30 EFL, and System One humidifier IFU (attached to this letter). The next revisions to the A40/A30 IFUs will incorporate the information provided in the attached addendums. | 2023-11-07 15:45
Fritz Stephan GmbH | SOPHIE | 28 | - | Fritz Stephan GmbH has been informed about a case where the measuring menu of a Sophie was opened x D without a user input x D A defective heating cable was identified as the cause of this behaviour If due to a defect an inadvertent contact x D of the live wire with the shielding occurs the signal inputs of the operating buttons that are no longer used in x D the Sophie can be triggered by the resulting disturbances This interference can cause the measuring menu x D to open and close by itself thereby impairing further operation of the unit x D During our internal investigation it was also found that in the event of a direct short circuit in the adapter cable x D of the disposable hose systems or the heating wire itself the power supply for the heating could toggle This x D can subsequently lead to a drop in the internal V voltage supply and thus to failure of the ventilator x D A direct short circuit in the adapter cable of the disposable tubing systems or the heating wire is very unlikely x D but cannot be completely ruled out x D We are not aware of any incidents in connection with such a direct short circuit | Fritz Stephan GmbH will provide a software update that intercepts the triggering of the unused signal inputs and makes them ineffective. At the same time, the internal error detection of the heating system will be adapted so that the heating no longer switches on independently in the event of a detected error. - If a fault of the heating system is detected and alarmed, the heating must be switched on again manually by the user in the heating menu after the fault has been remedied. | 2023-11-09 10:30
GE Hualun Medical Systems Co. Ltd | Definium Tempo, Definium Tempo Pro | 1 | - | Three of four Mounting bolts loose which hold Tube and collimator on Definium Tempo Pro This situation is clinically hazardous as the large mass could fall potentially resulting in catastrophic harm if it were to strike an individual on the head or trunk The associated hazardous situation is Cat Fall or tip of heavy large accessible object with significant impact or pressure on head or trunk | Inspect and correct all missing service steps identified as part of the investigation primary issue as well as all R/L issues identified on the Definium Tempo / Tempo Pro products. Update affected Service Documentation and physical copy to be left onsite with each system. | 2023-11-09 14:27
CareDX Pty Ltd | AlloSeq Assign | 1 | - | Expression variant display and reporting was not taken into account when developing the Ctrl G shortcut feature and P only or G only Report Display Options and is not included in the change control form The root cause can be considered a design deficiency | CareDx will update the IFU to remove references to the Ctrl+G shortcut and indicate that the P only and G only selections should not be used in the Report Display Options. CareDx will release a software update to address the design deficiency with relation to the Ctrl+G shortcut and P only and G only Report Display Options. | 2023-11-09 16:01
B.Braun Melsungen AG | Coroflex ISAR Neo | 1 | - | Based on customer feedback B Braun Melsungen AG Vascular Systems has become aware of incorrect labeling of at least one of the products in the batch Coroflex ISAR Neo x mm REF LOT F x D A product with the stent size x mm was distributed in a product box marked with the size x mm The sterile packaging and the product itself on its adapter are marked with the correct size x D | The incorrect labeling has so far only been confirmed for a single product. However, it cannot be excluded that other incorrectly labeled products were placed on the market. Furthermore, it cannot be excluded that there is a label with the size 2.5 x 19 mm on the sterile packaging of a 2.75 x 24 mm product. Therefore, both batches are considered affected and are recalled. | 2023-11-12 10:08
Cook Medical Europe Limited | Quantum TTC Biliary Balloon Dilator | 50 | - | Cook Endoscopy Wilson Cook Medical Inc is initiating a voluntary removal of affected x D QBD devices from the market because they are nonconforming They are manufactured x D correctly but do not comply the design requirements Although only incidents have x D occurred the likelihood of further incidents is probable to frequent and Cook is taking a x D proactive approach to remove the devices from the field and stop production until a root x D cause and corrective action is found x D x D It has been determined that nonconforming Quantum TTC Biliary Balloon Dilators x D QBD were distributed A sampling of device manufactured after April x D identified the devices were not able to inflate to the clinical working pressure without x D leakage The failures were across all QBD sizes x D EU MDR defines a field safety corrective action as a corrective action taken x D by a manufacturer for technical or medical reasons to prevent or reduce the risk of a x D serious incident This field action is a removal of devices to reduce the risk of a serious x D incident Since the affected devices were distributed in the EU this is considered an x D FSCA in the affected EU member states | Identify, Quarantine and Return Device. | 2023-11-12 15:19
AB Sciex | IonDrive Turbo V source | 2 | - | Leak of a highly flammable solvent at the ion source probe | SCIEX will be changing the electrode adjustment nut material. | 2023-11-12 15:21
AB Sciex | Turbo V source | 2 | - | Leak of a highly flammable solvent at the ion source probe | SCIEX will be changing the electrode adjustment nut material. | 2023-11-12 15:21
Leica Biosystems Nussloch GmBH | HistoCore PEGASUS // HistoCore Pegasus Plus | 1 | - | As part of our post market surveillance we have been made aware of an issue regarding poorly processed and or damaged biopsy tissue specimens on the HistoCore PEGASUS HistoCore PEGASUS Plus resulting from incorrect assignment of the carryover value set for created or edited protocols If the setup of carryover is lower than the actual carryover it may potentially cause tissue damage mainly under processed tissue | As an immediate action, please ensure you are selecting the correct setting of the carryover value for any created or edited protocols, following the instructions in the Instruction for Use (IFU), chapter 6, "Protocol Setup", section 6.1.5, "Carryover setting": "When a reagent drains out of a retort, a certain amount of the reagent remains in the retort and mixes with the next reagent in the protocol. The carryover setting is an estimate of the amount of a reagent that is carried over from one protocol step to the next. The reagent management system takes the carryover setting into consideration when determining reagent concentrations. The residual reagent carried over can come from: • Reagent retained on retort walls and baskets • Reagent on and infiltrated into tissue • Reagent on and infiltrated into small-tissue carriers (e.g. biopsy pads, biopsy cassettes, wraps, etc.)." "Because biopsy pads can have up to 10 times the carryover of standard cassettes, it is important to set a truly representative carryover value in your protocols (as it is to accurately enter the number of cassettes in each run). If you set too high a carryover value, the system calculates an excessive degradation of reagents, which will require you to replace them sooner than necessary. With too low a setting, the system will consider that reagents are purer than they are. Thus, you will use reagents beyond their optimal effectiveness, resulting in poor processing quality." Also, please note section 6.1.4, "Protocol validation" which states, "Supervisors creating or editing protocols (or copying pre-defined protocols) can validate them in the software. This serves as a sign that the protocols have passed the laboratory's validation tests and can be used for regular clinical processing. Supervisors can also make valid protocols invalid." | 2023-11-13 14:47
Philips Medical Systems (Cleveland), Inc. | Brilliance CT 64 Slice | 30 | - | Philips has identified two software issues affecting the performance of Ingenuity CT and Brilliance x D CT systems with software version x D Issue Ingenuity CT systems cannot be registered to the Remote Service Network x D These systems cannot be connected to the Philips Remote Service Network RSN resulting in x D no access to the remote services such as software distribution or remote desktop connection x D abilities x D Issue Customer service intellectual property tool cyber security vulnerability x D This issue impacts both Ingenuity CT and Brilliance CT systems with version If an x D unauthorized user utilizes counterfeit service tool security certificates they may gain access to x D patient data the service manual and restricted service functions of CT devices x D Philips has not received any reports of adverse events associated with these issues as of October x D | Philips will contact customers to schedule a time for a Philips Field Service Engineer (FSE) to visit site and install the solution (refer to FCO72800759) to improve system performance. | 2023-11-14 11:07
Philips Medical Systems (Cleveland), Inc. | Ingenuity CT | 30 | - | Philips has identified two software issues affecting the performance of Ingenuity CT and Brilliance x D CT systems with software version x D Issue Ingenuity CT systems cannot be registered to the Remote Service Network x D These systems cannot be connected to the Philips Remote Service Network RSN resulting in x D no access to the remote services such as software distribution or remote desktop connection x D abilities x D Issue Customer service intellectual property tool cyber security vulnerability x D This issue impacts both Ingenuity CT and Brilliance CT systems with version If an x D unauthorized user utilizes counterfeit service tool security certificates they may gain access to x D patient data the service manual and restricted service functions of CT devices x D Philips has not received any reports of adverse events associated with these issues as of October x D | Philips will contact customers to schedule a time for a Philips Field Service Engineer (FSE) to visit site and install the solution (refer to FCO72800759) to improve system performance. | 2023-11-14 11:07
Maquet Cardiopulmonary GmbH | CARDIOHELP-i | 84 | - | The quality issue under investigation is applicable to the Connection cable venous probe L m material accessory as well as spare part The part causing the nonconformance under investigation is responsible for data x D | Replacement of the Venous Probe cable. | 2023-11-14 16:07
SCHILLER Medical SAS | Defigard touch 7 | 2 | - | Please find attached FSN letter | Please find attached FSN letter | 2023-11-15 13:05
AT-OS SRL | AWD655 | 8 | - | The reason is due to an incorrect part installed on some of our machines x D Due to the growing difficulty of finding materials in order to satisfy market demands in a x D reasonable time one of our suppliers used a component on our machines which today we discover x D is not suitable For that reason we immediately inform you of an important recall action to be x D scheduled | replace the fuseholders | 2023-11-15 15:31
Covidien LLC | Kangaroo Epump Enteral Feed and Flush Pump with Pole Clamp | 15900 | - | The issue relates to air entry into the feed set tubing | FOLLOW the steps described in Attachment A to assess and monitor for entry of air into the feed set. Retain a copy of this notice with the product. 2) COMMUNICATE with all personnel that utilize these feeding sets of the potential risk of air entering feeding set tubing. 3) NOTIFY any customers to whom you may have distributed or forwarded affected product about this field correction. 4) REVIEW, complete, sign and return the enclosed Acknowledgement Form by following the directions on the Form. 5) MAINTAIN awareness of this notice until all affected product has been utilized. | 2023-11-15 20:58
Olympus Winter & lbe GmbH | ENDOEYE HD II | - | - | Olympus has identified a total of complaints including adverse events since September to August related to pink or green coloration of the image including cases with reported delays of treatments and or prolonged surgery Olympus investigated the complaints and identified that the CCD chip that produces the image became damaged The CCD chip is sensitive to heat and mechanical shocks Olympus s risk management is taking this into account and the instructions for use mention necessary steps to prepare the video telescope prior to usage x D This Field Corrective Action reminds customers to follow the steps in the instructions for use especially to inspect the image prior to a clinical procedure and always have a spare laparoscope available x D Olympus is currently investigating technical solutions to address this issue x D | Please make sure that all medical personnel are completely knowledgeable and thoroughly trained in the ENDOEYE HD II Instructions For Use (IFU). The purpose of this Field Corrective Action is to reinforce the requirement in the IFU that users inspect the image prior to a clinical procedure (as described in Section 7.4 Testing), always have a spare laparoscope available (as described in Section 2.5 General Dangers, Warnings and Cautions) and to check the image for the spare equipment. | 2023-11-17 16:11
B.Braun Melsungen AG | SeQuent Neo | 30 | - | Based on customer feedback B Braun Melsungen AG Vascular Systems has become aware of incorrect labeling of at least one of the products from the batch SeQuent Neo x mm REF D LOT F x D A product with the balloon size x mm was in a product box marked x mm The sterile packaging and the product itself on its adapter are marked with the correct size | Actions to be taken by the local B. Braun organization / distributor: 1. Please identify all affected customers in your country. 2. Please adapt the attached FSN template to your local language / contact information. 3. Please inform your customers about this FSCA using the FSN and collect affected products on their stock. 4. Please add your National contact point information in section 5 of the attached FSCA report form and submit the report to your authority. | 2023-11-20 09:10
B.Braun Melsungen AG | SeQuent® Please NEO | 9 | - | Based on customer feedback B Braun Melsungen AG Vascular Systems has become aware of incorrect labeling of at least one of the products in the batch SeQuent Please Neo x mm REF LOT H x D A product with the balloon size x mm was in a product box marked x mm The sterile packaging and the product itself on its adapter are marked with the correct size | 1) Please identify all affected customers in your country. 2) Please adapt the attached FSN template to your local language / contact information. 3) Please inform your customers about this FSCA using the FSN and collect affected products on their stock. 4) Please add your National contact point information in section 5 of the attached FSCA report form and submit the report to your authority. | 2023-11-20 09:20
Siemens Healthcare GmbH | ARTIS icono biplane | 4 | - | In Electrophysiology EP catheter laboratories there are many video sources to be displayed in one split screen layout x D on the Large Display In case of shrinking the recording system video source artefacts and or loss of ECG IECG details x D might occur | A dedicated scaling algorithm for ARTIS icono is in development which is meant to improve the ECG/IECG visualization quality of recording systems video outputs displayed in shrink split screen window size on the Large Display. | 2023-11-20 15:20
Siemens Healthcare GmbH | ARTIS icono floor | 4 | - | In Electrophysiology EP catheter laboratories there are many video sources to be displayed in one split screen layout x D on the Large Display In case of shrinking the recording system video source artefacts and or loss of ECG IECG details x D might occur | A dedicated scaling algorithm for ARTIS icono is in development which is meant to improve the ECG/IECG visualization quality of recording systems video outputs displayed in shrink split screen window size on the Large Display. | 2023-11-20 15:20
Philips Medical Systems Nederland B.V. | Panorama HFO | 4 | - | Philips has been informed of an event in which the structural integrity of the Panorama T HFO system components failed due to unintended excessive pressure buildup of helium gas during a magnet quench x D x D During a quench a large amount of helium evaporates and is vented outside the building through a venting system If an unknown blockage is present in the venting system and the pressure exceeds design limits the structural integrity of the system may be compromised The magnet may experience an unsolicited quench during normal use or will quench when initiated in an emergency situation by the operator pressing the Magnet EMERGENCY STOP button x D x D Philips has received one complaint of system and property damage after quenching the magnet on Panorama T systems as of November There was no report of injury or serious harm | Philips is requesting that customers immediately stop using the MR systems until an inspection can be completed. Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to inspect your system starting in December 2023. (Reference FCO78100572) | 2023-11-20 16:16
Medtronic MiniMed | Guardian 4 | 490 | - | The manufacturing team reported a situation with Filmetrics equipment where it was noticed that more than values were displayed in the Filmmaper screen Filmetrics equipment software Per the Filmetric Measurement Process the system is configured to measure GOX thickness on three working electrodes Upon finishing substrate plate measurements values are expected per substrate plate x D Upon further investigation it was noted that the data from the previously analyzed substrate plate was duplicated in the output excel file for the first following substrate plate This resulted in a missing data event for the subsequent substrate plate x D | Determine if you have sensors that are potentially impacted, please follow the steps below: 1. Gather all your Guardian™ 4 sensors. 2. Find the LOT number on the individual sensor package or sensor box. a. Check your sensor LOT number(s) on our website at https://www.medtronic-diabetes.com/FA1379 or in the list below to see if your sensors are impacted. a. If your LOT number(s) ARE NOT impacted, you can continue to use your sensors. b. If your LOT number(s) ARE impacted, do not use these sensors and request free replacement sensors by contacting Helpline contact at 8002472288 c. If you cannot identify the LOT number of the sensor you are currently wearing, remove the sensor and replace your sensor with a non-impacted sensor. 3. If you only have impacted sensors remaining, call us immediately so that we can ensure you receive replacement sensors as soon as possible. You should revert to your backup plan as recommended by your healthcare professional while you await your replacement sensors. 4. Dispose of all impacted sensors. | 2023-11-20 20:34
ResMed | ResMed masks with magnets | - | - | Description of Change x D ResMed is updating its existing contraindications and warnings to further inform patients and x D healthcare professionals on the safe use of ResMed masks with magnets This is in response to x D new information obtained through post market surveillance and industry practices related to possible x D magnetic interference when in close proximity to certain medical devices x D Update to Contraindication x D In an update to the previously contraindicated metallic aneurysm hemostatic clips in the head x D and metallic splinters in one or both eyes following a penetrating eye injury ResMed is further x D contraindicating the use of its masks with magnets for patients where they or anyone in close x D physical contact while using the mask e g bed partner have the following x D Active medical implants that interact with magnets i e pacemakers implantable x D cardioverter defibrillators ICD neurostimulators cerebrospinal fluid CSF shunts x D insulin infusion pumps x D Metallic implants objects containing ferromagnetic material i e aneurysm clips flow x D disruption devices embolic coils stents valves electrodes implants to restore x D hearing or balance with implanted magnets ocular implants metallic splinters in the x D eye x D Note not all models or manufactured variants of medical devices listed in the contraindications x D are affected by external magnetic fields Only those that interact with magnets or contain x D ferromagnetic materials present the potential for magnetic interference ResMed Pty Ltd x D Elizabeth Macarthur Drive Bella Vista NSW Australia ResMed com x D Update to Warning x D The safe distance to the magnets has been increased from inches mm and patients are x D now warned to keep the mask magnets at a safe distance of inches mm away from x D implants or medical devices that may be adversely impacted by magnetic interference This x D applies to patients or anyone in close physical contact with the mask magnets | ResMed is updating the contraindications and warning sections of the Instructions for Use (IFU) in the user guides of affected masks with magnets. In addition, further information related to this notice will be made available at www.resmed.com/magnetupdate. ResMed will coordinate with customers for a replacement alternative mask without magnets for contraindicated patients. Actions to be taken by Homecare providers, Hospitals and Distributors 1. Complete and return the acknowledgment form online (where available) or at the end of this notice. 2. Immediately provide a copy of this notice and the physician letter to prescribing physicians, and/or any other relevant healthcare professionals, notifying them of the updated labelling (contraindications and warning). 3. Immediately provide a copy of the patient letter to all patients currently using a ResMed mask with magnets, notifying them of the updated labelling (contraindications and warning). Note, patients are instructed in the patient letter to contact their mask provider for a replacement mask if they are now contraindicated. 4. Provide a replacement mask for contraindicated patients to an alternative mask without magnets in a timely manner. Where an alternative mask is not available, inform patients to consult with their physician. 5. Instruct patients to consult their physician and/or manufacturer of their implant / other medical device if they require additional information on the potential adverse effects of magnetic fields for their particular device (as described in warnings). | 2023-11-21 11:55
Covidien llc | Valleylab™ FT10 | 217 | - | Medtronic received a complaint where investigation identified that recent versions of the Valleylab FT FT Energy Platform generator with specific software versions and exhibited an error which may incorrectly prohibit use of select LigaSure instruments with the generator The specific error displayed on the Generator as E Usage Limit may occur when the Generator performs an authentication check to confirm that the LigaSure instrument has not been used previously While this prohibition is intended to prevent re use of single use surgical devices the software error residing within the impacted generators may mis classify some instruments which had not been previously used as used In this case the E error occurs when the instrument is initially connected to the Generator and may require a different instrument to be selected for use | Advice on actions to be taken by the distributor and the user: • Immediately notify all personnel in all care environments in which the Valleylab™ FT10 Energy Platform is used about this notice. • Update ValleylabTM FT10 Energy Platform to software version 4.0.4 to eliminate this issue. o For customers already familiar with the software update process through VLEx, software may be directly updated. o For customers not familiar with the software update process through VLEx, your Medtronic representative will assist in updating your ValleylabTM FT10 Energy Platform to software version 4.0.4. Your Medtronic representatives will schedule servicing to update the software within the coming weeks. • Until the software is updated, the Valleylab™ FT10 Energy Platform and LigaSure™ devices can continue to be used as instructed in the User Guide and per your facility protocols. However, please note that the error messages precluding use of LigaSure™ devices may be encountered until software version 4.0.4 is installed. • Complete the attached Customer Acknowledgement Form and return it as directed to confirm your receipt and understanding of this information. | 2023-11-21 20:39
Respironics, Inc. | Trilogy Evo O2, International | 224 | - | During an internal review of the Trilogy Evo and Trilogy Evo O Ventilators Instructions for Use IFUs manuals Philips Respironics determined that the Contraindications Statement was incomplete This URGENT Field Safety Notice is intended to inform you of all contraindications for the devices Please note the devices can continue to be safely used in line with the mitigations described in this document There have been no safety issues reported from the field in relation to this issue x D The following contraindications apply to the Trilogy Evo and Trilogy Evo O Ventilators x D Instructions for Use Contraindications x D If the patient has any of the following conditions consult the patient s health care professional before using noninvasive ventilation x D An inability to maintain a patent airway or adequately clear secretions x D At risk to aspirate gastric contents x D Acute sinusitis or otitis media x D Epistaxis causing pulmonary aspiration of blood x D Hypotension x D The AVAPS AE therapy mode is contraindicated for invasive use and patients less than kg x D The AVAPS feature is contraindicated for patients less than kg x D x D The contraindications listed above in bold are either omitted or listed in other locations within the Instructions for Use and Clinical Manual These have been added to the attached manual addendum for insertion into your current manuals x D | Philips Respironics is correcting the IFUs related to the Trilogy Evo and Trilogy O2 Ventilator referenced in this letter. An addendum to your current manual has been provided with this letter. | 2023-11-22 10:38
Olympus Medical Systems Corp. | HIGH FLOW INSUFFLATION UNIT | - | - | Olympus has become aware of patients experiencing complications from over insufflation including arrythmias reported as short cardiac arrests gas embolism and one death during surgical procedures where UHI s were used | •Discontinue use of the UHI-4 until the root cause investigation is completed and you receive additional instructions from Olympus. •If your facility does not have alternatives or is unable to obtain alternatives, you may choose to use the UHI-4 while exercising extreme caution, after weighing the potential benefits of the procedure versus the potential risk to health of over-insufflation described above. Please Read Before Use: It is recommended to use the lowest intraabdominal pressure allowing adequate visualiza􀆟on of the opera􀆟ve field for each procedure to help reduce risk of complica􀆟ons related to over insuffla􀆟on. Complications related to over insufflation include gas embolism, arrhythmias (bradycardia, asystole, or cardiac arrest), prolonged or more complex procedures, delay to treatment, pneumothorax, hypoxia, subcutaneous emphysema, kidney or urinary problems. These complica􀆟ons could poten􀆟ally lead to death. | 2023-11-23 12:09
Bard Peripheral Vascular Inc. | Bard® Mission™ | - | - | BD received complaints for the product regarding mismatch between the coaxial and the needle in Mission kit devices Based on the event reported the internal diameter of the coaxial cannula may be smaller or larger than the external diameter of the biopsy needle and the length of the cannula may exceed the stated length on the label As a result the biopsy needle may not fit properly into the coaxial cannula preventing access to the target tissue x D x D | Advice for Clinical Users Healthcare facilities and providers should cease use of the affected product and discard. BD Actions BD has investigated this issue and actions are in place, including 100% inspection of these products, to prevent recurrence of this issue. Customer Actions • Cease use of any unused affected Bard® Mission® Disposable Core Biopsy Instrument Kits. • Identify and quarantine all unused affected Bard® Mission® Disposable Core Biopsy Instrument Kits. o Make a note of the lot numbers and destroy all unused affected units. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 11th December 2023. • Circulate this notice to all those who need to be aware within your organization or to any organization where the potentially affected products have been transferred. • If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: • Cease distribution. • Identify, quarantine, making a note of the lot numbers then destroy all unused affected Bard® Mission® Disposable Core Biopsy Instrument Kits. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 11th December 2023. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. | 2023-11-23 13:00
CardioRenal Systems, Int. | RenalGuard | 150 | - | leakage in the connection between the foley catheter connector and the set tubing | Items Reserve in isolation area in our warehouse in Riyadh | 2023-11-23 13:10
Beckman Coulter Biomedical GmbH | DxA Automation System | 1 | - | Your DxA has been updated with DxA Software Version Beckman Coulter has become aware that the issue described below might still occur x D The DxA Software has configuration settings to identify tests that are not to be performed based upon a specific tube type The DxA does NOT communicate this to the Remisol middleware software quickly enough to prevent the test to be run | Please continue to follow the “Action” written below until the follow up visit from your Field Service Engineer is completed. Inspect tubes routed to error regions and confirm that the tube type was appropriate per the assay IFU for the test performed. | 2023-11-23 21:22
Cristalens Industrie | ARTIS T PLE | 0 | - | We hereby inform you that we are recalling lots as a precautionary measure x D Cristalens Industrie has been informed of cases of intraocular lenses for which the power of the implanted lens does not correspond to the power desired by the practitioner In both situations the postoperative refractive results are not satisfactory x D After reviewing the manufacturing data of potentially suspected batches Cristalens Industrie did not detect any anomalies in the manufacturing traceability data nor in the manufacturing process x D Impacted products Intraocular lens Model ARTIS PLE Dioptry D to D ARTIS T PL E diopter D | As a precautionary measure, Cristalens Industrie has identified in the manufacturing process, the close production batches that may correspond to the calculated power and is recalling these batches. Users of potentially suspected batches have been notified | 2023-11-25 13:21
Dragerwerk AG & Co. KGaA | Atlan | 203 | - | Damage to the battery can lead to a reduction in capacity and thus to a reduced runtime in battery mode | 1) Send the safety notice to all customers as per the location list. 2) Send a supplement for the instructions for use of Atlan together with the „customer letter accompanying the supplement“. | 2023-11-26 11:48
Olympus Medical Systems Corp. | Evis Exera III Duodenovideoscope Olympus | 45 | - | This is voluntary corrective action of all TJF Q V Q V Q V duodenoscopes after becoming aware of recent reports of infections and positive cultures Olympus is reminding users to closely follow reprocessing instructions including periodic and annual inspections by Olympus Service as well as regular microbiological checks and informing users about updated manuals intended to improve comprehension of existing reprocessing steps Olympus is also reminding users that deviating from the reprocessing manual steps or use of a damaged or contaminated endoscope can present an infection risk to patients | 1.Carefully read the content of this Field Safety Notice (FSN). 2.Inspect your inventory of duodenoscopes and identify all TJF-Q190V/Q290V/Q170V duodenoscopes. 3.Ensure all reprocessing personnel are completely knowledgeable and thoroughly trained on the reprocessing instructions in the new Reprocessing Manuals. Meticulous cleaning of the forceps elevator recesses and attention to following all reprocessing instructions are required. 4.Please indicate on the enclosed reply form that you have received this notification and taken the required actions by filling out and returning the completed enclosed Reply Form back to your local Olympus representative latest by 08.12.2023. | 2023-11-26 16:41
Medtronic, Inc. | SYNCHROMED® II | 531 | - | Electromagnetic interference from an MRI on a SynchroMed II pump can cause the pump to enter a continuous telemetry active status known as Telemetry Mode Telemetry Mode is known to interfere with motor stall processing and logging but does not prevent the pump from resuming drug delivery after leaving the MRI magnetic field This issue is a different variation of existing Telemetry Mode x D If Telemetry Mode begins in an MRI while the SynchroMed II pump is sounding an alarm the pump will not resume drug delivery once removed from the MRI In this case drug delivery will only resume once Telemetry Mode is resolved by performing a follow up interrogation of the device after the MRI x D This is inconsistent with the SynchroMed II MRI Guidelines for Medtronic implantable infusion systems x D | Patient Management Recommendations: • Upon completion of an MRI scan, interrogate the pump with the Clinician Programmer (or Personal Therapy Manager) to end telemetry mode and resume drug delivery. • Consult the MRI Guidelines for additional information on MRI preparation and post-examination review, and motor stall recovery timing (see MRI Guidelines at www.manuals.medtronic.com). • Remind your patients about the importance of interrogating the SynchroMed™ II pump after an MRI to ensure continuation of therapy. • Educate patients, caregivers, and family members to recognize the signs and symptoms associated with intrathecal drug therapy underdose or withdrawal. Patients receiving intrathecal baclofen therapy (e.g., Lioresal Intrathecal) are at higher risk for adverse events, as baclofen withdrawal can lead to a life-threatening condition if not treated promptly and effectively. Customer Required Actions: • Share this notice with all those who need to be aware of this issue within your organization and maintain a copy of this notice in your records. • Please complete and return the Customer Acknowledgment Form enclosed with this letter, acknowledging that you have received this information. | 2023-11-26 18:05
Biomerieux | BIOFIRE TORCH Base System | 28 | - | When turning on BIOFIRE TORCH systems arcing inside of the power switch may result in carbon x D build up on the switch contacts The carbon build up may lead to excess heating inside of the power x D switch subsequently leading to deformation of the power switch case The deformation can create x D an opportunity for a power switch electrical short In addition to an electrical short deformation may x D result in an open circuit causing the power switch to fail This event would only occur after the x D product is in use | In this context, we request you to take the following actions. Please: - If a user determines that the BIOFIRE TORCH Base requires complete shutdown, each individual BIOFIRE TORCH module must be powered down prior to turning on the BIOFIRE TORCH Base, - ‘Reset’ the BIOFIRE TORCH Base weekly using the ‘Reset’ button, - Distribute this information to all appropriate personnel in your laboratory, retain a copy in your files, and forward this information to all parties that may use this product, including others to whom you may have transferred our product. - Complete the Acknowledgement Form in Attachment A and return it to your local bioMérieux representative (to be adapted at local level) to confirm receipt of this notice. It is important that you return the acknowledgement form to bioMérieux even if you determine that this urgent product correction notice/information/recommendation (to be adapted based on issue) does not impact your facility. | 2023-11-26 19:00
Nipro Corporation | Single Patient Dialysis Machine | - | - | the effect that certain treatment conditions have on UF performance | The upgrade to software version 1.706 | 2023-11-27 17:23
ICU Medical | Plum360 Infusion System | 2 | بسبب عطل فنى بالجهاز. | - | Both devices were isolated to ensure they were not used. | 2023-11-28 15:29
Ambu A/S | Ambu® aView™ 2 Advance | - | - | Description of the problem x D Ambu has received information on two incidents where Ambu aView Advance caught fire when mounted on the VESA holder of the aCart Compact due to the screws used entering the lithium ion battery of the device No patients or staff members were harmed during the incidents x D VESA holder is a generic standard that specifies dimensions of screw holes for mounting Ambu aView Advance to a cart or stand Therefore even though aCart Compact is not available in your market any other brand medical workstations with VESA holder and longer screw lengths could be used in your market resulting in a similar fire hazard x D Packaging of aCart Compact includes several choices of screw lengths mm as VESA holder is usually used for external medical monitors In these two cases customers mounted Ambu aView Advance on the VESA holder using too long screws meant for external monitors x D Upon conducting a comprehensive investigation we have identified the root cause of these incidents The fire hazard was a direct consequence of a battery short circuit which in turn was triggered by the use of excessively long screws to secure the device on the VESA holder The longer screw lengths inadvertently penetrated the lithium ion battery leading to the fire hazard x D Since VESA option is available in the product design we want to inform our customers of the risk of penetrating the lithium ion battery when using too long screws for mounting At the same time we would like to inform our customers on how to securely fasten the Ambu aView Advance In case your Ambu aView Advance is placed on a table or mounted on an IV pole as recommended practice there is no risk of penetrating the battery and associated fired hazards x D We urge our customers to be aware of the below Warning applicable to the Instructions for Use of Ambu aView Advance x D Use only M screws with the length of mm when mounting Ambu aView Advance on a VESA interface Using longer screw lengths will penetrate the lithium ion battery and result in a fire hazard and battery leakage which can cause severe burns smoke inhalation and skin irritation Using shorter screw lengths could result in unsecure device fastening x D The information in this Field Safety Notice is relevant for all versions of Ambu aView Advance x D Please communicate this information to relevant personnel within your organization Included with the Field Safety Notice you will find an insert for the Instructions for Use for Ambu aView Advance The insert should be read and kept together with Instructions for Use you received together with your Ambu aView Advance The information is also included in Appendix of this notice x D Ambu A S is not removing any Ambu aView Advance from the field devices remain available for use | Advise on actions to be taken by user: Within one month of receipt of this letter, please return confirmation of receipt of this Field Safety Notice (appendix 1). Please familiarize yourself with the information in the insert of Instructions for Use and keep the insert together with the Instructions for Use booklet. Transmission of this Field Safety Notice: This notice needs to be passed on to all those who this might concern within your organization or to any organization where the devices could have been transferred. Please transfer this notice to other organizations on which this action has an impact. Please maintain awareness of this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. Patient safety remains our highest priority. If you have additional questions regarding this information, please contact your local Ambu sales representative | 2023-11-28 16:41
Randox Laboratories | Serology ORCH igM POSITIVE CONTROL | - | - | Randox Laboratories can confirm there has been a decrease in the reactivity of HSV Type x D IgM in the Serology ToRCH IgM Positive Control SR lot SR when tested on the x D DiaSorin Liaison XL While still testing positive for HSV Type IgM we advise that all x D customers cease the usage of the control for this marker as a precaution The other markers x D within this lot are not affected Please discard all copies of the IFU and download the latest x D version from www randox com | - Discuss the contents of this notice with your Medical Director. • Complete and return the response form 12187-QA to technical.servicesPrandox.com within five working days. • Please discard all copies of the IFU and download the latest version from www.randox.com. | 2023-11-28 17:16
Nobel Biocare AB | IPS CaseDesigner | 57 | - | KLS Martin Germany discovered internally an issue with the NHP Natural Head Positioning in IPS CaseDesigner When adjusting the NHP by a large number of rotations rotation steps the NHP adjustments suffer from a numerical instability and the visualisation is distorted The issue is identified as a bug in the software If not detected during the planning phase could lead to a longer surgery | Nobel Biocare advise that a new software release version will be made available in which the NHP adjustments are improved. As soon as this new version is available and the software is started with an internet connection, a dialog will be shown requiring the user to download and update to this new version. After download the installation process will be started which guides the user through the required steps. Administrator privileges are required during the installation of the new software version. | 2023-11-29 11:58
STERIS | Hookup Accessories | - | - | PLEASE FIND ATTACHED FSN | PLEASE FIND ATTACHED FSN | 2023-11-29 12:58
Philips Medical Systems | CareEvent | 2 | - | It has been discovered that background notifications for the Care Assist application fail to send from PIC iX Event Notification and CareEvent under certain conditions This failure is applicable to Apple iOS devices only x D Specifically when a user places the Care Assist application into the background on an iOS device by opening a different application or locking the device etc the user is registered as being logged off from the device Under this condition push notifications are no longer sent to the user If the Care Assist application is in the foreground i e actively open on the user device the events are correctly sent to the connected iOS device x D | A Philips representative will contact you to schedule a visit to install a software patch to your device which will correct the current issue. Patch PIC iX 4.2.2 will be provided for PIC iX 4.X and patch CareEvent C.03.07 will be provided for CareEvent C.03.X. | 2023-11-30 10:53
Siemens Shanghai Medical Equipment Ltd. | Luminos Fusion FD VE10 | 19 | - | The user mush check that the footrest is locked into position and attached firmly every time before allowing a patient to sit or stand on it x D The footrest must only be attached and removed when the table is horizontal | This advisory letter is distributed to all affected customers together with an addendum to the operator manual. The addendum contains additional details and illustrations on how to securely attach the footrest to further improve the understanding of handling the foot rest. | 2023-12-03 12:00
Siemens Shanghai Medical Equipment Ltd. | Luminos Fusion | 19 | - | The user mush check that the footrest is locked into position and attached firmly every time before allowing a patient to sit or stand on it x D The footrest must only be attached and removed when the table is horizontal | This advisory letter is distributed to all affected customers together with an addendum to the operator manual. The addendum contains additional details and illustrations on how to securely attach the footrest to further improve the understanding of handling the foot rest. | 2023-12-03 12:00
Siemens Shanghai Medical Equipment Ltd. | Luminos RF Classic | 19 | - | The user mush check that the footrest is locked into position and attached firmly every time before allowing a patient to sit or stand on it x D The footrest must only be attached and removed when the table is horizontal | This advisory letter is distributed to all affected customers together with an addendum to the operator manual. The addendum contains additional details and illustrations on how to securely attach the footrest to further improve the understanding of handling the foot rest. | 2023-12-03 12:00
Siemens Shanghai Medical Equipment Ltd. | LUMINOS Impulse | 19 | - | The user mush check that the footrest is locked into position and attached firmly every time before allowing a patient to sit or stand on it x D The footrest must only be attached and removed when the table is horizontal | This advisory letter is distributed to all affected customers together with an addendum to the operator manual. The addendum contains additional details and illustrations on how to securely attach the footrest to further improve the understanding of handling the foot rest. | 2023-12-03 12:00
Siemens Shanghai Medical Equipment Ltd. | Luminos Fusion II VE10 | 19 | - | The user mush check that the footrest is locked into position and attached firmly every time before allowing a patient to sit or stand on it x D The footrest must only be attached and removed when the table is horizontal | This advisory letter is distributed to all affected customers together with an addendum to the operator manual. The addendum contains additional details and illustrations on how to securely attach the footrest to further improve the understanding of handling the foot rest. | 2023-12-03 12:00
SAUDI MAIS CO FOR MEDICAL PRODUCTS | 3-Way Stopcock with/witout Extension Tube | 50538 | - | leakage due to issue in male luer connector | recalling the product | 2023-12-04 09:11
Philips | BrightView XCT | 18 | - | While using Pre Programmed Motion PPM during an extrinsic quality assurance scan a gap is created between the patient support and the detector This gap presents a potential extremity entrapment hazard for patients while the system detectors and patient support are in motion x D Philips has received one report of an adverse event associated with this issue In this reported event the operator positioned a patient on the patient support and initiated the total body Pre Programmed Motion PPM While the patient support and detectors were in motion the patient straightened their leg due to a cramp causing their foot to extend and become entrapped between the patient support and the detector The patient sustained a foot fracture | Philips will contact you to schedule a time for a Philips Field Service Engineer (FSE) to visit your site and install a technical solution to resolve this issue (refer to FCO88200537). | 2023-12-04 14:44
Philips | BrightView | 18 | - | While using Pre Programmed Motion PPM during an extrinsic quality assurance scan a gap is created between the patient support and the detector This gap presents a potential extremity entrapment hazard for patients while the system detectors and patient support are in motion x D Philips has received one report of an adverse event associated with this issue In this reported event the operator positioned a patient on the patient support and initiated the total body Pre Programmed Motion PPM While the patient support and detectors were in motion the patient straightened their leg due to a cramp causing their foot to extend and become entrapped between the patient support and the detector The patient sustained a foot fracture | Philips will contact you to schedule a time for a Philips Field Service Engineer (FSE) to visit your site and install a technical solution to resolve this issue (refer to FCO88200537). | 2023-12-04 14:44
Philips | BrightView X | 18 | - | While using Pre Programmed Motion PPM during an extrinsic quality assurance scan a gap is created between the patient support and the detector This gap presents a potential extremity entrapment hazard for patients while the system detectors and patient support are in motion x D Philips has received one report of an adverse event associated with this issue In this reported event the operator positioned a patient on the patient support and initiated the total body Pre Programmed Motion PPM While the patient support and detectors were in motion the patient straightened their leg due to a cramp causing their foot to extend and become entrapped between the patient support and the detector The patient sustained a foot fracture | Philips will contact you to schedule a time for a Philips Field Service Engineer (FSE) to visit your site and install a technical solution to resolve this issue (refer to FCO88200537). | 2023-12-04 14:44
Ortho Clinical Diagnostics | ORTHO VISION Analyzer | 126 | - | Root Cause x D The wash protocol for ORTHO Sera Anti Jkb on the ORTHO VISION VISION Max BioVue Analyzer does not sufficiently clean the internal surface of the probe to prevent false positive test results due to carryover of ORTHO Sera Anti Jkb An effective wash protocol has been identified and will be implemented in an upcoming modification x D At this time Ortho s investigation is ongoing and to date has not identified carryover issues related to any of the remaining ORTHO Sera specificities when tested on the ORTHO VISION VISION Max BioVue Analyzers | Until the wash protocol is modified. •The continued use of ORTHO Sera Anti-Jkb on the ORTHO VISION/VISION Max Analyzer must only be done using the directions provided in the Resolution Section of this letter. •Review the content of this communication with your Medical Director and retain this letter for your laboratory. •Complete the enclosed Confirmation of Receipt form no later than December 15, 2023. | 2023-12-05 11:39
Ortho Clinical Diagnostics | ORTHO VISION Max Analyzer" | 126 | - | Root Cause x D The wash protocol for ORTHO Sera Anti Jkb on the ORTHO VISION VISION Max BioVue Analyzer does not sufficiently clean the internal surface of the probe to prevent false positive test results due to carryover of ORTHO Sera Anti Jkb An effective wash protocol has been identified and will be implemented in an upcoming modification x D At this time Ortho s investigation is ongoing and to date has not identified carryover issues related to any of the remaining ORTHO Sera specificities when tested on the ORTHO VISION VISION Max BioVue Analyzers | Until the wash protocol is modified. •The continued use of ORTHO Sera Anti-Jkb on the ORTHO VISION/VISION Max Analyzer must only be done using the directions provided in the Resolution Section of this letter. •Review the content of this communication with your Medical Director and retain this letter for your laboratory. •Complete the enclosed Confirmation of Receipt form no later than December 15, 2023. | 2023-12-05 11:39
Ortho Clinical Diagnostics | VITROS Immunodiagnostic Products Anti-HBs Reagent Pack | - | - | customers using the lots of VITROS Anti HBs Reagent Pack Calibrators listed above may experience an increase in calibration failures or an increase in imprecision for the negative quality control and patient samples if a successful calibration is performed QuidelOrtho also received one customer complaint regarding a falsely elevated result which was generated using an affected lot | Immediately discontinue using, render unusable, and discard all affected lots of VITROS Anti-HBs Reagent Pack/Calibrator in your inventory. •Ensure your laboratory is following the guidance provided in the Instructions For Use and/or local regulatory requirements. •Complete the enclosed Confirmation of Receipt form no later than December, 15 2023. Upon receipt of your completed Confirmation of Receipt form, QuidelOrtho will provide credit or replacement of your discarded inventory. •Upon receipt of your replacement lot of VITROS Anti-HBs Reagent Pack and Calibrator, re-calibrate your VITROS System(s). •Save this notification with your User Documentation or post this notification by each VITROS ECi/ECiQ/3600/5600/XT 7600 System until this issue has been resolved. •Please forward this notification if the affected product was distributed outside of your facility. | 2023-12-05 12:05
Ortho Clinical Diagnostics | VITROS Immunodiagnostic Products Anti-HBs Calibrators | - | - | customers using the lots of VITROS Anti HBs Reagent Pack Calibrators listed above may experience an increase in calibration failures or an increase in imprecision for the negative quality control and patient samples if a successful calibration is performed QuidelOrtho also received one customer complaint regarding a falsely elevated result which was generated using an affected lot | Immediately discontinue using, render unusable, and discard all affected lots of VITROS Anti-HBs Reagent Pack/Calibrator in your inventory. •Ensure your laboratory is following the guidance provided in the Instructions For Use and/or local regulatory requirements. •Complete the enclosed Confirmation of Receipt form no later than December, 15 2023. Upon receipt of your completed Confirmation of Receipt form, QuidelOrtho will provide credit or replacement of your discarded inventory. •Upon receipt of your replacement lot of VITROS Anti-HBs Reagent Pack and Calibrator, re-calibrate your VITROS System(s). •Save this notification with your User Documentation or post this notification by each VITROS ECi/ECiQ/3600/5600/XT 7600 System until this issue has been resolved. •Please forward this notification if the affected product was distributed outside of your facility. | 2023-12-05 12:05
Cardiac Pacemaker Inc – A wholly owned subsidiary of Guidant Corporation – A wholly owned subsidiary of Boston Scientific Corporation | 3120 Programmer | - | - | Starting on January the clock parameter used by x D the Model software application on the Model ZOOM programmer will display x D certain diagnostic dates incorrectly | Notification Please keep this letter with all Model 3120 programmers at your center or distribute this letter to all HCPs who perform pacemaker checks using a Model 3120 ZOOM programmer. When using the Model 2892 software application on the Model 3120 ZOOM programmer, if a date includes a year in the 1990s, HCPs may: o Determine a precise date by interrogating an ALTRUA, INSIGNIA, or NEXUS pacemaker with a Model 3300 LATITUDE programmer; or o Determine an approximate date (within a two-day range) | 2023-12-05 16:52
Cardiac Pacemaker Inc – A wholly owned subsidiary of Guidant Corporation – A wholly owned subsidiary of Boston Scientific Corporation | INGENIO EL and CRT-P | - | - | potential x D for the approximately remaining INGENIO family of DR EL pacemakers and CRT Ps to x D exhibit a high battery impedance later in device life and initiate Safety Mode None of these affected x D devices remain available for implant | Notification for HCP as follow: 1- Individual patient evaluation. Identify patients who are at risk of harm due to Safety Mode’s nonprogrammable parameters. 2- Replacement:  If a device enters Safety Mode, perform emergent replacement for patients who are at risk of harm. For other patients, non-emergent replacement is recommended. When choosing a replacement interval, do not rely on previously reported battery time remaining estimates which do not account for Safety Mode’s increased outputs nor the battery’s high impedance state.  General prophylactic replacement is not recommended. For patients who are at risk of harm, device replacement is recommended as follows: – For DR EL pacemakers, schedule replacement when the longevity remaining is 4 years or less. – For CRT-Ps, schedule replacement when the longevity remaining is 3 years or less. 3- Follow-up interval: Perform system follow-up in accordance with the instructions for use:  Perform a system follow-up via remote or in-office interrogation at least every 12 months; and  When longevity reaches One-Year-Remaining, follow-up every three (3) months thereafter until replacement is indicated. 4- Medical records. For each patient with an affected device, append their medical record with this letter to maintain awareness of this topic for the remaining service life of the device | 2023-12-05 17:15
werfen | ACL TOP® Family 50 Series | 3 | - | A field corrective action was initiated by Werfen Regulatory Department in July Under this corrective action customers were provided a letter that included instructions on how to avoid a rare potential sample misidentification A letter informing users how to avoid this issue was also added to start up kits for all new instruments beginning in August There is a subset of customers that may have received a new instrument between the time of the field corrective action and the inclusion of the letter into the start up kits | Communicate to all accounts that have received a unit with a serial number in this range above and provide them with the attached letter. If they have already been contacted by Werfen Regulatory regarding the Field Corrective Action, then they must follow the instructions previously provided. The attached letter DOES NOT replace the instructions under the Field Corrective Action. | 2023-12-06 11:49
Becton, Dickinson and Company, BD Biosciences | BD Calibrite™ 3-color kit | - | - | BD internally identified the affected product above was errantly shipped to countries requiring the CE mark representing compliance with the EU In Vitro Diagnostic Medical Device Regulation There is no issue with the product composition or performance BD no longer markets these products as CE marked devices and has updated the product to remove the CE mark x D x D | Actions taken by BD BD has implemented appropriate measures to prevent recurrence of this issue. Customer Actions: • Cease use of any unused affected BD Calibrite™ 3-Color Kits and BD® Stem Cell Enumeration Kits. • Identify and quarantine all unused affected BD Calibrite™ 3-Color Kits and BD® Stem Cell Enumeration Kits. • Make a note of the lot numbers and destroy all unused affected units. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 15th January 2024. • Circulate this notice to all those who need to be aware within your organization or to any organization where the potentially affected products have been transferred. • If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: • Cease distribution. • Identify, quarantine, making a note of the lot numbers then destroy all unused affected BD Calibrite™ 3-Color Kit and BD® Stem Cell Enumeration Kit. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 15th January 2024. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. | 2023-12-06 14:45
Sentinel.CH.SpA | CALiaGold® Calibrator | 32 | - | Based on the data collected from some customer sites after the calibration the Quality Control values may be found out of range Internal studies confirmed that using the above mentioned lots of CALiaGold Calibrator REF an inaccuracy of control material has been observed | 1. In case of remaining stock of the affected lots of CALiaGold® Calibrator REF 1151100, stop using the products. 2. Dispose the affected lots in accordance with your national and local safety and environmental regulations. 3. Fill in the “Field Safety Notice Receipt” and return it to Your local representative by January 20, 2023. 4. Review the content of this communication with your Medical Director and retain this letter for any future reference. | 2023-12-07 12:30
DiaSorin Inc. | LIAISON® 25 OH Vitamin D TOTAL Assay, LIAISON® 25 OH Vitamin D TOTAL Assay 200 | 2767 | - | DiaSorin has determined that the bias between plasma and serum samples is greater than stated in the Instructions for Use IFU for the LIAISON OH Vitamin D TOTAL Assays REF and REF The IFU includes a statement in Section Specimen Collection and Preparation that Human plasma collected in EDTA or lithium heparin tube types reported a mean bias of versus serum | •Do not test specimens collected in EDTA-plasma or lithium heparin plasma tube types using the LIAISON® 25 OH Vitamin D TOTAL kits. •Serum specimens can continue to be used to test and report patient results. •Please share this information with your Laboratory Medical Director and anyone in your facility that may be impacted by this notice. •Due to the low health risk, retesting of previously reported results from samples collected in EDTA-plasma or lithium heparin plasma tube types is not recommended. | 2023-12-07 18:44
KARL STORZ SE & Co. KG | bone surgery instruments | 132 | - | It was found that there is not insufficient evidence to show that the reprocessing method of the x D products was adequately validated This issue affects all lot numbers of the referenced KARL x D STORZ article numbers | Immediately quarantine and discontinue use of associated part numbers listed. 2. Pass on this urgent field safety notice to all users of the products listed above and all other persons who need to be aware within your organization. 3. If you have or may have distributed the devices listed, please identify and promptly notify those recipients, or provide KARL STORZ a list of customers who received/may have received the products listed. 4. Return the filled feedback form by Fax or E-Mail to the indicated contact within 15 calendar days from the date of receipt. 5. Get in touch with your KARL STORZ representative to return affected products. 6. Please report any incidents related to this issue to the manufacturer, dealer or local representative and, if applicable, to the national competent authority, as this is important feedback. | 2023-12-10 09:16
BD Switzerland Sàrl | Alaris™ GP Guardrails™ Plus | 161 | - | BD has identified that Performance Verification Procedures PVP were not completed by BD technicians for the specific models and serial numbers of BD Alaris System with Guardrails Suite MX listed in Appendix A | Corrective actions by BD BD is investigating this issue and corrective actions will be implemented to prevent recurrence. Actions to be taken by BD BD will be sending a technician to your site to verify the necessary checks and complete the documentation. Actions to be taken by Clinical Users Continue to use your BD Alaris™ System with Guardrails™ Suite MX devices in line with the applicable Directions for Use. Actions to be taken by Biomedical Departments 1. Circulate this Field Safety Notice to all those within your organisation that may use the BD Alaris™ System with Guardrails™ Suite MX devices listed in Appendix A. 2. If you have further distributed the devices, please identify those users and notify them at once of this Field Safety Notice. 3. Locate the serial numbers listed in Attachment A and conduct the necessary preparations to make them available for a site visit by the BD technician. 4. Complete the Customer Response Form on page 3 and return it to dl-ksa-inf-tech-support@bd.com as soon as possible or no later than the 30th of December, 2023. o A BD technician will be in contact to schedule a site visit. | 2023-12-10 15:36
Renishaw Mayfield Sarl | neuromate system | 1 | - | During a review of IFU it was identified that for the neuroinspire neurolocate optional module in the software some images x D used within the documents are incorrect x D The text is correct in the IFUs but the images do not make sense with the text The images should provide the user an x D explanation of the buttons to use and the screens they should expect to see in the software The incorrect images do not make x D sense in relation to the text and they are obviously incorrect | A review should be immediate to identify incorrect IFUs | 2023-12-10 17:29
.Siemens Healthcare Diagnostics Inc | Atellica IM FBHCG | - | - | Reason for Correction x D The purpose of this communication is to inform you of an issue with the product indicated in Table above and provide instructions on actions that your laboratory must take x D Siemens Healthcare Diagnostics Inc has confirmed through internal investigation that the Free Beta Human Chorionic Gonadotropin FBHCG assay does not meet the Limit of Quantitation LOQ Limit of Blank LoB and Limit of Detection LoD as stated in the Instructions for Use IFU for the Atellica IM Analyzers x D Siemens has defined an interim LOQ for the FBHCG assays for customers to utilize for the lots listed in Table and future lots Furthermore future lots will be aligned to these interim values Refer to Table for the interim LOQ values x D Since the LoB and LoD fall below the assay measuring interval as defined by the LOQ the establishment of an interim LoB and LoD was not conducted nor indicated | Actions to be Taken by the Customer Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. Create a linearity range utilizing the interim LOQ values provided in Table 2 with instructions available in the Atellica Solution Operator’s Guide and/or Online Help. All Atellica IM FBHCG assay kit lots can continue to be utilized with the interim LOQ until further notice. Results below the newly defined linearity range are flagged as ‘Below Linearity’ and may be reported as less than the interim LOQ. Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. If you have received any complaints of illness or adverse events associated with the products listed in Table 1, immediately contact your local Siemens Healthineers Customer Care Center or your local Siemens Healthineers technical support representative. Please retain this letter with your laboratory records and forward this letter to those who may have received this product. | 2023-12-10 17:43
BIOMERIEUX S.A. | VIDAS CK-MB | -10 | - | Based on complaints for invalid calibration due to S value out of range high reported from the field with VIDAS CK MB Ref on lot bioM rieux initiated an investigation INV to confirm the product issue and identify the root cause x D According to the product Instruction For Use IFU Calibration using the calibrator provided in the kit must be performed each time a new lot of reagents is opened after the master lot data have been entered Calibration should then be performed every days The calibrator identified by S must be tested in duplicate see User s Manual The calibrator value must be within the set RFV Relative Fluorescence Value If this is not the case recalibrate x D Therefore an invalid calibration is not considered as a product malfunction x D To be noted that when the calibration is invalid a message is displayed on the VIDAS report showing that the calibration is invalid and the customer should not release patient results if they are preceded by an invalid calibration Therefore there is no impact on patient sample results and no risk of false results The only risk related to an invalid calibration is a potential delayed result | Considering the overall risk MINOR for delayed result and since a risk of back-order has been ruled out, the Field Action Board determined that a Field Safety Corrective Action FSCA #5801 is appropriate to inform customers about the issue and instruct them to stop using and discard the impacted lot 1009900940 of VIDAS® CK-MB (Ref. 30421) remaining in their inventory. The FSCA notification with the associated Urgent Field Safety Notice (customer letter) will be sent to the impacted subsidiaries and distributors. - At subs/distributors levels: o Destroy impacted lot 1009900940 already under PSS #5794 remaining in their inventory. o Translate and send the customer letter (urgent field safety notice) to customers having received the impacted lot 1009900940. - At customers level: o Stop using and destroy the impacted lot 1009900940. | 2023-12-10 19:00
B.Braun Melsungen AG | Combitrans Monitoring Set | - | - | Combitrans Combidyn Monitoring Sets are invasive disposable monitoring systems for measuring physiological blood pressure and blood sampling They are supplied sterile and for single patient use x D As users already know prolonged contact with organic solvents spray disinfectants can promote the formation of stress cracks in plastic tubing connectors and multidirectional stopcocks x D A corresponding warning for the Discofix multidirectional stopcocks integrated in the affected sets was already distributed by the manufacturer see attached document Discofix A IFU However it is not yet distributed with the sets as we were only recently informed about it by the manufacturer We would therefore like to inform you that in the future the affected Combitrans Combidyn monitoring sets will include also a separate warning notice for the Discofix multidirectional stopcocks x D | 1. Please identify all customers in your country who are currently using the affected products. 2. If you have active customers: a. Please finalize the attached Field Safety Notice and Appendices (provided additionally as MS Word file in English language) with your local contact information and signature. b. Please inform the identified customers with the Field Safety Notice including the Appendices. c. Please make sure that the receipt of the Field Safety Notice is confirmed by each customer on Appendix 3 of the FSN. d. If you require further information on correct product use or training, please contact your responsible B. Braun marketing/sales representative. e. Please report this Field Safety Corrective Action to your local authority, if required. Attached you can find a reporting template (used in European Union). Please add the name of your authority in section 1 and your local contact details in section 5. | 2023-12-11 10:34
B.Braun Melsungen AG | Exadyn Combitrans | - | - | Combitrans Combidyn Monitoring Sets are invasive disposable monitoring systems for measuring physiological blood pressure and blood sampling They are supplied sterile and for single patient use x D As users already know prolonged contact with organic solvents spray disinfectants can promote the formation of stress cracks in plastic tubing connectors and multidirectional stopcocks x D A corresponding warning for the Discofix multidirectional stopcocks integrated in the affected sets was already distributed by the manufacturer see attached document Discofix A IFU However it is not yet distributed with the sets as we were only recently informed about it by the manufacturer We would therefore like to inform you that in the future the affected Combitrans Combidyn monitoring sets will include also a separate warning notice for the Discofix multidirectional stopcocks x D | 1. Please identify all customers in your country who are currently using the affected products. 2. If you have active customers: a. Please finalize the attached Field Safety Notice and Appendices (provided additionally as MS Word file in English language) with your local contact information and signature. b. Please inform the identified customers with the Field Safety Notice including the Appendices. c. Please make sure that the receipt of the Field Safety Notice is confirmed by each customer on Appendix 3 of the FSN. d. If you require further information on correct product use or training, please contact your responsible B. Braun marketing/sales representative. e. Please report this Field Safety Corrective Action to your local authority, if required. Attached you can find a reporting template (used in European Union). Please add the name of your authority in section 1 and your local contact details in section 5. | 2023-12-11 10:34
B.Braun Melsungen AG | Exadyn Monitoring Set | - | - | Combitrans Combidyn Monitoring Sets are invasive disposable monitoring systems for measuring physiological blood pressure and blood sampling They are supplied sterile and for single patient use x D As users already know prolonged contact with organic solvents spray disinfectants can promote the formation of stress cracks in plastic tubing connectors and multidirectional stopcocks x D A corresponding warning for the Discofix multidirectional stopcocks integrated in the affected sets was already distributed by the manufacturer see attached document Discofix A IFU However it is not yet distributed with the sets as we were only recently informed about it by the manufacturer We would therefore like to inform you that in the future the affected Combitrans Combidyn monitoring sets will include also a separate warning notice for the Discofix multidirectional stopcocks x D | 1. Please identify all customers in your country who are currently using the affected products. 2. If you have active customers: a. Please finalize the attached Field Safety Notice and Appendices (provided additionally as MS Word file in English language) with your local contact information and signature. b. Please inform the identified customers with the Field Safety Notice including the Appendices. c. Please make sure that the receipt of the Field Safety Notice is confirmed by each customer on Appendix 3 of the FSN. d. If you require further information on correct product use or training, please contact your responsible B. Braun marketing/sales representative. e. Please report this Field Safety Corrective Action to your local authority, if required. Attached you can find a reporting template (used in European Union). Please add the name of your authority in section 1 and your local contact details in section 5. | 2023-12-11 10:34
B.Braun Melsungen AG | Haemofix-Combitrans Monitoring Kit | - | - | Combitrans Combidyn Monitoring Sets are invasive disposable monitoring systems for measuring physiological blood pressure and blood sampling They are supplied sterile and for single patient use x D As users already know prolonged contact with organic solvents spray disinfectants can promote the formation of stress cracks in plastic tubing connectors and multidirectional stopcocks x D A corresponding warning for the Discofix multidirectional stopcocks integrated in the affected sets was already distributed by the manufacturer see attached document Discofix A IFU However it is not yet distributed with the sets as we were only recently informed about it by the manufacturer We would therefore like to inform you that in the future the affected Combitrans Combidyn monitoring sets will include also a separate warning notice for the Discofix multidirectional stopcocks x D | 1. Please identify all customers in your country who are currently using the affected products. 2. If you have active customers: a. Please finalize the attached Field Safety Notice and Appendices (provided additionally as MS Word file in English language) with your local contact information and signature. b. Please inform the identified customers with the Field Safety Notice including the Appendices. c. Please make sure that the receipt of the Field Safety Notice is confirmed by each customer on Appendix 3 of the FSN. d. If you require further information on correct product use or training, please contact your responsible B. Braun marketing/sales representative. e. Please report this Field Safety Corrective Action to your local authority, if required. Attached you can find a reporting template (used in European Union). Please add the name of your authority in section 1 and your local contact details in section 5. | 2023-12-11 10:34
B.Braun Melsungen AG | Combitrans Monitoring EC | - | - | Combitrans Combidyn Monitoring Sets are invasive disposable monitoring systems for measuring physiological blood pressure and blood sampling They are supplied sterile and for single patient use x D As users already know prolonged contact with organic solvents spray disinfectants can promote the formation of stress cracks in plastic tubing connectors and multidirectional stopcocks x D A corresponding warning for the Discofix multidirectional stopcocks integrated in the affected sets was already distributed by the manufacturer see attached document Discofix A IFU However it is not yet distributed with the sets as we were only recently informed about it by the manufacturer We would therefore like to inform you that in the future the affected Combitrans Combidyn monitoring sets will include also a separate warning notice for the Discofix multidirectional stopcocks x D | 1. Please identify all customers in your country who are currently using the affected products. 2. If you have active customers: a. Please finalize the attached Field Safety Notice and Appendices (provided additionally as MS Word file in English language) with your local contact information and signature. b. Please inform the identified customers with the Field Safety Notice including the Appendices. c. Please make sure that the receipt of the Field Safety Notice is confirmed by each customer on Appendix 3 of the FSN. d. If you require further information on correct product use or training, please contact your responsible B. Braun marketing/sales representative. e. Please report this Field Safety Corrective Action to your local authority, if required. Attached you can find a reporting template (used in European Union). Please add the name of your authority in section 1 and your local contact details in section 5. | 2023-12-11 10:34
B.Braun Melsungen AG | Combitrans EC | - | - | Combitrans Combidyn Monitoring Sets are invasive disposable monitoring systems for measuring physiological blood pressure and blood sampling They are supplied sterile and for single patient use x D As users already know prolonged contact with organic solvents spray disinfectants can promote the formation of stress cracks in plastic tubing connectors and multidirectional stopcocks x D A corresponding warning for the Discofix multidirectional stopcocks integrated in the affected sets was already distributed by the manufacturer see attached document Discofix A IFU However it is not yet distributed with the sets as we were only recently informed about it by the manufacturer We would therefore like to inform you that in the future the affected Combitrans Combidyn monitoring sets will include also a separate warning notice for the Discofix multidirectional stopcocks x D | 1. Please identify all customers in your country who are currently using the affected products. 2. If you have active customers: a. Please finalize the attached Field Safety Notice and Appendices (provided additionally as MS Word file in English language) with your local contact information and signature. b. Please inform the identified customers with the Field Safety Notice including the Appendices. c. Please make sure that the receipt of the Field Safety Notice is confirmed by each customer on Appendix 3 of the FSN. d. If you require further information on correct product use or training, please contact your responsible B. Braun marketing/sales representative. e. Please report this Field Safety Corrective Action to your local authority, if required. Attached you can find a reporting template (used in European Union). Please add the name of your authority in section 1 and your local contact details in section 5. | 2023-12-11 10:34
Philips Medical Systems Nederland B.V. | AZURION 3 M12 | 74 | - | Philips has identified a potential issue which can cause the system to continuously restart restart loop This is likely to occur when the patient database gets too big over time studies If this issue occurs a potential for data loss may be expected x D | Philips is working on a software release that will correct this issue (reference: FCO72200525, FCO72200528, FCO72200535, and FCO72200548). | 2023-12-11 10:49
Philips Medical Systems Nederland B.V. | AZURION 3 M15 | 74 | - | Philips has identified a potential issue which can cause the system to continuously restart restart loop This is likely to occur when the patient database gets too big over time studies If this issue occurs a potential for data loss may be expected x D | Philips is working on a software release that will correct this issue (reference: FCO72200525, FCO72200528, FCO72200535, and FCO72200548). | 2023-12-11 10:49
Philips Medical Systems Nederland B.V. | Azurion 5 M20 | 74 | - | Philips has identified a potential issue which can cause the system to continuously restart restart loop This is likely to occur when the patient database gets too big over time studies If this issue occurs a potential for data loss may be expected x D | Philips is working on a software release that will correct this issue (reference: FCO72200525, FCO72200528, FCO72200535, and FCO72200548). | 2023-12-11 10:49
Philips Medical Systems Nederland B.V. | AZURION 7 B12 | 74 | - | Philips has identified a potential issue which can cause the system to continuously restart restart loop This is likely to occur when the patient database gets too big over time studies If this issue occurs a potential for data loss may be expected x D | Philips is working on a software release that will correct this issue (reference: FCO72200525, FCO72200528, FCO72200535, and FCO72200548). | 2023-12-11 10:49
Philips Medical Systems Nederland B.V. | AZURION 7 B20 | 74 | - | Philips has identified a potential issue which can cause the system to continuously restart restart loop This is likely to occur when the patient database gets too big over time studies If this issue occurs a potential for data loss may be expected x D | Philips is working on a software release that will correct this issue (reference: FCO72200525, FCO72200528, FCO72200535, and FCO72200548). | 2023-12-11 10:49
Philips Medical Systems Nederland B.V. | AZURION 7 M12 | 74 | - | Philips has identified a potential issue which can cause the system to continuously restart restart loop This is likely to occur when the patient database gets too big over time studies If this issue occurs a potential for data loss may be expected x D | Philips is working on a software release that will correct this issue (reference: FCO72200525, FCO72200528, FCO72200535, and FCO72200548). | 2023-12-11 10:49
Philips Medical Systems Nederland B.V. | AZURION 7 M20 | 74 | - | Philips has identified a potential issue which can cause the system to continuously restart restart loop This is likely to occur when the patient database gets too big over time studies If this issue occurs a potential for data loss may be expected x D | Philips is working on a software release that will correct this issue (reference: FCO72200525, FCO72200528, FCO72200535, and FCO72200548). | 2023-12-11 10:49
GME German Medical Engineering | Flexsys | 6 | حسب توجيهات المصنع | - | The updated User Manual was shared with customers, and they were informed of the updates. | 2023-12-11 14:19
Siemens Healthcare GmbH | SENSIS Vibe Combo | 1 | - | If the sensis documentation functionality is used it may happen that during adding once per study reporting events Type event as further defined in the administrator manual that the application crashes | The software will be updated to correct the issue. | 2023-12-11 15:07
.Siemens Healthcare Diagnostics Inc | IRON | - | - | Reason for Customer Notification x D The purpose of this communication is to inform you of an update to the product listed in Table above and provide instructions on actions that your laboratory must take x D Siemens Healthineers internal testing has determined that the IRON assay for use on the Dimension Clinical Chemistry System has lipemia interference claims in the Instructions For Use IFU that require updates as shown in Table While Siemens has not received any customer complaints for this issue grossly lipemic samples have the potential to falsely elevate iron results on the Dimension Clinical Chemistry System x D Updated versions of the Dimension IRON Assay Instructions For Use will be available on the Siemens Document Library | Actions to be Taken by the Customer Please perform the following steps: • Download the updated Dimension IRON IFU revision 09 SMN 10865227 from Siemens Document Library. NOTE: If your laboratory does not have the HIL feature [Hemolysis (H), Icterus (I), Lipemia/Turbidity (L)] activated on your Dimension system, no action is required. • Please perform the following steps if the HIL feature is activated on your Dimension Clinical Chemistry System. o Sign into the Dimension System Software with a User Account that has privileges to edit HIL settings. o From the main operating screen Select F5: PROCESS CTRL > F8: MORE OPTIONS. o Select F6: HIL SETUP. o Press the IRON key or position the cursor on the row for the IRON method and press Enter. o Update the L (Lipemia/Turbidity) from 0 (no interference) to 4. o Select F2: STORE. o Please refer to your Dimension Systems Operator’s Guide for further information and reference. Note: Each laboratory has the option of entering HIL indices based on the laboratory policies with considerations on acceptable medical decision errors, impact on clinical interpretations and ability to mitigate preanalytical specimen interfering substances. Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. If you have received any complaints of illness or adverse events associated with the product listed in Table 1, immediately contact your local Siemens Healthineers Remote Services Center or your local Siemens Healthineers Technical Support Representative. Please retain this letter with your laboratory records and forward this letter to those who may have received this product. | 2023-12-11 17:02
Randox | Calibration Serum Level 3 | - | - | There have been transcription errors on the IFU for the Calibration Serum Level CAL Additionally Standardisation is off by to the reference material for CK NAC x D x D The target for ALP for the AMP optimised to IFCC oC method has been mis assigned for the RX Series instruments in Calibration Serum Level CAL lot UE by approximately | Notifying the customers by email and implementing IFU or labelling change | 2023-12-12 10:31
Randox | Assayed Urine Control Level 2 | - | - | Randox can confirm that there has been a transcription error for the target and range values for Creatinine for all methods in the IFU | The IFU has been updated and available on www.randox.com please discard the incorrect IFU and download the updated IFU | 2023-12-12 12:49
Abbott | Proclaim™ XR SCS and Proclaim™ Elite SCS System | - | - | During standard use a Bluetooth x D x D paired Patient Controller is used to place a patient s implantable pulse generator in MRI x D mode which in turn disables the delivery of therapy Upon completion of the MRI the paired Patient Controller is then x D used to exit MRI mode and allow therapy to resume MRI mode is intended to protect the patient and the device during an x D MRI scan Before any MRI the Patient Controller PC should be used to enter MRI mode After the MRI the Patient x D Controller should be used to exit MRI mode x D | Following the guidance currently available in the Instructions for Use and this communication, if your patients need to receive an MRI scan: • Advise patients to contact the physician managing their implant prior to receiving an MRI for any reason • Ensure patients have upgraded their Patient Controller to the latest version of the application from the Apple‡ App Store‡ before attempting to place their device into MRI mode. • Advise patients to disable automatic updates for the Patient Controller app and iOS‡ software before placing the IPG in MRI mode. • Advise patients to set the IPG to MRI mode within a reasonable timeframe before the MRI procedure: o For SCS/DRG, this should be done within a day prior to arrival for the MRI procedure. o For DBS, this should be done no more than 24 hours in advance, and as close to the MRI procedure as possible, based on the patient’s ability to tolerate being off therapy. • Advise patients to not update, install, or delete the Patient Controller app or Bluetooth® connection while in MRI mode. • Advise patients not to alter, damage, update, or lose their Patient Controller handheld while in MRI mode. • After the MRI scan, use the Patient Controller to exit MRI mode and restore therapy. Patients may require the assistance of your Abbott representative or your medical team to do so. | 2023-12-12 13:22
Getinge | MAQUET CARDIOSAVE Hybrid and MAQUET CARDIOSAVE Rescue | - | - | Datascope Getinge has received complaints reporting unexpected shutdown of the Cardiosave x D IABP while providing therapy in very rare instances An internal investigation of the complaints x D determined an unexpected shutdown may be due to blood entering into the Cardiosave IABP x D when therapy is provided with a perforated intra aortic balloon catheter | A review of our records indicates that you may have a Cardiosave Hybrid and/or Cardiosave Rescue Intra-Aortic Balloon Pump (IABP) in your facility. • Please examine your inventory immediately to determine if you have any Cardiosave Hybrid and/or Rescue IABPs. • If a blood back event is suspected, clinicians are to remove the Cardiosave from patient use, and report the event to appropriate biomedical engineering staff for inspection of the safety disk prior to next patient use. • Please ensure that all Cardiosave Intra-Aortic Balloon Pump users at your facility are aware of this notice and actions to perform. o Distribute the above clinical guidance to users per your institution policy. | 2023-12-12 13:30
Olympus (Gyrus ACMI, Inc) | Single-use Laser Fibers | 239 | - | Olympus conducted an internal review of the SOLTIVE Single use Fibers and identified inconsistencies in the IFU Olympus is providing you with the attached labeling addendum which identifies the errors in the current instructions and the corrections that will be implemented in the IFU | A. Cleaning the Fiber Tip: Instructions have been updated to remove hydrogen peroxide because it was not validated, as follows: Prior to cleaning the tip, the laser must be placed in the STANDBY mode. During the treatment the laser fiber can be cleaned with a sterile pad immersed in hydrogen peroxide or sterile water. B. Fiber Cleaving and Stripping Instructions: Instructions have been clarified and replaced with the following: 1. Reference Table 3 to select the appropriate fiber stripper for the fiber size. 2. Strip the fiber so this action results in approximately 20 mm of exposed fiber per the fiber stripper Instructions for Use. 3. Holding the fiber cleaver, per the fiber cleaver Instructions for Use, score the fiber so the final exposed glass tip is 3-5mm, as recommended for each fiber size in Table 3. 4. Confirm proper stripping cleaving of the fiber by checking the aiming beam as described in Figure 1A. 5. Repeat the cleaving procedure from step 2 until a satisfactory aiming beam pattern is achieved. C. Accessories:Table 3 has been corrected to reflect 200 μm core diameter instead of 240 μm. | 2023-12-12 16:54
Olympus Medical Systems Corp. | ENDOEYE FLEX DEFLECTABLE VIDEOSCOPE | 2 | - | During an evaluation of user implementation of the Instructions for Use related to the reprocessing of the LTF S videoscope Olympus observed deviations from the following reprocessing steps detailed in the respective Reprocessing Manual Instructions These steps are required for the proper reprocessing of these videoscopes Therefore Olympus is reminding users that following the instructions in these reprocessing manuals LTF S Reprocessing Manual and LTF S Reprocessing Manual is required Olympus is offering educational on site support to review and train on the reprocessing instructions for the LTF S and LTF S The LTF S is similar in design to the LTF S and is therefore included in this corrective action | 1. Carefully read the content of this Field Safety Notice (FSN). 2. Please make sure that all medical personnel are completely knowledgeable and thoroughly trained on the reprocessing instructions contained in Sections 5.5, 5.6, and 5.7 of the LTF Reprocessing Manuals, especially adhering to the immersion time for the entire endoscope, and the brushing or wiping steps. 3. Indicate on the Reply Form that you have received and understood this Field Safety Notice by filling out and returning the completed enclosed Reply Form back to your local Olympus representative latest by 19.12.2023. 4. If you have distributed these devices outside your facility, please notify your customers of this matter immediately by forwarding them this Field Safety Notice. Please appropriately document your notification process and let us know the end-customer feedback accordingly | 2023-12-13 10:01
Abbott Ireland Diagnostics Division | Alinity i HBsAg Reagent Kit | 5 | - | Please check attached FSN letter | Please check attached FSN letter. | 2023-12-14 15:08
Bio-Rad | Monolisa HCV Ag-Ab Ultra V2 | - | - | few additional days of delay for HCV antigen positivity are observed | customer notification | 2023-12-17 10:45
werfen | HemosIL® AcuStar ADAMTS13 Activity | 42 | - | Customer complaints were reported regarding patient sample results below the normal range with a subset activity and erroneously low results with HemosIL AcuStar ADAMTS Activity In some cases these patient samples were later tested with FRET or ELISA assays with results reporting higher percent activity than the HemosIL AcuStar ADAMTS Activity assay | To address the erroneous low results, the manufacturer is releasing a mandatory service installable Assay Definition (AD) V9.0 which includes flags to identify issues related to sample mis-aspiration (such as films, bubbles, and clots). AD V9.0 is a mandatory service-installable upgrade for ACL AcuStar/BIO-FLASH systems with Windows 7 and software version 3.0.2 or higher that will be performed as mandatory upgrade. | 2023-12-17 14:29
werfen | HemosIL® AcuStar ADAMTS13 Activity | 42 | - | Customer complaints were reported regarding patient sample results below the normal range with a subset activity and erroneously low results with HemosIL AcuStar ADAMTS Activity In some cases these patient samples were later tested with FRET or ELISA assays with results reporting higher percent activity than the HemosIL AcuStar ADAMTS Activity assay | To address the erroneous low results, the manufacturer is releasing a mandatory service installable Assay Definition (AD) V9.0 which includes flags to identify issues related to sample mis-aspiration (such as films, bubbles, and clots). AD V9.0 is a mandatory service-installable upgrade for ACL AcuStar/BIO-FLASH systems with Windows 7 and software version 3.0.2 or higher that will be performed as mandatory upgrade. | 2023-12-17 14:36
Vygon | POLYFILM adhesive dressing | 200 | - | Some units of these batches may have been manufactured with a potential defective sealing | Recall product | 2023-12-17 15:55
Wallac OY | NeoBase 2 Non-derivatized MSMS kit | - | - | The purpose of the letter is to inform you that Revvity Wallac Oy is voluntarily initiating a field safety corrective action of NeoBase Non derivatized MSMS kit x D Reason for the Voluntary field safety corrective action x D There is a minor inadequacy in the instructions for use for product NeoBase Internal Standards IS vial reconstitution and further Extraction Working Solution EWS preparation steps contain minor unclarity Note that there are no changes in the in use stability of EWS x D The extended storage of reconstituted NeoBase Internal Standards vial at C may possibly cause some loss of C LPC IS and or C IS from solution due to their observed irreversible binding to the inner vial surface Therefore these new clarified IS vial handling instruction should be followed | Actions to be taken by the customer: • The use of the NeoBase 2 Non-derivatized MSMS kit can continue. However, please note that the instructed single Extraction Working Solution (EWS) should be prepared immediately after reconstitution step of NeoBase 2 Internal Standards vial. Do not store reconstituted Internal Standards vial. After the EWS preparation (see: Reagent Preparation, step 1), it can be used for 2 weeks when stored at +2–+8 °C protected from light. When not in use, store the EWS protected from light in a glass laboratory bottle with firmly sealed screw cap closure at +2–+8 °C. • Complete the Response Form and return it to Wallac Oy / Revvity. Please contact your local Revvity representative for further information. Actions to be taken by Wallac Oy / Revvity: • The Instructions for use will be updated and released once approved by applicable regulatory authorities. Other Information: Please inform those affected in your organization accordingly. Kindly review the atatched UFSN | 2023-12-18 11:22
Megadyne Medical Product, Inc. | MEGADYNE™ MEGA SOFT™ Universal and Universal Plus Reusable Patient Return Electrodes | 4 | - | Reason for the Voluntary Correction x D Megadyne Medical Products Inc Megadyne has received reports of patient burns identified after surgical procedures in which Mega Soft pads were used Megadyne is taking this corrective action to mitigate the potential risk to health in the population of children under years of age We have conducted a thorough investigation and have not identified any design or manufacturing defects nor have we determined definitive root cause for the reports x D We are initiating updates to the Instructions for Use IFU and product labeling to reflect that these product codes should not be used in patients under years old The IFU update will be made available electronically at www e ifu com Users should continue to follow the current Mega Soft Instructions for Use IFU except for this new limitation in population of intended use We will notify customers if we identify any additional actions that may help to ensure safe use of the products x D | Actions Required – 1. Share this notification update with all users of Mega Soft Universal and Universal Plus pads. 2. Confirm that personnel using the Mega Soft Universal and Universal Plus pads understand the intended use is changing to patients aged 12 years and older. 3. Post a copy of this communication to remind staff not to use the Mega Soft Universal and Mega Soft Universal Plus pads on patients under 12 years old. Although the current Mega Soft Universal and Mega Soft Universal Plus have printing of > 0.35 kg (0.8 lbs) on the pads, they should only be used for patients aged 12 years and older, and over labeling of the pad is not required. 4. If any subject product has been forwarded to another facility, contact that facility to share this information. Please share a copy of this notification when communicating. 5. Complete the Business Reply Form (BRF) Attachment A confirming receipt of this notice and fax or email it to Evgeniya Skorotetskaya [eskorote@its.jnj.com] within three (3) business days. 6. As a reminder, it is important to follow proper cleaning, placement and setup steps for the Mega Soft pad. Failure to follow the Mega Soft pad IFU may contribute to patient burns. Copies of Cleaning and Care Visual Aid and Placement and Setup Visual Aid are available by contacting Evgeniya Skorotetskaya [eskorote@its.jnj.com]. 7. If you have any further questions related to this notice or if you need any additional communications, please contact Evgeniya Skorotetskaya [eskorote@its.jnj.com]. | 2023-12-18 12:47
Philips | BrightView | 17 | - | A detector support component may fail due to unexpected wear x D Scenario If the detector is below the center of the gantry bore and the detector support component fails the detector may descend downward unexpectedly and potentially contact the patient x D Scenario If the detector is above the center of the gantry bore and the detector support component fails the detector will not be able to move to complete the imaging x D Philips has received a complaint associated with this issue however there are no reports of injury or serious harm | Philips will contact you to schedule a time for a Philips Field Service Engineer (FSE) to visit your site and to inspect and correct the system if necessary (FCO 88200538). | 2023-12-18 13:35
Philips | BrightView XCT | 17 | - | A detector support component may fail due to unexpected wear x D Scenario If the detector is below the center of the gantry bore and the detector support component fails the detector may descend downward unexpectedly and potentially contact the patient x D Scenario If the detector is above the center of the gantry bore and the detector support component fails the detector will not be able to move to complete the imaging x D Philips has received a complaint associated with this issue however there are no reports of injury or serious harm | Philips will contact you to schedule a time for a Philips Field Service Engineer (FSE) to visit your site and to inspect and correct the system if necessary (FCO 88200538). | 2023-12-18 13:35
Philips | BrightView X | 17 | - | A detector support component may fail due to unexpected wear x D Scenario If the detector is below the center of the gantry bore and the detector support component fails the detector may descend downward unexpectedly and potentially contact the patient x D Scenario If the detector is above the center of the gantry bore and the detector support component fails the detector will not be able to move to complete the imaging x D Philips has received a complaint associated with this issue however there are no reports of injury or serious harm | Philips will contact you to schedule a time for a Philips Field Service Engineer (FSE) to visit your site and to inspect and correct the system if necessary (FCO 88200538). | 2023-12-18 13:35
Randox | RX Daytona Plus (with ISE) | - | - | The RX Daytona Plus Parameter settings have been updated in Parameter version v The rerun settings sample volume Low High and dilutions used for Auto reruns have been disabled Randox are in the process of generating and updating existing data to comply with In vitro Diagnostic Regulations IVDR For samples above the assay range a manual dilution should be performed and results multiplied by the dilution factor x D The RX Daytona plus operator manual has also been updated with Technical Bulletin RXTB detailing that C wash solution RX is pre diluted no additional dilutions are required Insert RXTB into the operator manual to ensure all within your organisation are aware of this update | Download and update parameters, or manually enter updates to avoid losing any previously obtained data, for assays performed on your Rx Daytona Plus. | 2023-12-18 16:05
Randox | RX Daytona Plus (without ISE) | - | - | The RX Daytona Plus Parameter settings have been updated in Parameter version v The rerun settings sample volume Low High and dilutions used for Auto reruns have been disabled Randox are in the process of generating and updating existing data to comply with In vitro Diagnostic Regulations IVDR For samples above the assay range a manual dilution should be performed and results multiplied by the dilution factor x D The RX Daytona plus operator manual has also been updated with Technical Bulletin RXTB detailing that C wash solution RX is pre diluted no additional dilutions are required Insert RXTB into the operator manual to ensure all within your organisation are aware of this update | Download and update parameters, or manually enter updates to avoid losing any previously obtained data, for assays performed on your Rx Daytona Plus. | 2023-12-18 16:05
Tandem Diabetes Care | t:slim X2, Control-IQ | 1 | - | A system check was performed and no issues were identified with the pump cartridge infusion set tubing infusion set cannula or the insulin in use at the time of event | The customer was hospitalized, and treated with insulin. Customer was released from the hospital on September 21st, 2023, with the issue resolved. Customer did not sustain any permanent damage. | 2023-12-19 11:27
Randox | Copper | - | - | Randox released an update to the Carryover Avoidance Technical Bulletin to detail that Copper should not be run directly after Total Protein on the RX series instruments testing order If reagent carryover persists avoid running Copper on analysers that run Total Protein Please discard all previous version of the Carryover Avoidance Technical Bulletin and replace with the updated version | - Review your instrument testing order in line with the Carryover Avoidance Technical Bulletin (RXTB-0148) - Update the RX user manual with the updated Carryover Avoidance Technical Bulletin (RXTB-0148) and ensure all operators are aware of the recommendations. - Discuss the contents of this notice with your Medical Director. - Complete and run the response form 12187-QA to technical.services@randox.com within five working days. | 2023-12-19 13:02
LifeGlobal Group, LLC | global® | 1 | - | CooperSurgical has become aware of a sudden increase in complaints regarding the aforementioned lots of this x D product While we do not know the cause of the performance concern due to the high volume of customer x D complaints for the three associated lots we wish to proactively address any possible issue with our products x D while we continue to investigate | Our records indicate that you may have purchased the affected product from CooperSurgical. Please take the following steps to ensure proper quarantine and safe return of the affected product(s) for additional testing and CooperSurgical will issue credit for the return. 1) Inspect your inventory, identify, and quarantine global® Media (Part Numbers: LGGG-100, LGGG-050, and LGGG-020, Lots: 231020-018743, 231020-018742, and 231020-018741) 2) If you are a Customer, complete page 3 of this communication, also labeled Customer Acknowledgement Form and return to Recall@coopersurgical.com or fax to +1 203.601.9870, ATTN: Recall. Be sure to document information clearly to prevent delays. 3) If you are a Distributor, complete page 4 of this communication, also labeled Distributor Acknowledgement Form and return to Recall@coopersurgical.com or fax to +1 203.601.9870, ATTN: Recall. Be sure to document information clearly to prevent delays. 4) As a regulatory requirement, even if you do not have any affected product in your inventory, please complete and return the form so that we may document confirmation and receipt of this Field Safety Notice. | 2023-12-19 14:29
Medtronic Navigation, Inc. | SOFTWARE STEALTHSTATION CRANIAL | 20 | - | Medtronic has identified an anomaly with the StealthStation System Cranial software Biopsy Depth Gauge graphical display which presents during navigation in a Cranial Biopsy Procedure When this anomaly presents the software enters a state where the Biopsy Depth Gauge is no longer synchronized with the navigational information displayed on the screen This causes the position of the biopsy needle to be displayed inaccurately This defect is present in the Navigation task of the Biopsy Procedure in Synergy Cranial Software Versions and StealthStation Cranial Software Versions and x D x D Update April x D Medtronic has identified a new software anomaly in StealthStation Cranial model version and therefore is ceasing installation of this version pending an updated software version Synergy Cranial model version is unaffected and installation of this software version will continue In addition to the above change in software version installation this notification is also intended to inform affected customers about an additional software anomaly impacting the StealthStation S and i systems running Cranial Software Version This anomaly can under specific conditions result in potential for inaccurate conflicting information to be displayed during Tumor Resection Shunt Placement and Nexframe DBS procedures x D x D Under certain situations the Distance to Target D T text that displays on the screen may not display accurately This can impact Tumor Resection Shunt Placement and Nexframe DBS procedures x D x D Medtronic will attach a warning and instructional placard about the software anomaly to the impacted systems until a software fix is available x D | 1. Please review this information with all physician users. If you have any questions related to this issue, please contact your Medtronic field representative 2. Please confirm via the enclosed Customer Acknowledgment Form that you understand Medtronic will be performing a software update on your impacted StealthStation™ system(s), , providing the updated IFU(s) upon software update completion, and removing the warning and instructional placard and that this notification has been communicated within your facility with all physician users. Send the completed Customer Acknowledgment Form to Medtronic 3. This notice needs to be passed on to all those who need to be aware of within your organization or to any organization where the potentially affected devices have been transferred. Please maintain a copy of this notice in your records. | 2023-12-19 17:56
Cordis Cashel | OptEase® Retrievable Vena Cava Filter and introduction kit | - | - | Overview x D This letter provides important information concerning the decision by Cordis to update the Warning statement in the Instructions for Use IFU for the Cordis OPTEASE Retrievable Vena Cava Filter and OPTEASE Retrieval Catheter x D This Field Safety Notice FSN is being conducted solely to inform end users of clarifications to the previously defined timeframe the OPTEASE Retrievable Vena Cava Filter device is implanted before retrieval x D Therefore Cordis has updated the Warning statement in accordance with regulatory requirements as follows x D The OPTEASE Retrievable Filter has not been studied for long term implantation and must be retrieved within days after placement x D Previous Warning statement x D The OPTEASE Retrievable filter can be retrieved up to and including days after placement The OPTEASE Retrievable filter is considered a permanent implant if it is not retrieved within the specified time period x D The statement below was removed from the device description for the OPTEASE Retrieval Catheter IFU in accordance with the updates made to the Warning statement in the IFU for the Cordis OPTEASE Retrievable Vena Cava Filter and OPTEASE Retrieval Catheter x D The OPTEASE Retrievable Filter can be retrieved within a specified period after implantation refer to the OPTEASE Retrievable Filter Instructions for Use or remain implanted as a permanent filter x D Please share this information with any of your staff involved in the use of this device | Actions requested on your part: 1.Read this Urgent Field Safety Notice. 2.Sign and return the enclosed Acknowledgement Form in accordance with the directionson the form. 3.Share this notification with anyone in your facility that needs to be informed. 4.Contact any other facilities that have been provided with units of the affected catalogcodes (all unexpired lots). 5.Retain a copy of this notice with the product. Kindly review the attached UFSN | 2023-12-20 10:57
Cordis Corporation | OPTEASE® Retrieval Catheter | - | - | Overview x D This letter provides important information concerning the decision by Cordis to update the Warning statement in the Instructions for Use IFU for the Cordis OPTEASE Retrievable Vena Cava Filter and OPTEASE Retrieval Catheter x D This Field Safety Notice FSN is being conducted solely to inform end users of clarifications to the previously defined timeframe the OPTEASE Retrievable Vena Cava Filter device is implanted before retrieval x D Therefore Cordis has updated the Warning statement in accordance with regulatory requirements as follows x D The OPTEASE Retrievable Filter has not been studied for long term implantation and must be retrieved within days after placement x D Previous Warning statement x D The OPTEASE Retrievable filter can be retrieved up to and including days after placement The OPTEASE Retrievable filter is considered a permanent implant if it is not retrieved within the specified time period x D The statement below was removed from the device description for the OPTEASE Retrieval Catheter IFU in accordance with the updates made to the Warning statement in the IFU for the Cordis OPTEASE Retrievable Vena Cava Filter and OPTEASE Retrieval Catheter x D The OPTEASE Retrievable Filter can be retrieved within a specified period after implantation refer to the OPTEASE Retrievable Filter Instructions for Use or remain implanted as a permanent filter x D Please share this information with any of your staff involved in the use of this device | Actions requested on your part: 1.Read this Urgent Field Safety Notice. 2.Sign and return the enclosed Acknowledgement Form in accordance with the directionson the form. 3.Share this notification with anyone in your facility that needs to be informed. 4.Contact any other facilities that have been provided with units of the affected catalogcodes (all unexpired lots). 5.Retain a copy of this notice with the product. Kindly review the attached UFSN | 2023-12-20 10:57
Maquet Critical Care AB | FLOW-i | 126 | - | Internal tests at Getinge have revealed a cybersecurity related vulnerability which potentially could lead to Denial of Service DoS attack tampering or remote code execution via remote login to the Flow Family Anesthesia systems which could lead to hazardous situations | For users with system version 4.8 and above: If you have a Getinge Connect module (X10/X20) connected to the Ethernet port on the Flow Family Anesthesia system, no action is required. If your Flow Family Anesthesia system is connected to Connected Services directly via the Ethernet port on the Flow Family Anesthesia system: Please follow the steps described in Annex 1, Please contact your local Getinge representative if you wish to activate Connected Services with a Getinge Connect module. For users with system version 4.7 and below: If your Flow Family Anesthesia system is not connected to Remote Services, no action is required. Do not connect an Ethernet cable to use Remote Services. If your Flow Family Anesthesia system is connected to Remote Services: Please follow the steps described in Annex 1 A Getinge Connect module is not available for system version 4.7 and lower. Please contact your local Getinge representative if you wish to activate Connected Services in the future with a Getinge Connect module (requires system version 4.8 and above). | 2023-12-21 11:14
Lowenstein Medical | Elisa 300 | - | - | To ensure stress free therapy without constant alarms during bronchoscopy the elisa VIT ventilators automatically extend the alarm limits when the bronchoscopy maneuver is started x D When the maneuver is completed the originally set alarm limits are restored and displayed x D Due to an error constellation in the buffer memory the audible and visual alarms are not triggered in accordance with the displayed alarm limits after the bronchoscopy maneuver is deactivated | Required action by the operator/user: Clinics that have configured the bronchoscopy maneuver may no longer use this maneuver as a preventive measure. We expressly recommend that the bronchoscopy maneuver option be permanently deactivated in the configuration by your person which is responsible for the devices! All other clinics are not affected. Action by the manufacturer: An update will be made available for the two affected software levels (version 2.10.7 and 2.13.3), in which this error constellation is rectified and the correct alarm parameters are copied into the alarm system. The above-mentioned software update will be carried out for all devices of software levels 2.10.x and 2.13.x at the latest as part of the next annual maintenance or due service call. Clinics that had previously configured the bronchoscopy manoeuvre as active and would like to use the function again soon, please send an email to: klinik@loewensteinmedical.com with reference to this FSCA2023023 in order to prioritize the implementation of the measure. The software update does not require instruction, provided that the update is carried out at the identical software level. An update of all elisa intensive care ventilators with affected software version is generally required. | 2023-12-21 13:32
Lowenstein Medical | Elisa 500 | - | - | To ensure stress free therapy without constant alarms during bronchoscopy the elisa VIT ventilators automatically extend the alarm limits when the bronchoscopy maneuver is started x D When the maneuver is completed the originally set alarm limits are restored and displayed x D Due to an error constellation in the buffer memory the audible and visual alarms are not triggered in accordance with the displayed alarm limits after the bronchoscopy maneuver is deactivated | Required action by the operator/user: Clinics that have configured the bronchoscopy maneuver may no longer use this maneuver as a preventive measure. We expressly recommend that the bronchoscopy maneuver option be permanently deactivated in the configuration by your person which is responsible for the devices! All other clinics are not affected. Action by the manufacturer: An update will be made available for the two affected software levels (version 2.10.7 and 2.13.3), in which this error constellation is rectified and the correct alarm parameters are copied into the alarm system. The above-mentioned software update will be carried out for all devices of software levels 2.10.x and 2.13.x at the latest as part of the next annual maintenance or due service call. Clinics that had previously configured the bronchoscopy manoeuvre as active and would like to use the function again soon, please send an email to: klinik@loewensteinmedical.com with reference to this FSCA2023023 in order to prioritize the implementation of the measure. The software update does not require instruction, provided that the update is carried out at the identical software level. An update of all elisa intensive care ventilators with affected software version is generally required. | 2023-12-21 13:32
Lowenstein Medical | Elisa 600 | - | - | To ensure stress free therapy without constant alarms during bronchoscopy the elisa VIT ventilators automatically extend the alarm limits when the bronchoscopy maneuver is started x D When the maneuver is completed the originally set alarm limits are restored and displayed x D Due to an error constellation in the buffer memory the audible and visual alarms are not triggered in accordance with the displayed alarm limits after the bronchoscopy maneuver is deactivated | Required action by the operator/user: Clinics that have configured the bronchoscopy maneuver may no longer use this maneuver as a preventive measure. We expressly recommend that the bronchoscopy maneuver option be permanently deactivated in the configuration by your person which is responsible for the devices! All other clinics are not affected. Action by the manufacturer: An update will be made available for the two affected software levels (version 2.10.7 and 2.13.3), in which this error constellation is rectified and the correct alarm parameters are copied into the alarm system. The above-mentioned software update will be carried out for all devices of software levels 2.10.x and 2.13.x at the latest as part of the next annual maintenance or due service call. Clinics that had previously configured the bronchoscopy manoeuvre as active and would like to use the function again soon, please send an email to: klinik@loewensteinmedical.com with reference to this FSCA2023023 in order to prioritize the implementation of the measure. The software update does not require instruction, provided that the update is carried out at the identical software level. An update of all elisa intensive care ventilators with affected software version is generally required. | 2023-12-21 13:32
Lowenstein Medical | Elisa 800 | - | - | To ensure stress free therapy without constant alarms during bronchoscopy the elisa VIT ventilators automatically extend the alarm limits when the bronchoscopy maneuver is started x D When the maneuver is completed the originally set alarm limits are restored and displayed x D Due to an error constellation in the buffer memory the audible and visual alarms are not triggered in accordance with the displayed alarm limits after the bronchoscopy maneuver is deactivated | Required action by the operator/user: Clinics that have configured the bronchoscopy maneuver may no longer use this maneuver as a preventive measure. We expressly recommend that the bronchoscopy maneuver option be permanently deactivated in the configuration by your person which is responsible for the devices! All other clinics are not affected. Action by the manufacturer: An update will be made available for the two affected software levels (version 2.10.7 and 2.13.3), in which this error constellation is rectified and the correct alarm parameters are copied into the alarm system. The above-mentioned software update will be carried out for all devices of software levels 2.10.x and 2.13.x at the latest as part of the next annual maintenance or due service call. Clinics that had previously configured the bronchoscopy manoeuvre as active and would like to use the function again soon, please send an email to: klinik@loewensteinmedical.com with reference to this FSCA2023023 in order to prioritize the implementation of the measure. The software update does not require instruction, provided that the update is carried out at the identical software level. An update of all elisa intensive care ventilators with affected software version is generally required. | 2023-12-21 13:32
Lowenstein Medical | Elisa 800VIT | - | - | To ensure stress free therapy without constant alarms during bronchoscopy the elisa VIT ventilators automatically extend the alarm limits when the bronchoscopy maneuver is started x D When the maneuver is completed the originally set alarm limits are restored and displayed x D Due to an error constellation in the buffer memory the audible and visual alarms are not triggered in accordance with the displayed alarm limits after the bronchoscopy maneuver is deactivated | Required action by the operator/user: Clinics that have configured the bronchoscopy maneuver may no longer use this maneuver as a preventive measure. We expressly recommend that the bronchoscopy maneuver option be permanently deactivated in the configuration by your person which is responsible for the devices! All other clinics are not affected. Action by the manufacturer: An update will be made available for the two affected software levels (version 2.10.7 and 2.13.3), in which this error constellation is rectified and the correct alarm parameters are copied into the alarm system. The above-mentioned software update will be carried out for all devices of software levels 2.10.x and 2.13.x at the latest as part of the next annual maintenance or due service call. Clinics that had previously configured the bronchoscopy manoeuvre as active and would like to use the function again soon, please send an email to: klinik@loewensteinmedical.com with reference to this FSCA2023023 in order to prioritize the implementation of the measure. The software update does not require instruction, provided that the update is carried out at the identical software level. An update of all elisa intensive care ventilators with affected software version is generally required. | 2023-12-21 13:32
BALT EXTRUSION SAS | SONIC | 43 | - | The investigation of the returned products confirmed that the products dimensions do not match the specifications the outer diameter of the HYBRID guidewires inside the pouch does not correspond to the outer diameter mentioned on the labels x D The investigation revealed that the root cause of these complaints was an isolated human error in production Confirmation testing is being performed on our remaining inventory to confirm the scope of this issue is limited to the lot subject to this FSN | Procedure to be applied by distributors: • Inform your customers and your local competent authority about this notice (outside EEA, UK, Switzerland, and Turkey). • Complete and return the “Notice Receipt form” below (Appendix section) as soon as possible to the e-mail address: Claim@baltgroup.com. • Contact BALT Extrusion SAS for any additional information. Procedure to be applied by the hospital staff: • Communicate this information to staff within the hospital that may use the above-mentioned references and lots (see above for details) or any other person if deemed necessary. • Complete and return the “Notice Receipt form” below (Annex section) as soon as possible to the e-mail address: Claim@baltgroup.com. By returning the completed Notice Receipt form by e-mail or mail, you acknowledge that you have read and understood this Field Safety Notice. • Contact Balt Extrusion SAS or your local distributor for any additional information. | 2023-12-21 14:56
Olympus Medical Systems Corp. | Single Use Distal Cover MAJ-2315 | - | - | Olympus has become aware that the Single Use Distal Cover MAJ may unexpectedly detach from the duodenoscope during use Olympus received complaints of the MAJ distal cover falling off within or outside patients This included events considered to be Serious Injury which involved foreign body retrieval and or additional treatment Consequences of a detached cover could include the risk of mucosal injury tissue damage bleeding perforation or obstruction due to retained distal cover in the gastrointestinal tract that would require urgent medical intervention for removal and or burns from the use an uncovered duodenoscope distal end Detachment of a distal cover in the oral cavity may result in aspiration airway obstruction or respiratory distress and would require immediate medical intervention for removal Appropriate medical intervention management should be based on the clinical circumstance x D Olympus is reinforcing the instructions for MAJ distal cover attachment and user verification of distal cover attachment This Field Safety Notice instructs you on the correct method for MAJ distal cover attachment You will be contacted by an Olympus representative to schedule on site in service training at your user facility on MAJ distal cover attachment and verification of cover attachment to the duodenoscope s distal end x D As required in the Instruction for Use it is important to ensure that the distal cover is attached correctly and checked carefully before conducting a procedure As a reminder we are including with this letter the below illustrations on how to correctly attach the MAJ distal cover and confirm distal cover attachment x D Additionally as a reminder do not apply anti fogging products such as olive oil or products containing petroleum based substances to the MAJ distal cover or the endoscope These products may cause cracks in the MAJ distal cover which could contribute to distal cover detachment thermal injury from electric current leaks when performing high frequency cauterization treatment or mucosal injury from the sharp edges due to cracks on the distal cover x D | 1. Carefully read the content of this Field Safety Notice. 2. Please ensure that all Olympus TJF duodenoscope users at your facility are knowledgeable on MAJ-2315 distal cover attachment and attachment verification to ensure the safe use of these devices. 3. Please ensure that the reference guide is distributed to all parties at your facility who will be involved in the process of attaching the MAJ-2315 distal cover. 4. indicate on the Reply Form that you have received and understood this Field Safety Notice by filling out and returning the completed enclosed Reply Form back to your local Olympus representative latest by 10.01.2024. 5. in-service training at your user facility on MAJ-2315 distal cover attachment and attachment verification to ensure the safe use of these devices. | 2023-12-21 17:18
Maquet Cardiopulmonary GmbH | HLS Set | 80 | - | Getinge discovered narrow channels in the seal of the sterile bags used to package the Emergency Priming Line Getinge receives this sterile bag from a supplier that performs tests to verify sterility Incomplete sealing can potentially lead to a breach of the sterile barrier x D x D x D x D Therefore Emergency Priming Lines affected by this Field Safety Notice may only be used immediately after opening an HLS Set or else disposed of if not needed x D x D x D x D This testing demonstrated that three supplier batches of sterile bags were affected Therefore this Field Safety Notice is limited to products that contain sterile bags from the three affected supplier batches | to Identify Device and to Return Device | 2023-12-24 16:43
Siemens Healthcare GmbH | ARTIS icono floor | 1 | - | During transversal table movement it could happen that the rolling bearing may get dislocated it is part of the linear x D guide rail and enables the transversal table movement As a consequence of this unlikely event the rolling bearing may x D fall out of the linear guide rail As a further consequence the mechanical connection between table top and table base x D may get lost x D It has not yet been reported to us that this issue occurred during clinical use of the system | An inspection will be performed to check the correct positioning of the linear guide rails of the Siemens Healthineers VE OR Table. An updated Artis Addendum will be distributed with the corrective action. | 2023-12-26 10:33
GUERBET | secufill | 39647 | - | Some particulates have the potential to detach from the device and in rare circumstances might enter the fluid x D path and be injected into the patients vascular system with the saline contrast medium solution This may cause x D an embolism that could damage some organs thus resulting in serious injury to the patients To date MEDEX x D has not received any report of adverse events nor patient injury related to this matter and is issuing this x D notification out of an abundance of caution | MEDEX distribution records indicate that the product(s) described in the section 4 was/were shipped to you. Therefore, please: a) Discontinue immediately the use of the products listed in Appendix 1. b) Check your inventory and quarantine all units from the products listed in Appendix 1 at your facility, and contact your Local Sales Representative to organize the return of the affected products. c) Make sure that all caregivers and users of these products are made aware of this Field Safety Notice. d) Complete and sign the enclosed acknowledgment form (in Appendix 2) and return it to the Local Sales Representative (their contact point is indicated in Appendix 2). MEDEX has notified the Competent/ Regulatory Authorities of the countries concerned by this Field Safety Notice and informed them about this quality defect and the above-written actions. Please, report all device-related incidents to your Local Sales Representative and/or to your Competent/ Regulatory Authority if appropriate, as this provides important feedback. MEDEX has implemented additional control during the manufacturing process to ensure that only batches meeting the approved predefined specifications for particulate formation are released on the market. We sincerely apologize for any inconvenience that this problem may cause. MEDEX is committed to providing the highest level of support, and we thank you for assisting with this process. Please, contact your Local Sales Representative with any questions regarding this notification. Sincerely, | 2023-12-26 12:19
GE Healthcare Finland Oy | Centricity™ High Acuity Clinical Notification Module | 13 | - | Patient allergies that are imported from external systems via interface messaging to CCC CA and CHA systems could be out of sync and users may not be aware of this Once a patient allergy is imported via the interface the data will remain in the system Allergy deletion is not supported via interface messaging this needs to be done manually in the application x D | A new User Manual supplement for all affected customers to clarify the limitations in the inbound allergy interface. Customers having CHA versions 5.0 and later, in addition to the User Manual supplement: Field Engineers implement configuration changes in CHA System interface and User Interface to enable end users to validate and reconcile the allergy data. A new service manual detailing the FMI 38011 configuration changes in System Interface and User Interface. | 2023-12-27 12:21
GE Healthcare Finland Oy | Centricity High Acuity Anesthesia 5.X | 13 | - | Patient allergies that are imported from external systems via interface messaging to CCC CA and CHA systems could be out of sync and users may not be aware of this Once a patient allergy is imported via the interface the data will remain in the system Allergy deletion is not supported via interface messaging this needs to be done manually in the application x D | A new User Manual supplement for all affected customers to clarify the limitations in the inbound allergy interface. Customers having CHA versions 5.0 and later, in addition to the User Manual supplement: Field Engineers implement configuration changes in CHA System interface and User Interface to enable end users to validate and reconcile the allergy data. A new service manual detailing the FMI 38011 configuration changes in System Interface and User Interface. | 2023-12-27 12:21
Siemens Shanghai Medical Equipment Ltd. | Multix Fusion Max | - | - | During system installation the mounting of the BABIX holder was possibly performed incorrectly | Service engineer will check the mounting for the BABIX holder on the Bucky wall stand and if necessary, mount it correctly. | 2023-12-27 13:09
Siemens Shanghai Medical Equipment Ltd. | MULTIX Impact C | - | - | During system installation the mounting of the BABIX holder was possibly performed incorrectly | Service engineer will check the mounting for the BABIX holder on the Bucky wall stand and if necessary, mount it correctly. | 2023-12-27 13:09
Welch Allyn, Inc. | Welch Allyn®X Scribe™ | - | - | Baxter Healthcare Corporation is issuing a Correction for the Q Stress and XScribe Cardiac x D Stress Testing Systems Q Stress version or higher and XScribe version or higher due to a x D potential change in the QRS amplitude identified in electrocardiogram ECG readings when x D the Source Consistency Filter SCF is enabled The SCF is a configurable feature with x D multichannel recordings which aims to reduce noise and remove inconsistent signals without x D distorting the ECG signal associated with stress testing | Due to the potential impact on the QRS complex, Baxter is developing a software update for Q-Stress version 6 or higher and XScribe version 6 or higher to resolve the issue. Baxter will contact you when the software update is available. a team has been gathered to develop the software upgrade. A planning camp is scheduled for January 2024, and we are hoping to have a timeline for software release established by the end of January 2024. Once a timeline is established for the release of the software update, Baxter will establish a target completion date for this FSCA. | 2023-12-28 12:30
Meril Life Sciences Pvt. Ltd. | Myval Transcatheter Heart Valve System | 14 | استخدام صمام القلب (Myval Transcatheter Heart Valve System) لغير الغرض المصنع من أجله بالموضع ( Pulmonary position ) من قبل عدد 2 من مراكز أمراض وجراحة القلب بالمملكة. | Myval Transcatheter Heart Valve System Pulmonary position | - | 2023-12-28 23:37
Philips Medical Systems Nederland B.V. | INGENIA AMBITION X | 1 | - | Philips has identified that an incorrect electrical shock warning label may have been placed on the Liquid Cooling Cabinet LCC of the following MR systems with a F LR transformer component Achieva T Achieva T Conversion Achieva T Ingenia Ambition X Intera T and SmartPath to dStream for T systems x D The caution symbol on the incorrect label has a white background instead of yellow x D The label is missing the high voltage warning symbol x D | A Philips representative will contact you to schedule a time for a Field Service Engineer (FSE) to verify the electrical shock warning label and apply the correct label if applicable (Reference FCO78100574). | 2023-12-29 10:21
Datex-Ohmeda, Inc. 3030 Ohmeda Drive, PO Box 7550, Madison, 53707-7550, USA | Evair | 90 | - | GE HealthCare has become aware of elevated levels of formaldehyde when the EVair or EVair Jun Air optional compressors are used with the CARESCAPE R or Engstr m Carestation Pro ventilators respectively These elevated results were observed in preliminary testing that was conducted at an elevated room temperature of oC oF at the lowest possible flow condition of L min worst case minimum bias flow with no additional ventilation and all of the gas being supplied from the compressor i e FiO of no supplemental oxygen GE HealthCare is continuing to evaluate the root cause for these elevated levels of formaldehyde x D x D While these test conditions are not representative of typical clinical use conditions GE HealthCare is taking this action to further reduce the potential for patient exposure The amount of exposure to formaldehyde during typical clinical use and the potential risks to health from such exposure if any are currently unknown However in an unlikely scenario in which the compressors are used at the conditions described above the formaldehyde levels may lead to adverse pulmonary affects such as the potential for transient reversible airway irritation or inflammation that could lead to airway hyperresponsiveness e g asthma in neonates or infants resulting in additional medical intervention e g bronchodilator administration adjustment of ventilator settings increased duration or degree of ventilatory support and or oxygen support x D x D Compressors are optional accessories for ventilators and are only used when wall air is not available which is not typical in most hospital settings x D x D GE HealthCare has not received any reports of patient injury or adverse effects related to potential exposure to formaldehyde from the use of the compressors with ventilators x D | GE HealthCare has provided the attached Addendum with the updated operating conditions (30oC (86oF) maximum room air temperature) and indications for use specifying that the compressors are not recommended for use with neonates and infants (0-2 years of age). | 2023-12-31 13:37
HUMAN Gesellschaft fȕr Biochemica & Diagnostica mbH | AUTOCAL | 350 | - | caused by decreasing enzyme activities and altered internal traceability procedures | inform the customers about the issue of the affected Lots, based on this Urgent Field Safety Notice | 2023-12-31 15:06
HUMAN Gesellschaft fȕr Biochemica & Diagnostica mbH | HumaTrol N | 350 | - | caused by decreasing enzyme activities and altered internal traceability procedures | inform the customers about the issue of the affected Lots, based on this Urgent Field Safety Notice | 2023-12-31 15:06
HUMAN Gesellschaft fȕr Biochemica & Diagnostica mbH | HumaTrol P | 350 | - | caused by decreasing enzyme activities and altered internal traceability procedures | inform the customers about the issue of the affected Lots, based on this Urgent Field Safety Notice | 2023-12-31 15:06
HUMAN Gesellschaft fȕr Biochemica & Diagnostica mbH | Serodos | 350 | - | caused by decreasing enzyme activities and altered internal traceability procedures | inform the customers about the issue of the affected Lots, based on this Urgent Field Safety Notice | 2023-12-31 15:06
HUMAN Gesellschaft fȕr Biochemica & Diagnostica mbH | SERODOS® PLUS | 350 | - | caused by decreasing enzyme activities and altered internal traceability procedures | inform the customers about the issue of the affected Lots, based on this Urgent Field Safety Notice | 2023-12-31 15:06
Philips Medical Systems Nederland B.V. | Allura Xper FD10 | 135 | - | Three components of the PCs may not perform as intended due to manufacturing issues The PCs within your system may have one or more of the impacted components | Philips will be replacing the three (3) components in all affected systems. Component replacements will be prioritized to those customers that have experienced or may be experiencing a component issue, followed by age of the PC. Philips will contact you to schedule a visit to replace the affected components (reference 2023-IGT-BST-027). | 2024-01-02 09:29
Philips Medical Systems Nederland B.V. | Allura Xper FD10 OR Table | 135 | - | Three components of the PCs may not perform as intended due to manufacturing issues The PCs within your system may have one or more of the impacted components | Philips will be replacing the three (3) components in all affected systems. Component replacements will be prioritized to those customers that have experienced or may be experiencing a component issue, followed by age of the PC. Philips will contact you to schedule a visit to replace the affected components (reference 2023-IGT-BST-027). | 2024-01-02 09:29
Philips Medical Systems Nederland B.V. | Allura Xper FD10/10 | 135 | - | Three components of the PCs may not perform as intended due to manufacturing issues The PCs within your system may have one or more of the impacted components | Philips will be replacing the three (3) components in all affected systems. Component replacements will be prioritized to those customers that have experienced or may be experiencing a component issue, followed by age of the PC. Philips will contact you to schedule a visit to replace the affected components (reference 2023-IGT-BST-027). | 2024-01-02 09:29
Philips Medical Systems Nederland B.V. | Allura Xper FD20 | 135 | - | Three components of the PCs may not perform as intended due to manufacturing issues The PCs within your system may have one or more of the impacted components | Philips will be replacing the three (3) components in all affected systems. Component replacements will be prioritized to those customers that have experienced or may be experiencing a component issue, followed by age of the PC. Philips will contact you to schedule a visit to replace the affected components (reference 2023-IGT-BST-027). | 2024-01-02 09:29
Philips Medical Systems Nederland B.V. | Allura Xper FD20 OR Table | 135 | - | Three components of the PCs may not perform as intended due to manufacturing issues The PCs within your system may have one or more of the impacted components | Philips will be replacing the three (3) components in all affected systems. Component replacements will be prioritized to those customers that have experienced or may be experiencing a component issue, followed by age of the PC. Philips will contact you to schedule a visit to replace the affected components (reference 2023-IGT-BST-027). | 2024-01-02 09:29
Philips Medical Systems Nederland B.V. | Allura Xper FD20/10 | 135 | - | Three components of the PCs may not perform as intended due to manufacturing issues The PCs within your system may have one or more of the impacted components | Philips will be replacing the three (3) components in all affected systems. Component replacements will be prioritized to those customers that have experienced or may be experiencing a component issue, followed by age of the PC. Philips will contact you to schedule a visit to replace the affected components (reference 2023-IGT-BST-027). | 2024-01-02 09:29
Philips Medical Systems Nederland B.V. | Allura Xper FD20/15 | 135 | - | Three components of the PCs may not perform as intended due to manufacturing issues The PCs within your system may have one or more of the impacted components | Philips will be replacing the three (3) components in all affected systems. Component replacements will be prioritized to those customers that have experienced or may be experiencing a component issue, followed by age of the PC. Philips will contact you to schedule a visit to replace the affected components (reference 2023-IGT-BST-027). | 2024-01-02 09:29
Philips Medical Systems Nederland B.V. | AZURION 3 M12 | 135 | - | Three components of the PCs may not perform as intended due to manufacturing issues The PCs within your system may have one or more of the impacted components | Philips will be replacing the three (3) components in all affected systems. Component replacements will be prioritized to those customers that have experienced or may be experiencing a component issue, followed by age of the PC. Philips will contact you to schedule a visit to replace the affected components (reference 2023-IGT-BST-027). | 2024-01-02 09:29
Philips Medical Systems Nederland B.V. | AZURION 3 M15 | 135 | - | Three components of the PCs may not perform as intended due to manufacturing issues The PCs within your system may have one or more of the impacted components | Philips will be replacing the three (3) components in all affected systems. Component replacements will be prioritized to those customers that have experienced or may be experiencing a component issue, followed by age of the PC. Philips will contact you to schedule a visit to replace the affected components (reference 2023-IGT-BST-027). | 2024-01-02 09:29
Philips Medical Systems Nederland B.V. | AZURION 7 B12 | 135 | - | Three components of the PCs may not perform as intended due to manufacturing issues The PCs within your system may have one or more of the impacted components | Philips will be replacing the three (3) components in all affected systems. Component replacements will be prioritized to those customers that have experienced or may be experiencing a component issue, followed by age of the PC. Philips will contact you to schedule a visit to replace the affected components (reference 2023-IGT-BST-027). | 2024-01-02 09:29
Philips Medical Systems Nederland B.V. | AZURION 7 B20 | 135 | - | Three components of the PCs may not perform as intended due to manufacturing issues The PCs within your system may have one or more of the impacted components | Philips will be replacing the three (3) components in all affected systems. Component replacements will be prioritized to those customers that have experienced or may be experiencing a component issue, followed by age of the PC. Philips will contact you to schedule a visit to replace the affected components (reference 2023-IGT-BST-027). | 2024-01-02 09:29
Philips Medical Systems Nederland B.V. | AZURION 7 M12 | 135 | - | Three components of the PCs may not perform as intended due to manufacturing issues The PCs within your system may have one or more of the impacted components | Philips will be replacing the three (3) components in all affected systems. Component replacements will be prioritized to those customers that have experienced or may be experiencing a component issue, followed by age of the PC. Philips will contact you to schedule a visit to replace the affected components (reference 2023-IGT-BST-027). | 2024-01-02 09:29
Philips Medical Systems Nederland B.V. | AZURION 7 M20 | 135 | - | Three components of the PCs may not perform as intended due to manufacturing issues The PCs within your system may have one or more of the impacted components | Philips will be replacing the three (3) components in all affected systems. Component replacements will be prioritized to those customers that have experienced or may be experiencing a component issue, followed by age of the PC. Philips will contact you to schedule a visit to replace the affected components (reference 2023-IGT-BST-027). | 2024-01-02 09:29
Philips Medical Systems Nederland B.V. | Allura Centron | 135 | - | Three components of the PCs may not perform as intended due to manufacturing issues The PCs within your system may have one or more of the impacted components | Philips will be replacing the three (3) components in all affected systems. Component replacements will be prioritized to those customers that have experienced or may be experiencing a component issue, followed by age of the PC. Philips will contact you to schedule a visit to replace the affected components (reference 2023-IGT-BST-027). | 2024-01-02 09:29
Philips | Allura Xper FD20 Biplane | 135 | - | Three components of the PCs may not perform as intended due to manufacturing issues The PCs within your system may have one or more of the impacted components | Philips will be replacing the three (3) components in all affected systems. Component replacements will be prioritized to those customers that have experienced or may be experiencing a component issue, followed by age of the PC. Philips will contact you to schedule a visit to replace the affected components (reference 2023-IGT-BST-027). | 2024-01-02 09:29
GE KLINIK SISTEMLER TICARET A.S Mustafa Kemal Mah.2158.Sok No:9 Eskisehir Yolu 06520 Ankara/Turkiye | All Care Plus, Care Plus® Model 1000, 2000, 3000 and 4000 Incubator, Lullaby Incubator, Lullaby TR Incubator, and Lullaby XP Incubators | 100 | - | GE HealthCare has become aware that the bedside panel of Care Plus Care Plus Model and Incubator Lullaby Incubator Lullaby TR Incubator and Lullaby XP Incubator can be upright and look closed but not be latched if the red indicator is visible The portholes also can look closed when not latched when the portholes are obstructed e g hood covers If a patient impacts the x D bedside panel or porthole that is not latched the panel can disengage and fall open no longer protecting the patient from falling | Customer letter will be shipped to all affected consignees to inspect all their units and if affected unit is located, customer need share the unit serial number to ship the kit Kit with Label and IFU will be shipped to active affected units to apply the label and place IFU addendum with manual | 2024-01-03 12:15
Smith medical | Syringe Infusion Pump | 25 | - | Pump may not detect over infusion or under infusion of medication | Software upgrade | 2024-01-03 13:41
Immunotech S.A.S a Beckman Coulter Company | Access hsTnI | 62 | - | Beckman Coulter has determined that a residual risk for intra assay carryover into an Access hsTnI reagent pack remains possible after upgrading to system software and above This risk is present when using the optional automated Access hsTnI Onboard Dilution OBD assay x D x D Clinically significant carryover into a reagent pack may occur when a sample with a cardiac Troponin I cTnI concentration greater than the S calibrator pg mL is diluted using the Access hsTnI OBD assay on a UniCel DxI or UniCel DxI immunoassay analyser | Beckman Coulter will implement a design change that removes DxI system onboard dilution information from the Access hsTnI reagent (IFU) and removes the OBD assay from the UniCel DxI 600 and UniCel DxI 800 instruments. | 2024-01-07 08:34
Abbott | HeartMate Touch Communication System | - | - | During pump priming do not disconnect the pump until the timer reaches zero the pump stops and the Priming x D is complete message appears as instructed per the IFU x D x D x D The following scenarios can occur if the STOP PUMP sequence is not finished on a HeartMate Touch System x D After the pump is primed if the same HeartMate Touch System is then connected to the pump during the implant x D procedure the LVAD may unexpectedly start pumping x D When the user connects the HeartMate Touch App System to the same or a different HeartMate System Controller x D with a running LVAD the pump will stop If the pump stops the HeartMate System controller will alarm Pump Off x D Alarm x D | Abbott is in the process of developing and releasing a software update to prevent the sequence of events described above. This product correction will be included in an upcoming release of the HeartMate Touch software. An Abbott Representative will contact you for a software upgrade beginning Q2 2024 | 2024-01-07 16:58
Philips | SmartPath to dStream for 1.5T | 21 | - | The Quadrature Body Coil QBC seal adhesive may fail creating sharp edges that may come in contact with patients The QBC seal may become loose as the patient table travels in a horizontal motion in and out of the system bore The QBC seal Figure is a rubber seal that is glued between the cone cover and QBC cover and functions to prevent sharp edges of the QBC cover from contacting patients during an examination | Philips will contact you to schedule a time for a Field Service Engineer (FSE) to visit your site and replace your system’s QBC Seal (reference FCO78100573). Philips plans to start implementing corrections in Q3 2024. | 2024-01-08 13:43
Philips | Intera Achieva 1.5T Pulsar | 21 | - | The Quadrature Body Coil QBC seal adhesive may fail creating sharp edges that may come in contact with patients The QBC seal may become loose as the patient table travels in a horizontal motion in and out of the system bore The QBC seal Figure is a rubber seal that is glued between the cone cover and QBC cover and functions to prevent sharp edges of the QBC cover from contacting patients during an examination | Philips will contact you to schedule a time for a Field Service Engineer (FSE) to visit your site and replace your system’s QBC Seal (reference FCO78100573). Philips plans to start implementing corrections in Q3 2024. | 2024-01-08 13:43
Philips Medical Systems Nederland B.V. | Achieva 3.0T | 21 | - | The Quadrature Body Coil QBC seal adhesive may fail creating sharp edges that may come in contact with patients The QBC seal may become loose as the patient table travels in a horizontal motion in and out of the system bore The QBC seal Figure is a rubber seal that is glued between the cone cover and QBC cover and functions to prevent sharp edges of the QBC cover from contacting patients during an examination | Philips will contact you to schedule a time for a Field Service Engineer (FSE) to visit your site and replace your system’s QBC Seal (reference FCO78100573). Philips plans to start implementing corrections in Q3 2024. | 2024-01-08 13:43
Philips | Intera 1.5T Achieva Nova-Dual | 21 | - | The Quadrature Body Coil QBC seal adhesive may fail creating sharp edges that may come in contact with patients The QBC seal may become loose as the patient table travels in a horizontal motion in and out of the system bore The QBC seal Figure is a rubber seal that is glued between the cone cover and QBC cover and functions to prevent sharp edges of the QBC cover from contacting patients during an examination | Philips will contact you to schedule a time for a Field Service Engineer (FSE) to visit your site and replace your system’s QBC Seal (reference FCO78100573). Philips plans to start implementing corrections in Q3 2024. | 2024-01-08 13:43
Philips Medical Systems Nederland B.V. | Achieva 1.5T | 21 | - | The Quadrature Body Coil QBC seal adhesive may fail creating sharp edges that may come in contact with patients The QBC seal may become loose as the patient table travels in a horizontal motion in and out of the system bore The QBC seal Figure is a rubber seal that is glued between the cone cover and QBC cover and functions to prevent sharp edges of the QBC cover from contacting patients during an examination | Philips will contact you to schedule a time for a Field Service Engineer (FSE) to visit your site and replace your system’s QBC Seal (reference FCO78100573). Philips plans to start implementing corrections in Q3 2024. | 2024-01-08 13:43
Philips Medical Systems Nederland B.V. | Ingenia 1.5T CX | 21 | - | The Quadrature Body Coil QBC seal adhesive may fail creating sharp edges that may come in contact with patients The QBC seal may become loose as the patient table travels in a horizontal motion in and out of the system bore The QBC seal Figure is a rubber seal that is glued between the cone cover and QBC cover and functions to prevent sharp edges of the QBC cover from contacting patients during an examination | Philips will contact you to schedule a time for a Field Service Engineer (FSE) to visit your site and replace your system’s QBC Seal (reference FCO78100573). Philips plans to start implementing corrections in Q3 2024. | 2024-01-08 13:43
Becton Dickinson and Company | BD BBL™ Staphyloslide™ Latex Test | 13 | - | BD confirmed through a recent evaluation of BD BBL Sensi Disc product that twenty eight out of thirty antimicrobial discs showed reproducibility accuracy and or Quality Control QC failures when tested with Haemophilus spp Testing with other bacterial pathogens where indicated is not affected | BD Actions: • BD is investigating root cause and will implement appropriate corrective and preventative measures to prevent recurrence. • BD will issue credit to customers that have catalogue number 291270, following receipt of the completed Customer Response Form. • Future lots of BD BBL™ Sensi-Disc™ provided (excluding catalogue number 291270) will be labelled with a sticker notifying customers “This product should not be used for the semi-quantitative in vitro susceptibility testing of Haemophilus influenzae”. Actions for Clinical Users • Users should refrain from conducting AST testing for Haemophilus species with affected BD BBL™ Sensi-Disc™. The affected product can continue to be used with other bacterial pathogens (where indicated). • Cease use of catalogue number 291270, which is solely intended for AST testing for Haemophilus influenzae. • Previous test results do not need to be reviewed and no additional clinical actions are recommended. Customer Actions: For impacted lots of catalogue number 291270 (only): • Identify and quarantine all unused affected units of impacted lots of catalogue number 291270. • Make a note of the lot numbers, then destroy the affected units of catalogue number 291270. • NOTE: Other catalogue numbers with lot numbers in scope of this action can continue to be used with other bacterial pathogens (where indicated). For all other catalogue numbers in scope of this action: • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 16th February 2024. • Circulate this notice to all those who need to be aware within your organization or to any organization where the potentially affected products have been transferred. • If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: For impacted lots of catalogue number 291270 (only): • Cease distribution of catalogue number 291270 (only). • Identify, quarantine, making a note of the lot numbers then destroy all affected units of catalogue number 291270 (only). • Identify the facilities where you have distributed all affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 16th February 2024. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. For all other catalogue numbers in scope of this action: • Identify the facilities where you have distributed all affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 16th February 2024. • Complete and return the Customer Response Form following completion of your reconciliation activities. • For future shipments of impacted lots from your inventory (excluding catalogue number 291270, which should be discarded), ensure this letter is forwarded for awareness, until your inventory has depleted. No further action is required for lots that are not listed, as these have been labelled with a statement to not use the affected product for Haemophilus influenza. • If you experience any issues, please report as a complaint as per your normal process. | 2024-01-08 16:46
Randox Laboratories | RX Monza | - | - | There has been no non conformance for this product Randox are discontinuing the product | Randox Laboratories are discontinuing the RX Monza Analyser, parts supply and service support effective Dec.2024. | 2024-01-10 07:48
ICU - Smith Medical | syringe infusion pump | 25 | - | Smiths Medical identified that certain Barrel Clamp Guides from the above lots may contain a molding defect that could potentially lead to slippage of the spring within the barrel clamp assembly If this occurs it could result in the inability of the pump to recognize a syringe or the pump may misidentify the size of the syringe loaded | 1- customer awareness 2- customer singed acknowledgment | 2024-01-10 12:15
smith medical | syringe infusion pump | 25 | - | The calibration of the force sensor used to detect occlusions may shift over time | perform the Annual Maintenance test inform customer and signed acknowledgment | 2024-01-10 15:11
Randox Laboratories | HUMAN ASSAYED MULTI-SERA - LEVEL 3 (HUM ASY CONTROL 3) | - | - | There has been a transaction error on the Additional Value Sheet for Human Assayed Multi Sera Level HE lot UE for lipase The method was incorrectly listed as Colorimetric Siemens Dimension LIPL kit C The correct method is Colorimetric Siemens Dimension LIP kit C | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organization. | 2024-01-15 10:51
Randox Laboratories | HUMAN ASSAYED MULTI-SERA - LEVEL 3 (HUM ASY CONTROL 3) | - | - | There has been a transaction error on the Additional Value Sheet for Human Assayed Multi Sera Level HE lot UE for lipase The method was incorrectly listed as Colorimetric Siemens Dimension LIPL kit C The correct method is Colorimetric Siemens Dimension LIP kit C | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organization. | 2024-01-15 10:59
Randox Laboratories | HUMAN ASSAYED MULTI-SERA - LEVEL 3 (HUM ASY CONTROL 3) | - | - | There has been a transaction error on the Additional Value Sheet for Human Assayed Multi Sera Level HE lot UE for lipase The method was incorrectly listed as Colorimetric Siemens Dimension LIPL kit C The correct method is Colorimetric Siemens Dimension LIP kit C | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organization. | 2024-01-15 11:09
Shenzhen Mindray Scientific | BeneFusion nSP Syringe Pumps | 322 | - | Mindray has identified a potential issue associated with the BeneFusion nSP Syringe Pumps linked with environmental specifications x D x D As an electronic device the syringe pump generates heat when it is used The device may affect the temperature of the fluid inside the syringe associated with the comprehensive factors For your reference in the ambient temperature of C In the use case of channel pumps i e channel BeneFusion nSP stack operating simultaneously at a flow rate of mL h or less we observed a maximum temperature rise of C on the third channel at which point the drug solution in the syringe reaches its maximum temperature x D x D This increase in temperature can be aggravated by certain factors in particular when the ambient temperature is high the drug flow rate is low or ventilation is inadequate | Advise on action to be taken by the Hospital administrator: 1. Please pass the Field Safety Notice (FSN) to all those who need to be aware within your organization or to any organization where the affected BeneFusion nSP Syringe Pumps have been delivered. 2. You can continue to use the device when you receive this field safety notice (FSN) and the updated user manual. During use, all users within your organization need to be aware of the following matter: The device may affect the temperature of the fluid inside the syringe associated with the comprehensive factors including the high ambient temperature, extremely low drug flow rate and poor air circulation in room. Keep the vents of the device unblocked during use. 3. For drugs sensitive to temperature, please follow the instructions or manufacturer's guidelines, and carefully evaluate the applicable temperature and flow rates of the drugs. Also, consider a faster replacement of syringes for shorter exposure to heat. 4. Our local Mindray Service Representatives or cooperative distributors will contact you as soon as possible and will send the product manual of the new version to you. Advise on action to be taken by the distributor: 1. Please pass the Field Safety Notice (FSN) to all those who need to be aware within your organization or to any organization where the affected BeneFusion nSP Syringe Pumps have been delivered. 2. If any BeneFusion nSP Syringe Pumps in your facility is on the affected list, Mindray's Service Representatives will contact you as soon as possible and send the product manual of the new version to you. | 2024-01-15 12:32
Ortho Clinical Diagnostics | VITROS Immunodiagnostic Products B•R•A•H•M•S PCT Controls | 11 | - | VITROS B R A H M S PCT Controls Lot was assigned the incorrect expiration date of Jul during release analysis The correct true expiration date for Lot is Aug The incorrect expiration date appears on the product labeling packaging Certificate of Analysis and any other associated documentation for this product Please note that the performance of Lot is not affected by this issue when used by the true expiration date | •Complete the enclosed Confirmation of Receipt form no later than January 22, 2024. Upon receipt of your completed Confirmation of Receipt form, QuidelOrtho will provide credit or replacement of your inventory of VITROS B•R•A•H•M•S PCT Controls, Lot 0380. •Upon receipt of the replacement lot discontinue using, render unusable, and discard any remaining inventory of VITROS B•R•A•H•M•S PCT Controls, Lot 0380. •Save this notification with your User Documentation or post this notification by each VITROS ECi/ECiQ/3600/5600/XT 7600 System until the issue has been resolved. | 2024-01-16 09:15
Randox Laboratories | HUMAN ASSAYED MULTI-SERA - LEVEL 3 (HUM ASY CONTROL 3) | - | - | There has been a transaction error on the Additional Value Sheet for Human Assayed Multi Sera Level HE lot UE for lipase The method was incorrectly listed as Colorimetric Siemens Dimension LIPL kit C The correct method is Colorimetric Siemens Dimension LIP kit C | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organization. | 2024-01-17 11:42
Alba Bioscience Limited | ORTHO™ Sera Anti – Lea | 205 | - | A Customer Complaint investigation has confirmed that the lots of ORTHOTM Sera Anti Lea stated above show unexpected negative results with Lea antigen positive red cells This phenomenon was noted during testing of these lots against reagent red blood cells and EDTA red cells positively expressing the Lea antigen | • The reagents listed above should be removed from use and discarded. Quotient is recommending that a review of previous donor and patient testing performed with these lots of antisera should be conducted. • Repeat testing should be conducted using an alternative lot of antisera if previously obtained results appear anomalous. • All required regulatory notifications are being completed. We request that you help us meet our compliance requirements. Please perform the required actions: complete and sign the attached form and return to us at your earliest opportunity by email: Vigilance.Notifications@quotientbd.com. | 2024-01-17 13:04
GE Healthcare Austria GMBH & CO. OG | IC9-RS | 85 | - | GE HealthCare has become aware of an issue with certain IC RS ultrasound probes see Affected Products List that can result in a double image artifact creating a ghost image with realistic features An unrecognized artifact may lead to misdiagnosis | Repair or replace affected probes that exhibit a double image artifact. | 2024-01-17 14:19
3M Health Care | 3M™ Surgical Clipper with Pivot Head | 20 | - | Description of the problem and potential hazard and risk for the patient user x D This corrective action has been initiated to communicate an additional caution statement to x D customers using the above named products By applying this additional caution the safe usage x D of the products will be ensured This additional caution statement is provided below x D | • Do not use on injured skin. • Do not leave clipper running unused in air without applying to skin for more than 1 minute as blade temperature may exceed 41°C (105.8°F). • Do not leave the clipper blade applied to the same location on the patient’s skin for longer than 1 second | 2024-01-18 19:26
Maquet Critical Care AB | Base Unit Servo-N | 37 | - | This field action was initiated to inform users about a limitation of the intended performance of Servo n during high frequency oscillatory ventilation HFOV when using the Fisher Paykel neonatal circuit N and the humidifier FP | New information, including a warning, will be added into Servo-n IFU for System Version 4.4 and later software versions regarding the following topics: 1. ET Tube 2. Weight Ranges 3. Clarification why ‘Set Pampl cannot be reached’ alarm may occur Kindly refer to the attached documents for detailed explanation of this FSN. | 2024-01-18 19:52
Abbott GmbH. & Co. KG | Alinity s System | 19 | - | Two potential issues that may be exhibited when operating the Alinity s System with the above software version x D Issue Misassigned sample pipettor aspiration locations at the LAS can cause results to be wrongly attributed to the incorrect unique sample identifier SID x D Issue Pipettor positioning may be off target if a pipettor probe is not straight or homed correctly due to unintended contact between the pipettor or pipettor probe with another component during instrument operation x D | Please continue to operate and maintain the Alinity s System as described within the Alinity s System Operations Manual, refer to “Necessary Actions” in the attachment for more details. Abbott is releasing Alinity s System software version 2.9.0 (LN 04U76‐18) to prevent above potential issues and improve the system. | 2024-01-19 20:36
SoClean, Inc. | SoClean® 2 | 106 | - | SoClean has received approximately complaints equating to of all SoClean and SoClean units sold resulting from improper set up unauthorized device modifications and use by individuals for whom the device is not recommended as indicated in the user manual Of those complaints were FDA reportable adverse events and zero deaths were reported x D Common complaints include i the customer not being able to set up or turn on the device ii mildew smell in the hose iii excessive ozone smell iv cough and v exacerbation of preexisting condition s x D As stated in the user manual persons with underlying lung diseases such as asthma and chronic obstructive pulmonary disease also known as COPD which includes emphysema and chronic bronchitis and those with cardiovascular disease may be sensitive to ozone and should consult with their health care professional before using a SoClean or SoClean x D SoClean is revising its labeling to incorporate i additional warnings and contraindications and ii proper and consistent device set up instructions including use of a hose and mask adapter thereby reducing potential risks associated with the previous device design and labeling | - Obtain a copy of the new version of the user manual. - If you recently purchased a SoClean filter and only want to obtain a no cost Hose and Mask Adapter, please contact our Customer Care team. - Confirmation of Receipt: please complete the form to acknowledge that you’ve read and understand this communication. | 2024-01-21 08:23
Balt USA, LLC | OPTIMA COIL SYSTEM | 37 | - | On November Balt USA became aware of a situation where discoloration along the delivery system pusher component of the Optima Coil System was identified during in house inspection of released finished good products Discoloration was observed on the outside of the PETjacket along the hypotube component | Procedure to be applied by the hospital staff: • Stop or cease use of the affected product. • Inform your hospital staff, such as safety officers, pharmacists, purchasing agent, head of neuroradiology and the neuroradiology staff, physicians including but not limited to endovascular neurosurgeons and interventional neurologists, as well as any other person if deemed necessary. • Identify and locate the Optima Coil System products concerned by this field safety notice (FSN). • Collect and put in quarantine the Optima Coil System products concerned by this FSN, then return them to Balt USA throughthe usual “RGA” (Return Goods Authorization) procedure by contacting Balt USA’s Customer Service department. • Keep informed by Balt USA about the status of every unit of Optima Coil System products concerned by this recall procedure. • Fulfill the “Notice of receipt” (refer to annex on pages 3 and 4), then return it to Balt USA via the indicated contact. • Contact Balt USA for any additional information. Procedure to be applied by distributors/subsidiaries: • Stop or cease processing of the affected recall product. • Inform your customers about this notice. • Identify and locate the Optima Coil System products concerned by this FSN. • Collect and put in quarantine the Optima Coil System products concerned by this FSN and then return them to your localdistributor as per their return procedure. • Keep informed by Balt USA about the status of every unit of Optima Coil System product concerned by this FSN. • Fulfill the “Notice of receipt” (refer to annex on pages 3 and 4), then return it to Balt USA via the indicated contact. • Contact Balt USA for any additional information. | 2024-01-21 11:32
Abbott Laboratories | Alinity s System | 19 | - | Hardware issue associated with the assembly of the Alinity s System incubation track x D | Please continue to operate and maintain the Alinity s System as described within the Alinity s System Operations Manual, refer to “Necessary Actions” in the attachment for more details. Your local Abbott representative will be performing an inspection of your Alinity s System(s) to confirm the five standoffs summarized in the attachment are correct. If any standoffs are identified as incorrect, your Abbott representative will work with your site to schedule the required service activities to replace these standoffs. | 2024-01-22 11:33
Siemens Healthcare Diagnostics Inc. | Atellica CH IgM_2 | - | - | Reason for Correction x D The purpose of this communication is to inform you of an issue with the product indicated in x D Table above and provide instructions on actions that your laboratory must take x D Siemens Healthineers has confirmed the potential for a negative bias with quality control QC and x D patient sample results when using the Atellica CH Immunoglobulin M IgM reagent The x D negative bias was observed after the IgM reagent was stored onboard the analyzer regardless x D of whether the reagent wells were punctured or unpunctured Unopened reagents stored x D refrigerated at C are unaffected See Tables and in the Additional Information section for x D observed results x D This correction is applicable to all future lots until further notice Siemens Healthineers is currently x D investigating the root cause of this issue x D Kindly review the attached UFSN | Actions to be Taken by the Customer Siemens Healthineers recommends batch testing samples for Atellica CH IgM_2 as follows: 1. Remove and discard any Atellica CH IgM_2 reagent packs onboard the analyzer. 2. Load a single fresh Atellica CH IgM_2 reagent pack onto the analyzer. 3. Perform a Lot calibration and process Quality Control (QC). 4. Immediately process a batch of patient samples and conclude with a repeat run of QC. Patient results should not be reported until the QC performed at the end of the batch run has been assessed. o If the QC results are within the established range, patient results can be reported. o If the QC results are not within the established range, do not report patient results and repeat steps 1 – 4 above. 5. Remove and discard the Atellica CH IgM_2 reagent pack at the end of the batch run. Siemens Healthineers does not recommend using the ADVIA IgM_2 reagent on the Atellica CH or Atellica CI Analyzers. Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. If you have received any complaints of illness or adverse events associated with the product listed in Table 1, immediately contact your local Siemens Healthineers Customer Care Center or your local Siemens Healthineers technical support representative. Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached UFSN | 2024-01-22 17:23
Maquet Cardiopulmonary GmbH | 782698 Bubble Sensor | 1 | - | Since a label change in a label template with a Unique Device Identifier UDI was used for the type label attached to the Bubble Sensor BS x L Article no itself The affected labels refers to UDI instead of | Inform all customers possessing the affected products promptly about this Field Action by sending the Field Safety Notice for Customers. | 2024-01-24 09:33
Medtronic, Inc. | Duet External Drainage and Monitoring System, Interlink Injection Sites | 3512 | - | Medtronic received multiple complaints alleging instances of the Duet catheter tubing disconnecting at the Luer connector Complaint trend analysis shows an increased rate of Duet EDMS catheter disconnection beginning October This trend analysis along with a review of product specifications indicates that this catheter disconnection concern is limited to Duet family EDMS products Further investigations to determine root cause of this issue are currently ongoing | Medtronic requests that you immediately take the following actions: • Identify and quarantine any unused impacted product(s). Refer to Appendix A - Affected Products for impacted products. Patient Management Recommendations: • As stated in the Duet External Drainage and Monitoring System Instructions for Use section titled, “System Setup,” check all components for damage and that all connections are secure and leak-free. • If a patient is currently connected to an impacted Duet EDMS and a leak or disconnection is detected, the device should be changed to a new alternative device utilizing a sterile technique. • It is not recommended to remove or replace a Duet system device that is connected to a patient and has been examined and found to be working as intended. • Used product should not be returned to Medtronic and should be disposed of by the healthcare facility in accordance with the healthcare facility’s policies and practices. Unused Product: • Return all unused and non-expired product(s) in your inventory to Medtronic. Your Medtronic Sales Representative can assist in returning any affected inventory, if applicable. • Complete the Customer Acknowledgement Form enclosed with this letter (even if you have no product to return), acknowledging that you have received this information. • This notice should be distributed to all others in your organization who should be aware, or to any organization where the potentially affected devices have been transferred. Please maintain a copy of this notice in your records. | 2024-01-27 17:11
Medtronic, Inc. | Duet External Drainage and Monitoring System, SmartSite Injection Sites, Ventricular Catheter | 3512 | - | Medtronic received multiple complaints alleging instances of the Duet catheter tubing disconnecting at the Luer connector Complaint trend analysis shows an increased rate of Duet EDMS catheter disconnection beginning October This trend analysis along with a review of product specifications indicates that this catheter disconnection concern is limited to Duet family EDMS products Further investigations to determine root cause of this issue are currently ongoing | Medtronic requests that you immediately take the following actions: • Identify and quarantine any unused impacted product(s). Refer to Appendix A - Affected Products for impacted products. Patient Management Recommendations: • As stated in the Duet External Drainage and Monitoring System Instructions for Use section titled, “System Setup,” check all components for damage and that all connections are secure and leak-free. • If a patient is currently connected to an impacted Duet EDMS and a leak or disconnection is detected, the device should be changed to a new alternative device utilizing a sterile technique. • It is not recommended to remove or replace a Duet system device that is connected to a patient and has been examined and found to be working as intended. • Used product should not be returned to Medtronic and should be disposed of by the healthcare facility in accordance with the healthcare facility’s policies and practices. Unused Product: • Return all unused and non-expired product(s) in your inventory to Medtronic. Your Medtronic Sales Representative can assist in returning any affected inventory, if applicable. • Complete the Customer Acknowledgement Form enclosed with this letter (even if you have no product to return), acknowledging that you have received this information. • This notice should be distributed to all others in your organization who should be aware, or to any organization where the potentially affected devices have been transferred. Please maintain a copy of this notice in your records. | 2024-01-27 17:11
Beckman Coulter Ireland, Inc. | DxUm Series Automated Urine Microscopy Analyzer Instrument, DXU 850m Iris | 9 | - | Beckman Coulter has become aware that the cannula which is found in the Flowcell and the iQClear Rinse Waste Well Assembly may rust Rust from the cannula has the potential to generate erroneous results | Follow the actions below: Continue to run a complete Quality Control (QC) rack [including Iris Cleanser, Iris Diluent, iQ Focus and iQ Controls] as directed per iQ200 Series Instructions for Use (IFU) (300-4320 CE or 300-4321 EE) and DxU 850m and 840m Iris IFU (C49320 AB) If the QC passes: o Run patient samples as routine | 2024-01-28 11:31
Beckman Coulter Ireland, Inc. | DxUm Series Automated Urine Microscopy Analyzer Instrument, DXU 850m Iris | 9 | - | Beckman Coulter has become aware that the cannula which is found in the Flowcell and the iQClear Rinse Waste Well Assembly may rust Rust from the cannula has the potential to generate erroneous results | Follow the actions below: Continue to run a complete Quality Control (QC) rack [including Iris Cleanser, Iris Diluent, iQ Focus and iQ Controls] as directed per iQ200 Series Instructions for Use (IFU) (300-4320 CE or 300-4321 EE) and DxU 850m and 840m Iris IFU (C49320 AB) If the QC passes: o Run patient samples as routine | 2024-01-28 12:30
Siemens Healthcare GmbH | Mammomat Revelation | 1 | - | If an examination is interrupted due to any system error while x D the InSpect Projection View PV is active and x D a vacuum biopsy system is used x D the tube arm can only be moved to a position by using the rotation buttons instead of the position that is possible with other Projection Views x D x D The manual movement range allows easy removal of the vacuum biopsy system to release the patient from the unit The limitation to manual movement range may lead to difficulties in removing the vacuum biopsy system | This update delivers a customer safety advisory notice for customers with the option "tomo biopsy" due to a limited movement range of 6˚ for Inspect views. | 2024-01-28 16:55
Covidien llc | Sofsilk | 2 | - | Quality personnel discovered several different suture products product families from Sofsilk Monosof Ti Cron Steel and Surgilon that were re sterilized without approval The discovered lots were re sterilized at gamma doses that exceed the range approved in QD Maximum Sterilization Cycles and permitted dose ranges A total of lots are in the scope of this field action The affected re sterilized lots were released from manufacturing between September and March Based upon available data impact to tensile strength may be delayed but is expected to impact the strength profile of the suture throughout the shelf life of the product Through January Medtronic has received no complaints or reports of serious patient injury related to this issue | • Identify and quarantine all unused and non-expired affected sutures listed in Attachment A of the FSN. See Attachment B of the FSN for guidance on identifying potentially affected devices. • Return all unused product from the affected lot in your inventory to Medtronic. • If purchased from a distributor, contact your distributor directly to arrange for the return of the product back to your distributor. • In addition, please complete and return the Customer Confirmation Form enclosed with the FSN letter, acknowledging that you have received this information. • Pass on the FSN to all those who need to be aware within your organization or to any organization where the potentially affected sutures have been transferred or distributed. | 2024-01-31 08:38
SIEMENS HEALTHINEERS | NXS HOST | 25 | - | SOFTWARE PROBLEM | SOFTWARE UPDATE | 2024-01-31 09:43
siemens healtineers | epoc NXS Host. An in vitro diagnostic medical device (IVD) | 25 | - | software error | software update | 2024-01-31 09:58
Medtronic MiniMed | MiniMed 780G | - | - | A FSN will be delivered to affected HCP s patients and distributors to inform them that the MiniMed G insulin pump software version can now be upgraded from version to version via the Medtronic Diabetes Updater app Upgrading the insulin pump software to version will eliminate the potential occurrence of these errors so that you can safely use your pump for boluses over units at quick bolus speed Instructions for upgrading the pump to version are included with this letter | Upgrade your MiniMed™ 780G insulin pump software to version 6.7 using one of the three options below: • Follow the step-by-step instructions contained in the Software Update Guide attached to this communication. The Software Update Guide can also be found by visiting our website at: https://medtronicemea-my.sharepoint.com/personal/karavd3_medtronic_com/_layouts/15/onedrive.aspx?id=%2Fpersonal%2Fkaravd3%5Fmedtronic%5Fcom%2FDocuments%2FDIB%20MARKETING%20DELIVERIES%2FUploader%20App%20Guides%20%2850577%29&fromShare=true&ga=1 • Contact Helpline at 8002472288if you need assistance with performing the software update. | 2024-01-31 12:17
Yiwu Honggian Technology industry company | - | 896 | تتكون الوصلة من قطعنين قابلة للفك مع طول الاستخدام | - | Product recall | 2024-02-01 11:32
Osatu S.Coop. | - | 1 | - | In certain situations and depending on the operating temperature affected REANIBEX devices may indicate a battery charge level that appears sufficient for defibrillation therapy but when trying to deliver a shock the defibrillator may unexpectedly shut down x D Updating the device software to version or above will change the method for compensating the standby battery consumption and the indicated remaining battery value will be correct | Identification, quarantining and upgrade of the affected devices should be completed as soon as possible, after having received / being aware of the FSN | 2024-02-05 08:58
DiaSorin Inc. | - | 76 | - | DiaSorin has confirmed that the LIAISON Q S E T Device Plus may contain bent or broken wands or be missing the blue seal We are working with our supplier to resolve these issues In the interim DiaSorin has implemented a inspection of all finished devices prior to shipping x D x D A broken or bent wand or missing blue seal may cause an incorrect patient result however this is unlikely due to the obvious nature of the defects and the low defect rate of x D x D The risk to health is considered low as a falsely low or high test result is not likely to cause adverse health consequences Assay results should be used in conjunction with other clinical and laboratory data to assist the clinician in making individual patient management decisions x D | •Please inspect the LIAISON® Q.S.E.T. Device Plus prior to use for a broken or bent wand or missing blue seal and dispose of any defective devices. Replacement product will be provided. To assist you with the inspection of the devices pictures have been provided and are attached to this communication. •Due to low health risk, review of previously reported results is not recommended. | 2024-02-05 11:10
Philips | - | 4 | - | What the problem is and under what circumstances it can occur x D The Philips Azurion System contains a Certeray generator that may fail due to a potential short circuit in the Printed Circuit Board Assembly PCBA in PoInt EVR Power Invertor If a short circuit occurs it will cause the fuses to trip and the system will become non functional due to power loss In this case the source of electrical current will stop preventing further damage Prior to fuses tripping users may experience a burning smell coming from the generator x D | As a remedy, Philips will replace the Power Inverter in the Azurion generator cabinet in all affected Azurion systems through the implementation of Field Change Order (FCO72200544). | 2024-02-05 14:35
Beckman Coulter Inc. | - | 1 | - | Beckman Coulter has determined that the Access Total T reagent lots listed above demonstrate an assay imprecision that is higher than expected x D An unusual amount of precipitate may be observed in the Access Total T reagent packs |  Discontinue using the Access Total T4 reagent lots listed above. Discard all remaining reagent packs from these lots.  Beckman Coulter recommends performing a retrospective review on patient results generated with the affected reagent lots and assess the need for retesting.  Contact your local Beckman Coulter representative for replacement product requests. | 2024-02-06 15:52
NeuMoDx Molecular,Inc. | - | 2 | - | The results impacted by this change are those with HIV RNA viral loads of to IU mL that previously equated to copies mL with the existing conversion factor that now equate to results of to copies mL with the new conversion factor x D HIV management guidelines such as those from European AIDS Clinical Society EACS recommend four HIV viral load tests per year No harm is therefore expected due to these regular viral load monitoring intervals Transmission of HIV to other individuals a result of the existing conversion factor providing a result of copies mL carries a negligible risk Patient guidance such as from the US Centers for Disease Control and Prevention CDC refers to no risk of HIV transmission in individuals with undetectable HIV RNA in plasma the Terrence Higgins Trust as copies mL and aidsmap copies mL respectively x D | Review this notice with your laboratory/medical director. • Important: Forward this information to all individuals and departments within your organization using NeuMoDx HIV-1 Quant Test Strips. If you are not the end user, please forward this notice to the product end user. • Commercial partners: o Forward this notice to your customers. o Follow-up on Acknowledgements of Receipt Form with your customers • Complete “Acknowledgement of Receipt” attached to this letter as soon as possible. • Perform a review of all results obtained since implementation of the assay in your laboratory from individuals with HIV viral loads between 66.7 to 192.31 IU/mL. You should use the new conversion factor to recalculate the results in copies/mL [multiply IU/mL value by 0.75] and review with your HIV physicians as appropriate and identify any results that do not fit the clinical picture in order to determine the need for further evaluation/retesting. • The updated Instructions for Use and ADF will be provided as soon as they are available. Should you wish to update results generated by the system in the interim for copies/mL or log10 copies/mL please follow the guidance: | 2024-02-06 16:06
Ortho Clinical Diagnostics Inc. | - | 95 | - | the Sampling Center mechanism unintentionally disengaged from the UNIVERSAL SAMPLE TRAY TRAY resulting in the misalignment of the intended sample container to be aspirated From our investigation of this single occurrence QuidelOrtho determined that if this disengagement and misalignment occurs it may cause multiple patient samples to be aspirated from an unintended sample container within the same TRAY and report the results using the incorrect patient samples | • To reduce the risk of unintended sample aspiration, refer to the procedure provided in this notification titled, ‘Temporary TRAY Loading Instructions’. • QuidelOrtho recognizes that the ‘Temporary TRAY Loading Instructions’ procedure may not be tenable for all laboratories due to testing volumes. Review your laboratory risk management plan(s) to assess the need for this further mitigation. • Acknowledge your understanding of this notification by completing the enclosed Confirmation of Receipt form no later than February 8, 2024. • Save this notification with your User Documentation or post this notification by each VITROS 3600/5600/XT 7600 System until this issue has been resolved. • If you suspect that your laboratory has experienced the anomaly described in this notification and you have not already done so, please report the occurrence to your local Global Services Organization (formerly Ortho Care). | 2024-02-07 08:35
Illumina Inc. | - | 1 | - | Illumina has released a new software version that remediates the cybersecurity vulnerability and restores remote access to the LRM Web User Interface New versions of the NextSeq Operating Software NOS v and NextSeq Control Software NCS v which both include Local Run Manager v and v respectively are available for installation on the NextSeq Dx x D instruments | Illumina Actions: An Illumina Field Representative will contact the customer to schedule an on-site visit to install the new software. The pertinent local and international regulatory bodies, including the Competent Authorities, have been notified. Required Customer Actions: To complete the Verification Form to acknowledge that the customer has received this notice. And to be emailed the completed form to Illumina Technical Support techsupport@illumina.com within 5 business days. Alternatively, the customer may e-mail Illumina Technical Support techsupport@illumina.com to provide the information requested in the form. | 2024-02-07 09:13
Abbott Laboratories | - | 1022 | - | Please find attached FSN | Please find attached FSN | 2024-02-07 10:59
B. Braun Melsungen AG | - | 7000 | - | Please find attached FSN | Please find attached FSN | 2024-02-07 11:28
Philips Healthcare | - | 10 | - | The PIM Patient Interface Monitor cable delivers the ECG Electrocardiogram signal when performing cardiac scans with the Incisive CT system If the PIM cable s connector is plugged unplugged frequently as part of daily operation the connector and wire may disconnect from the welding point as shown in Figure in the letter x D This disconnection may result in intermittent loss of the ECG signal If the ECG signal is unavailable during a cardiac scan and the obtained image s are not sufficient for clinical diagnosis the operator may decide to rescan the patient x D x D No adverse events have been reported related to this issue as of Sep x D Intended Use x D The Incisive CT is a Computed Tomography X Ray System intended to produce images of the head and body by computer reconstruction of Xray transmission data taken at different angles and planes These devices may include signal analysis and display equipment patient and equipment supports components and accessories The Incisive CT is indicated for head whole body cardiac and vascular X ray Computed Tomography applications in patients of all ages x D x D | Philips will contact you to schedule time for a Field Service Engineer to visit your site to replace the PIM cable with an upgraded version (reference FCO72800782). | 2024-02-07 14:46
Medtronic XOMED, Inc | - | 920 | - | Medtronic has received complaint reports from customers experiencing issues with the NIM Standard Reinforced EMG Endotracheal Tube NIM CONTACT Reinforced EMG Endotracheal Tube during use It was reported that the cuff of the endotracheal EMG tube was over inflated causing it to herniate over the end of the tube leading to loss of ventilation due to airway blockage x D The EMG Endotracheal Tube are silicone tubes with the main shaft reinforced by a wire coil to prevent collapse while maintaining flexibility The cuffs are also manufactured with silicone Not following the Instructions for Use IFU and over inflating the cuff increases intra cuff pressure which can cause the silicone cuff to extend herniate or distort over the end of the tube and or the murphy eye potentially causing obstruction of the patient airway and loss of ventilation x D | Medtronic is asking that Anesthesia Care Providers / Users of these products take the following actions: 1. Review each of the safety notices (May 2022 and February 2024 FSN) with all Anesthesia Care providers / Users of these products. 2. On the enclosed acknowledgement form, a. I have received and read the Urgent Field Safety Notice, dated February 2024, from Medtronic regarding the Medtronic NIM™ Standard Reinforced EMG Endotracheal Tube & NIM CONTACT™ Reinforced EMG Endotracheal Tube follow up communication: i. Notifying me of the IFU updates (and) ii. New training material is being deployed and traced via Medtronic Academy to reinforce the safe usage of the EMG Tubes iii. Reviewed with all Anesthesia Care providers / Users of these products. And have taken the appropriate actions as indicated within the customer letter. 3. Please complete and return the customer acknowledgement form enclosed with this letter to acknowledge receipt of this information. Submit via email and maintain a copy of all records associated with this action. | 2024-02-07 17:29
Applied Medical Resources Corporation | - | - | - | The Epix Universal Clip Applier serves to provide ligation of tubular structure or vessels in laparoscopic and general surgical procedures The nonconformance may result in a clip not loading into the jaws of the device after the trigger is actuated The nonconformance does not impact the functionality of the clip when applied to a structure The clip itself will be properly formed and applied to the structure when the clip is loaded into the jaws of the device The potential hazard is unexpected medical intervention required to prevent bleeding or further aggravation of bleeding vessels in a patient when the clip applier is actuated on a vessel or structure with no clip in the jaws There were no reported events originating from Kingdom of Saudi Arabia associated with the nonconformance of a clip not loading into the jaws x D | Qur records indicate that you have received units from the affected lots. For recall effectiveness, we ask that you please complete the following actions: - Check your inventory for recalled product. - Complete the attached Customer Recall Notification Confirmation Form (Page 3) to acknowledge the recall and indicate if your facility is returning or has already used the lots listed above. ~ Please note that you must return the form even if you have no devices in inventory. - If no product is being returned, please indicate on the Customer Recall Notification Confirmation Form (Page 3). - Complete the attached Distributor Recall Notification Confirmation Form (Page 4) to acknowledge the recall and indicate if your facility is returning or has already used the lots listed above. - If no product is being returned, please indicate on the Distributor Recall Notification Confirmation Form (Page 4). - Please notify any facilities to which you distributed the affected product. Please also complete the Distributor Recall Notification Confirmation Report (Page 5). - Return the Recall Notification Confirmation Form to Applied Medical by emailing to CustomerSatisfaction- MENAT@appliedmedical.com | 2024-02-08 13:20
Beckman Coulter Inc. | - | 20 | - | Beckman Coulter has identified a misalignment between the expiry dates of the individual reagent components and the overall kit expiry date for UIBC REF OSR lot UIBC kits contain two R components R a and R Reagents and two R components R a and R Reagents The R R a and R component expiry dates exceed that of the R a component and the kit label The analyzer will assign the reagent expiry date from R or R Feb instead of correct expiry date from the kit box label Jan | For UIBC (OSR61205) lot 2736, unopened reagents and reagents that are on board the analyzer must be run in a fixed position so that the correct expiry date of 01 Jan 2025 can be manually entered on the analyzer. Please follow the instrument IFU for details on fixed position testing. If replacement kits are required, please contact your local Beckman Coulter representative. | 2024-02-10 23:26
Advanced Sterilization Products Inc. | - | 94 | - | UV lamp power supply that may experience a short circuit causing the power supply to fail x D | The STERRAD Sterilization Systems can continue to be used in accordance with the user’s guide. ASP is implementing a correction to eliminate the possibility of overheating and the potential for burning odor and smoke. | 2024-02-11 14:16
Abbott GmbH. & Co. KG | - | 3 | - | Please find attached FSN letter | Please find attached FSN letter | 2024-02-12 09:40
Philips | - | 20 | - | Philips has become aware of a software login issue associated with specific models of MobileDiagnost wDR systems running on Windows operating system preventing use of the system This issue occurred because the software certificate protecting the Eleva user login expired on May Therefore if you log out or reboot your system you will be prevented from logging in | A Philips representative will contact you to schedule a time for a Field Service Engineer (FSE) to visit your site and implement a software solution to resolve the issue (Reference FCO71200226 for Australia and New Zealand and FCO71200228 for all other impacted countries) | 2024-02-12 14:47
Arrow International LLC | - | 100 | - | Teleflex is initiating this voluntary FSCA for the above mentioned products due to reports received indicating a potentially defective component x D The complaints received relate to resistance of the guidewire handle chamber during use The possible immediate health consequences of the component issue are arterial vasospasm and vessel injury arising from multiple arterial punctures with repeated attempts Any long range consequences would be contingent on the severity of any immediate consequences If tested prior to use as directed by the information for use and good clinical practice the component issue should be discovered Nevertheless Teleflex is recalling the products listed in Appendix x D | 1-If you have affected product, mark the applicable checkbox on the Acknowledgement Form. 2-Cease use immediately and quarantine, the affected product. 3-Teleflex (or your local dealer) will issue a credit note upon receipt of the returned affected product. | 2024-02-13 10:57
Globus Medical INC. | - | 2 | - | Devices may not be rendered sterile due to an internal gap that may be present on some devices causing increased difficulty cleaning prior to steam sterilization | Customers were instructed to immediately return affected cannulas to Globus Medical and complete the letter's Response Card. | 2024-02-13 16:25
Roche Diagnostics GmbH | - | 5093 | - | The root cause investigation showed that the interfering agent is present in the pellet fraction after blood to plasma x D processing Since the issue was not observed with all K K EDTA plasma primary tubes tested it might be caused x D by micro clots even such clots not visible with the naked eye that seem to be predominantly present in specific x D primary tubes x D With this version of this Field Safety Notice we want to announce that the instructions for use IFU of Elecsys x D TnThs have been updated to add one additional centrifugation step for K K EDTA plasma samples to avoid any x D issue in the future The updated IFUs will be released with this notification | Customers using K2/K3 EDTA plasma for TnT hs quantification are required to perform the following additional preanalytical measure: Re-centrifuge K2/K3 EDTA plasma samples in a secondary tube for 5 min at 3’000 x g or 30 sec at 10’000 x g prior to measurement. This action is required and now part of the Instructions for Use. This additional preanalytical measure has been assessed by Roche internally and was proven effective with the samples tested. | 2024-02-14 10:58
Löwenstein Medical Technology GmbH + Co. KG | - | - | - | In rare cases an internal communication problem between the ventilator s controllers may result in ventilation being discontinued without the issue of an alarm From a technical perspective these communication problems may occur in all firmware versions x D available in the market In the market the occurrence was only observed in firmware version Four potentially serious incidents have become known in the market in which this problem has occurred To date L wenstein Medical Technology has not received any reports of harm to patients in this context | Please immediately inform your staff, affected patient, customers and users about the potential risk. You should update all LUISA, TIVAN LS, LifeOne (LM150TD) ventilators to a firmware version 1.9.0007 or higher within 6 months. Please use an additional monitoring system to monitor patients if the therapy of patients dependent on ventilation with the above-mentioned ventilators is to be continued until updating of the device. For example, continuous pulsoximetric (SpO2) monitoring or monitoring of the carbon dioxide level (CO,) on exhalation can be used here as an additional monitoring system. Where an affected ventilator with the above fault picture has failed, the therapy can be commenced again by restarting the device. After restarting, check that the home display is shown. Then hold down the on/off key until the therapy starts again. This process can take up to 30 seconds. After restarting the ventilator, there is no increased probability that the fault picture will re-occur multiple times with the same device. | 2024-02-15 15:01
Biomerieux | - | - | - | bioM rieux has identified a potential signal of increased false positive Norovirus results when using the BIOFIRE FILMARRAY Gastrointestinal GI Panel All unexpired product is potentially impacted | - If a positive Norovirus result is inconsistent with clinical presentation, the positive Norovirus result should be confirmed using another method | 2024-02-15 15:20
Randox Laboratories Ltd. | - | - | - | Randox Laboratories has identified that the Cystatin C Calibrator Series CYS is running with a negative bias compared to other methods We have realigned the target values in all lots within expiry to the IFCC Reference Standard In line with the realignment to the Cystatin C Calibrator Series we are updating the target and range for the Cystatin C controls CYS and CYS | In line with the Cystatin C Calibrator Series, Randox Laboratories is updating the target and range for the Cystatin C controls, CYS5019 and CYS5020. | 2024-02-18 10:57
Biomerieux | - | 13 | - | bioM rieux has identified a potential for false positive seasonal Coronavirus results when using the BIOFIRE FILMARRAY Pneumonia Panel | - Use another method to confirm positive results for seasonal Coronavirus from the BIOFIRE Pneumonia Panel or BIOFIRE Pneumonia Panel plus prior to reporting the result (e.g., consider results from testing a nasopharyngeal swab with a different BIOFIRE respiratory panel such as the BIOFIRE® Respiratory Panel 2.1). Alternatively, the BIOFIRE® Pneumonia Panel or BIOFIRE Pneumonia Panel plus seasonal Coronavirus positive result should not be reported. | 2024-02-18 14:43
Beckman Coulter, Inc. | - | 2447 | - | Beckman Coulter is aware that customers may have received Access Substrate bottles without labels |  Verify upon receipt that all Access substrate bottles have a label.  Discard all Access substrate bottles without labels, according to your local regulations.  Contact your Beckman Coulter representative and request a replacement order if your laboratory receives Access substrate bottles that are missing a label. | 2024-02-18 17:11
Philips | - | 33 | - | The DIVA Inch Widescreen LCD Touch Display is one of the Philips displays available for use with the PIC iX central monitoring station Use of this display allows the user to interact with the surveillance sectors It has been discovered that occassionally the display doesn t function properly and may generate unintended inputs without user interaction along the right and left edges of the screen Although there is a potential for the touch input to impact an area involving alarm acknowledgment the risk of this to happen is very low and the risk of harm is unlikely | A Philips representative will contact you to schedule a visit from a Philips Field Service Engineer who will replace the malfunctioning DIVA 24 Inch Widescreen LCD Touch Display. | 2024-02-19 16:09
Abiomed Inc. | - | 1215 | - | During an internal review Abiomed discovered that information on safe use of Impella x D pumps were issued with two technical bulletins a k a Impella Product Update but the x D product s IFU were not updated to include the same level of detail covered in the bulletins x D and one of the bulletins was not distributed to customers in Europe and the Middle East x D This includes x D Technical bulletin for operator mishandling of the Impella left sided devices resulting in x D iatrogenic ventricular wall perforation x D Impella Product Update for an issue with fibers entrapped in the impeller | -This notice needs to be passed on all those who need to be aware within your organisation or to any organisation where Impella pumps have been transferred. -Please maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action and keep this FSN together with the existing version of the product IFU. | 2024-02-19 23:44
Abiomed Inc | - | - | - | During an internal review Abiomed discovered that information on safe use of Impella x D pumps were issued with two technical bulletins a k a Impella Product Update but the x D product s IFU were not updated to include the same level of detail covered in the bulletins x D and one of the bulletins was not distributed to customers in Europe and the Middle East x D This includes x D Technical bulletin for operator mishandling of the Impella left sided devices resulting in x D iatrogenic ventricular wall perforation x D Impella Product Update for an issue with fibers entrapped in the impeller x D | This notice needs to be passed on to all those who need to be aware within your organization or to any organization where Impella pumps have been transferred. to ensure the effectiveness of the corrective action and keep this FSN together with the existing version of the product IFU. | 2024-02-20 00:18
Abbott | - | 129 | - | planned update to our Instructions for Use associated with observed outflow x D graft deformation known as Extrinsic Outflow Graft Obstruction EOGO associated with the HeartMate Left x D Ventricular Assist System LVAS and HeartMate II LVAS Significant EOGO will manifest clinically as a x D persistent low flow alarm under certain circumstances in some patients and in such cases may impair the ability x D of the HeartMate LVAS to provide adequate hemodynamic support Refer to Appendix A for a complete list of x D impacted product model numbers x D | Abbott will update the Instructions for Use (IFU) to include additional diagnostic recommendations related to persistent low flow and related risks associated with EOGO. | 2024-02-20 21:26
werfen | - | 15084 | - | x D with iQM listed in the attached Appendix GEM Premier X PAK with iQM may experience x D an increase in GEM CVP Level pO failures If the GEM PAK completes warm up and passes x D GEM CVP for pO the GEM PAK may remain in use x D iQM is an active quality process control program designed to provide continuous monitoring x D of the entire testing process and requires the use of GEM CVP Calibration Validation Product x D out of cartridge warm up prior to running patient samples When iQM disables or flags a x D parameter the program is ensuring there is no potential for patient harm x D x D Internal testing identified that the GEM PAKs cartridges for the GEM Premier with iQM listed in the attached Appendix GEM Premier PAK with iQM may experience an increase in GEM CVP Level pO failures If the PAK completes warm up and passes GEM CVP for pO the GEM PAK may remain in use x D iQM is an active quality process control program designed to provide continuous monitoring of the entire testing process and requires the use of CVP Calibration Validation Product out of cartridge warm up prior to running patient samples When iQM disables or flags a parameter the program is ensuring there is no potential for patient harm x D x D Internal testing identified that the GEM PAKs cartridges for the GEM Premier with iQM listed in the attached Appendix GEM Premier PAK with iQM may experience an increase in APV Auto PAK Validation pO failures If the PAK completes warm up for pO the GEM PAK may remain in use x D iQM is an active quality process control program designed to provide continuous monitoring of the entire testing process and requires the use of APV Auto PAK Validation out of cartridge warm up prior to running patient samples When iQM disables or flags a parameter the program is ensuring there is no potential for patient harm | Please take the following immediate actions:1-  Ensure the iQM is enabled for all the listed PAKs and GEM CVPs are run according to Instructions for Use.  Continue to use the GEM PAKs (cartridges) that passes GEM CVP out of warm-up for pO2.  Contact your local Werfen representative for a replacement GEM PAK, if GEM CVP does not pass pO2.  Share this information with your laboratory staff and follow your internal procedures.  Forward this notification to all affected locations within your facility.  Retain a copy of this notification for your records. 2- • Continue to use the GEM PAKs (cartridges) that passes GEM CVP out of warm-up for pO2. • Contact your local Werfen Representative for a replacement GEM PAK, if GEM CVP does not pass pO2. • Share this information with your laboratory staff and follow your internal procedures. • Forward this notification to all affected locations within your facility. • Retain a copy of this notification for your records. 3- • Continue to use GEM PAKs (cartridges) that pass APV out of warm-up for pO2. • Contact your local representative for a replacement GEM PAK, if needed. • Share this information with your laboratory staff and follow your internal procedures. • Forward this notification to all affected locations within your facility. • Retain a copy of this notification for your records. | 2024-02-21 14:00
Gyrus ACMI, Inc. | - | - | - | Areview of customer complaints and adverse events we have become aware that some clinicians are defaulting to the preset laser values for all clinical procedures when that is not consistent with the Instructions for Use x D The Soltive Laser is intended for incision excision resection ablation coagulation hemostasis and x D vaporization of soft tissue with or without an endoscope in urology lithotripsy gastroenterological x D surgery and gynecological surgery | 1. Carefully read the content of this Field Safety Notice. 2. Ensure all personnel, including clinical staff, are completely knowledgeable and thoroughly aware of the contents of this letter and the Instructions for Use. 3. Indicate on the Reply Form that you have received and understood this Field Safety Notice by filling out and returning the completed enclosed Reply Form back to your local Olympus representative latest by 04.03.2024. | 2024-02-23 15:10
Biomerieux | - | - | - | The cause for this risk is the presence of an increased level of DNA fragments from non viable Candida tropicalis targets in BD BACTEC blood culture vials Table The presence of DNA fragments does not compromise the intended function of the blood culture vials culturing viable microorganisms However the BIOFIRE BCID Panel detects nucleic acid from viable and non viable organisms alike A false positive result incorrect ID could lead to an inappropriate change in therapy The patient may remain on inappropriate therapy until the Candida tropicalis is confirmed or not | - If the BIOFIRE BCID2 Panel is used to test BD BACTEC™ blood culture vials (examples in Table 1), positive results for Candida tropicalis should be confirmed by another method prior to reporting the test results. - Distribute this information to all appropriate personnel in your laboratory, retain a copy in your files, post this letter in or near the laboratory, and forward this information to all parties that may use this product, including others to whom you may have transferred our product. | 2024-02-25 15:35
Siemens Medical Solution USA, Inc | - | 60 | - | The Cardiac DICOM SR viewer will NOT display the MINIMUM Min or MAXIMUM Max value and will instead display the LAST measured value under certain condition described in the attachment In that case the displayed measurement in the Cardiac DICOM SR viewer may NOT be the Minimum Min or MAXIMUM Max value measured during the examination x D | Please note to avoid potential misdiagnosis by an over- or underrepresented measured value, it is recommended to refer to the measurement values displayed on the ACUSON Juniper Ultrasound System Final Report when the configuration is to display the Min or Max value. It is recommended by Siemens Healthineers a review of previous cardiac ultrasound examination results where the clinical assessment was reviewed using a Cardiac DICOM SR viewer. This issue will be corrected by Siemens Healthineers with a software update, refer to attachment for more info. | 2024-02-26 14:01
UTAK Laboratories, Inc. | - | 51 | - | Incorrect product labeling on same vials of part drugs of abuse level lot number D some of the vials in the dugs abuse level box were mislabeled as part dugs of abuse level | Step 1: Identify which vials need to be relabeled. Remove from storage. The product can be stored in ambient temperature for up to 1 hour. Step 2: Confirm the lot number on the vials pulled and updated labels provided by UTAK is D3228. The only difference between the labels should be the part number (98818/98819) and level (Level 1/Level 2). Step 3: Peel the new label from the label sheet provided by UTAK and place the new label directly over the current label. The labels are the same size and will cover it completely. Ensure the old label is completely covered by the new label. Step 4: Return vials to their appropriate storage location and continue to use the product as planned. | 2024-02-27 11:07
Abbott Laboratories Diagnostics Division | - | 12 | - | Please find attched FSN letter | Please find attched FSN letter | 2024-02-28 10:13
Randox Laboratories | - | - | - | Randox Laboratories has identified that the Microalbumin Calibrator Series MA MA is running with a negative bias compared to other methods We have reassigned the target values in the above lot numbers in line with the ERM reference material Please refer to the table below for the updated calibrator targets You may experience a shift in Quality Control and patient sample recovery of up to | If a shift in Quality Control is experienced and patient sample recovery of up to +12%. Please discard all copies of the calibrator Instructions For Use (IFU)s and download the updated IFUs from www.randox.com. If further information is required, please contact technical.services@randox.com. | 2024-02-28 14:52
Illumina, Inc. | - | 5 | - | Illumina has released new software versions that remediates the cybersecurity vulnerability and x D restores remote access to the LRM Web User Interface New versions of the MiSeq Operating Software MOS v MiSeq Control Software MCS v which includes Local Run Manager v are available for installation on the MiSeqDx instruments | Illumina Actions: An Illumina Field Representative will contact you to schedule an on-site visit to install the new software. The pertinent local and international regulatory bodies, including the Competent Authorities, have been notified. Required Customer Actions: Please complete the Verification Form to acknowledge that you have received this notice. We kindly ask that you email the completed form to Illumina Technical Support techsupport@illumina.com within 5 business days. Alternatively, you may e-mail Illumina Technical Support techsupport@illumina.com to provide the information requested in the form. | 2024-02-29 09:17
ORIGIO a/s | - | - | - | CooperSurgical has internally become aware of a non conformity of Origo Sperm Wash Media related to the low concentration of sodium pyruvate for the aforementioned lots No complaints have been received to date related to the above listed part number lots | 1) Inspect your inventory, quarantine any affected product, and immediately stop using the product. 2) Complete the appropriate version of the attached form (Customer Acknowledgement Form or Distributor Acknowledgement Form). Once completed, return the form to CooperSurgical as indicated at the top of the form. Note: Even if you do not have any affected product in your inventory (or have not distributed any affected product to customers), please complete and return the enclosed form so that we may document receipt of this FSN. When the completed form is received by CooperSurgical, arrangements will be made for the return of any affected product at no additional cost to you. Credit will be applied to your account for product returned under this action. CooperSurgical is committed to high-quality, safe, and effective products. This corrective action has been initiated to ensure this failure mode does not reoccur and future potential patient harms can be avoided. | 2024-02-29 14:12
Mirion Technologies (Capintec), Inc. | - | 7 | - | Spring arm failure with Captus e Thyroid Uptake Systems can cause the collimator to fall downward to its lowest point of travel causing potential injury with a patient or operator x D | 1. Review the attachment and ensure that all affected personnel, including all operators of Captus 4000e Thyroid Uptake Systems, are aware of the contents. 2. Instruct affected personnel and operators to comply with the steps in the attachment in the interest of patient and operator safety. | 2024-02-29 17:28
Philips Medical Systems | - | 2 | - | It was found that the Care Assist mobile application version and earlier would not play custom ringtones when iOS device screens were locked with a passcode The issue occurred on devices running iOS and as well as multiple versions of iPhones The Care Assist mobile application is not intended to replace the patient s primary alarming system If users rely on the Care Assist application as their primary event notification alarming system the failure of custom ringtones could result in a delay in treatment for patients experiencing a change in clinical condition | 1. The action by Philips will be to send each known affected customer an URGENT Field Safety Notice. 2. The Care Assist version 4.1.2 mobile application will be made available immediately in the Apple App Store. | 2024-03-01 15:20
SAMSUNG ELECTRONICS Co., Ltd. | - | 3 | - | Mobile x ray systems moving arm frame welding issue | Samsung Electronics will carry out the field safety corrective action for all affected systems produced between October 2016 to April 2018. The corrective action is to fasten the arm with a high tensile bolt to eliminate the risk of the moving arm falling and preventing any potential risk of injury. Service engineers will correct all affected systems free of charge and contact you to arrange for the correction | 2024-03-03 19:46
Ortho Clinical Diagnostics Inc. | - | - | - | the Sampling Center mechanism unintentionally disengaged from the UNIVERSAL SAMPLE TRAY TRAY resulting in the misalignment of the intended sample container to be aspirated From our investigation of this single occurrence QuidelOrtho determined that if this disengagement and misalignment occurs it may cause multiple patient samples to be aspirated from an unintended sample container within the same TRAY and report the results using the incorrect patient samples | • To reduce the risk of unintended sample aspiration, refer to the procedure provided in this notification titled, ‘Temporary TRAY Loading Instructions’. • QuidelOrtho recognizes that the ‘Temporary TRAY Loading Instructions’ procedure may not be tenable for all laboratories due to testing volumes. Review your laboratory risk management plan(s) to assess the need for this further mitigation. • Acknowledge your understanding of this notification by completing the enclosed Confirmation of Receipt form no later than February 8, 2024. • Save this notification with your User Documentation or post this notification by each VITROS 3600/5600/XT 7600 System until this issue has been resolved. • If you suspect that your laboratory has experienced the anomaly described in this notification and you have not already done so, please report the occurrence to your local Global Services Organization (formerly Ortho Care). | 2024-03-04 11:02
Ortho Clinical Diagnostics Inc. | - | - | - | the Sampling Center mechanism unintentionally disengaged from the UNIVERSAL SAMPLE TRAY TRAY resulting in the misalignment of the intended sample container to be aspirated From our investigation of this single occurrence QuidelOrtho determined that if this disengagement and misalignment occurs it may cause multiple patient samples to be aspirated from an unintended sample container within the same TRAY and report the results using the incorrect patient samples | • To reduce the risk of unintended sample aspiration, refer to the procedure provided in this notification titled, ‘Temporary TRAY Loading Instructions’. • QuidelOrtho recognizes that the ‘Temporary TRAY Loading Instructions’ procedure may not be tenable for all laboratories due to testing volumes. Review your laboratory risk management plan(s) to assess the need for this further mitigation. • Acknowledge your understanding of this notification by completing the enclosed Confirmation of Receipt form no later than February 8, 2024. • Save this notification with your User Documentation or post this notification by each VITROS 3600/5600/XT 7600 System until this issue has been resolved. • If you suspect that your laboratory has experienced the anomaly described in this notification and you have not already done so, please report the occurrence to your local Global Services Organization (formerly Ortho Care). | 2024-03-04 11:10
DiaSorin Italia S.p.A. | - | 5 | - | Low reactivity of the Positive Control lot belonging to the control lot resulting in a value close to or below the lower label limit | ●Please stop using the affected lot. A following lot not affected by the same issue is available and can be ordered. DiaSorin will provide a credit note for the unused kits. ●No retesting of patient's samples is required. | 2024-03-04 17:23
Covidien llc | - | 5985 | - | Medtronic manufacturing site has received field complaints for the product Auto Suture Structural Balloon Trocar Three unsealed Structural Balloon Trocar instruments Model OMST SB Lot Numbers P A P J and P J were received at Medtronic Engineering for further evaluation of Seal damage component disengaged into the cavity Seal Disengaged and Seal Damaged The visual inspection found that the Main seal component had disengaged from its valve body assembly also traces of silicone adhesive were found around the circumference of the Seal | Advice on actions to be taken by the distributor and the user: • Inform all surgeons and clinicians who handle the preparation and/ or placement of a Mesh Device that utilize the balloon and blunt tip trocar devices. • Prior to using any Mesh device in conjunction with the following Trocars: Covidien Auto Suture™ Structural Balloon Trocar and Auto Suture™ Blunt Tip Trocar, please carefully review and adhere to the mesh manufacturer’s Instructions for Use (IFU) on proper insertion techniques. • Please complete and return the customer acknowledgement form enclosed with this letter acknowledging receipt of this information. • Please transfer this notice to other organizations on which this action has an impact and maintain a copy of this notice in your records. | 2024-03-04 18:34
OrthoPediatrics Corp,. | - | 17 | - | OrthoPediatrics received nine complaints in relation to items without manufacturer logos imprinted After investigation OrthoPediatrics has identified these wrenches may corrode after multiple sterilisation cycles | The affected product must be returned to Orthopediatrics. Replacement product will be provided for device part numbers 910-CW-10L, 915-CW-10 and C915-CW-07. Part Number 915-FW-10 will no longer be available. Customers are to contact customer service to arrange return of the affected product. No additional corrective actions need to be taken for product already implanted. | 2024-03-05 14:25
Cristalens Industrie | - | 87 | - | Cristalens Industrie has been informed of cases of intraocular lenses with unexpected refractive errors x D As a precaution Cristalens Industrie is recalling the batches potentially involved x D Users of batches concerned by the recall have already been informed x D Hazard description Risk of unsatisfactory post operative refractive results | You will find the references to the intraocular lenses concerned by this notice. Please follow the instructions below: 1- Cristalens & RMS has identified that affected devices have been provided to your establishment. Check your stock management systems and your customers' warehouses: Please do not use intraocular lenses with the serial numbers shown in Appendix 1. Please return intraocular lenses with the serial numbers shown in Appendix 1. Please provide us with the references of the intraocular lenses with the serial numbers shown in Appendix 1 that have been fitted and the date of implantation. 2- Immediately transmit this notice to all persons who should be informed of this notice in your establishment, and/or to any other establishment where the affected products have been transferred. 3- Any correspondence or discussions relating to this notice must mention the following reference: FSN_RC_00641 We inform you that this notice has been communicated to the competent national authorities. Cristalens Industrie & RMS are committed to providing you and your customers with the highest quality products, thank you for your cooperation and is fully aware of the inconvenience caused. | 2024-03-05 17:48
QIAGEN | - | 3 | - | An issue has been identified with the EZ Connect MDx instrument This is caused by insufficient welding between the heater cable and the heater plate on the heater system The insufficient welding leads to the heater cable becoming brittle and a consequent disintegration of the cable Failure of the cable if not spotted early by a field service technician during a service visit may cause the instrument to present with following error x D ERROR HEATER UNABLE TO REACH TEMPERATURE x D If the instrument presents with Error during a run there is a risk that the run will stop and the sample being processed will be lost In this event QIAGEN recommends that the instrument should not be used again until a field service engineer has repaired the unit x D x D x D x D | Actions to be taken by the Customer/User • Complete the Acknowledgment of Receipt promptly and return to quality.communications@qiagen.com • Should the error described above occurs, please perform the following actions: o Switch off the instrument and disconnect from the power supply. o Remove any onboard samples and perform decontamination procedure of the instruments as described in the user manual. o Call your local QIAGEN Technical Service or the Commercial Partner responsible for QIAGEN products in your country. Actions to be taken by the Commercial Partner/Distributor • Forward this notice to your customers. • Follow up on the Acknowledgment of Receipt with all your customers. • Confirm the completion of the follow up of the Acknowledgement of Receipt of your customers to quality.communications@qiagen.com. | 2024-03-06 09:07
Respironics, Inc. | - | 244 | - | What the problem is and under what circumstances it can occur x D All Trilogy Evo Trilogy Evo O and Trilogy EV devices are susceptible to this problem x D x D A software algorithm that calculates remaining battery life can malfunction and cause the device to either x D A Issue a Loss of Power alarm that stops CPAP or PSV therapy while operating on battery power alone x D B Issue a Battery Depleted alarm while continuing therapy if plugged into a permanent power source such as AC or DC power x D x D This can only happen if all of the following conditions are met x D The device is operating in CPAP or PSV mode x D The device is not able to detect the respiratory effort of the patient for at least ten minutes and fifty five seconds x D Patients most vulnerable to this issue include neonatal and pediatric patients patients recently removed from anesthesia or other patients with low inspiratory effort due to their potential for minimally detectable respiratory effort x D x D This malfunction is due to a software algorithm calculation error and is not a malfunction of the internal or detachable batteries An alarm which can stop therapy may occur even if there is sufficient battery life remaining x D x D This malfunction will not happen in ventilation modes other than CPAP and PSV x D x D | Philips Respironics is developing a software correction that will remedy this issue. It will be released for download by all device users in the second quarter of 2024. When available, you will receive an additional notification advising how to access and install the software on your devices. | 2024-03-06 15:02
Vyaire Medical, Inc. | - | 12 | - | Not provided | Onsite desturction | 2024-03-07 14:52
AGFA NV | - | 1 | - | Description has been provided by component supplier x D It is related to the steel cable installed inside the mobile column which supports the weight of the telescopic arm with the tube head x D The potential risk consists of an interference which could cause that the safety system which blocks the arm if the cable is cut off does not work properly x D The interference occurs between the steel cable and a metal plate of the arm s carriage due to the shape of the cable s end ref figure x D | Agfa/local business partner will inspect the cable of your device at the earliest opportunity and will replace it in case of any potential malfunction of the safety system. | 2024-03-07 16:30
Draeger Medical, Inc. | - | 85 | - | Damage to the battery can lead to a reduction in capacity and thus to a reduced runtime in battery mode | Solution: 1) Send the safety notice to all customers as per the location list. 2) Send a supplement for the instructions for use of Perseus together with the „customer letter accompanying the supplement“. (You will then be informed by Peter Küpper by email. The location list will be updated). Note! If a battery fails during testing, it must be replaced free of charge. Caution: Batteries that are older than 24 months are subject to a charge to be paid by the customer. | 2024-03-12 14:56
Masimo Corporation | - | 164 | - | Masimo identified certain Rad G devices powering off and on without pressing the power button Masimo s investigation identified an issue that can result in an unintentional change in the power state | • Please assess your inventory to confirm if you have Rad-G devices listed in Attachment 2-part number(s) and serial number(s). • If you identify affected Rad-G® devices, please promptly remove them from use and follow the steps below. • Contact Masimo Technical Services to request return of the affected part and serial number (s) for repair or replacement. | 2024-03-12 16:02
Synthes GmbH | - | 2 | - | The labels for two lots of CERVIOS chronOS Wedge Shaped Cage have incorrect expiry dates of August Lot L and October Lot L The actual expiry date is October for both lots To date no complaints have been received regarding the expiry date | The DPS QRB decided to initiate a field action (removal) for two (2) lots of class III products that were distributed without an incorrect expiration date (1 year longer than expected). In Saudi Market, the quantity is: 870.923S (Q:2) FSA ref: 2355156_2358304 Note: 870.925S is not in Scope for SA market. | 2024-03-12 22:40
Ortho Clinical Diagnostics Inc. | - | 529 | - | Ortho Clinical Diagnostics QuidelOrtho received complaints regarding mis matched x D diluent lots assembled into VITROS Performance Verifier Lots Q and R x D QuidelOrtho investigated the issue and confirmed that some sales units of VITROS x D Performance Verifier Lots Q and R contained the correct lyophilate but incorrect x D diluent | • To ensure your laboratory is using the appropriate diluent for reconstitution, inspect your inventory of VITROS Performance Verifier, Lots Q1174 and R1176. • If mis-matched diluent Lots M9815 or P9967 are found, render unusable and discard all sales units of VITROS Performance Verifier, Lots Q1174 and R1176. • Complete the enclosed Confirmation of Receipt form no later than March 22, 2024. Upon receipt of your completed Confirmation of Receipt form, QuidelOrtho will provide credit for, or replacement of, discarded product. • Review your laboratory’s ECO2 and Na+ QC results generated using VITROS Performance Verifier, Lots Q1174 and R1176, for any unexpected QC shift. • Please forward this notification if the affected product was distributed outside of your facility. • Save this notification with your User Documentation or Post this notification near your laboratory’s inventory storage area until an inspection of VITROS Performance Verifier, Lots Q1174 and R1176 has been performed. | 2024-03-13 09:55
FUJIFILM Corporation | - | - | - | Whenever the key switch is turned on the exposure conditions are set to the default values kV mAs approx mGy SID m which are intended for adult chest imaging Before making an exposure make sure that intended exposure conditions are set | we are providing the information in this letter which is in addition to the information contained in our operation manual. | 2024-03-13 13:34
Synthes GmbH | - | 2 | - | The subject products are being removed from the field because lot H was not sterilized | Synthes GmbH is initiating a field safety notice (removal) for selected lots of TFNA Femoral Nails. The TFNA system is intended for temporary fixation and stabilization of proximal femur fractures. FSA ref: 2348044 | 2024-03-13 23:20
Toshiba Medical Systems Corporation | - | 2 | - | TUS VC has a function when inputted a patient iD automatic y end the current study and automatically enter the next patients information from the stored study history in the unit to start a new study The function may not run properiy depending on the conditions and multiple patients study data may be mixed up in a study | 1) (CMSC) Will issue an FSI (customer notification) describing the phenomenon and method of prevention applied by the user before corrective measures are taken. 2) (CMSC) Will issue an FMi to fix the problem | 2024-03-14 03:57
DiaMed GmbH | - | 4103 | - | The purpose of this letter is to inform you about a quality issue we are facing with Bio Rad x D reagent intended for antibody screening x D Reason for the Field Safety Notice x D We have been able to confirm an increased level of complaints and adverse events related x D to weak non specific reactions when using ID System screening reagents in Indirect x D Antiglobulin Testing IAT for QC donor and patient samples x D Customers using both the manual and instrument methods reported weak positive reaction x D results identified as wR and sometimes on instruments instead of an x D expected clearly negative reaction in IAT x D These doubtful reactions or in some cases weak positive reactions may occur randomly x D among batches x D The investigations so far point out that a possible enhancer of the phenomenon is the x D combination of certain lots of reagents red blood cells with certain lots of associated IDCards x D Kindly review the attached UFSN | Action(s) to be taken by the Customer: The capacity of impacted reagents to detect clinically significant antibodies is not affected by the issue described above. For that reason, you may continue to use these products for their intended purpose. In case you would experience non-specific reaction impacting your ability to render results, Bio-Rad is requesting that customers affected by this notice take the following action: 1. If available in your laboratory, repeat the test with another lot of ID-Cards If the issue persists or if you do not have other lot of ID-Cards available, 2. Switch to another lot of cells that you have received within your standing order. We thank you to continue reporting any issues to our support team, where the team will liaise with the laboratory to collect the relevant information. This may include:  The Daily Journal with images  Information on the Instrument used or manual testing  Information on the reagents lots used (ID-Cards and ID-Cells)  The frequency of “wR”'s or weak reactions observed versus the total amount of screening tests per day  IH-QC samples impacted  Patient/Donor samples impacted Please ensure this notice is passed to all those who need to be aware within your organization or to any organization where the impacted devices have been transferred. Please complete and return the attached response form as soon as possible so that we are assured you have received this important communication. Kindly review the attached UFSN | 2024-03-14 10:55
Philips Medical Systems Nederland B.V. | - | 69 | - | The patient support floor plates for MR systems may not have been anchored to the floor according to Philips requirements causing the patient support to become unstable with the potential for tipping see Figure This movement can inadvertently lead to harm to the patient and or operator under certain circumstances x D x D There have been complaints reported to Philips associated with the patient support floor plate issue for MR systems and no reports of harm or injury as of February x D | Philips is providing this Field Service Notice (FSN) Letter which contains recommendations for continued use of the systems referenced in Section 4. A Philips representative will contact you to schedule time for an FSE to visit your site to perform an inspection regarding the patient support stability. (Reference FCO 78100570). | 2024-03-15 11:53
Philips | - | 69 | - | The patient support floor plates for MR systems may not have been anchored to the floor according to Philips requirements causing the patient support to become unstable with the potential for tipping see Figure This movement can inadvertently lead to harm to the patient and or operator under certain circumstances x D x D There have been complaints reported to Philips associated with the patient support floor plate issue for MR systems and no reports of harm or injury as of February x D | Philips is providing this Field Service Notice (FSN) Letter which contains recommendations for continued use of the systems referenced in Section 4. A Philips representative will contact you to schedule time for an FSE to visit your site to perform an inspection regarding the patient support stability. (Reference FCO 78100570). | 2024-03-15 11:53
Maquet Cardiopulmonary GmbH | - | 1 | - | During retrospective internal testing MCP Getinge discovered narrow channels in the seal of the sterile bags used to package the RF MCP Getinge receives this sterile bag from a supplier that performs tests to verify sterility Incomplete sealing can potentially lead to a breach of the sterile barrier x D To gauge the scope of affected products MCP Getinge re tested all lots of the supplier provided packaging used for products that are still within shelf life This testing demonstrated that three supplier batches of sterile bags were affected | Product Removal Inform all customers possessing the affected products promptly about this Field Action by sending the Field Safety Notice for Customers. • Arrange the return of the affected product and provide the customer with credit note. • Implementing additional corrective and preventive measures with the supplier in order to prevent the issue from happening again. • Verification testing for integrity of the sterile seal is now performed by MCP/Getinge in-house before the material is used in production. | 2024-03-17 14:41
Maquet Cardiopulmonary GmbH | - | 134 | - | During retrospective internal testing of HLS Set packaging MCP Getinge discovered narrow channels in the seal of the sterile bags used to package the Emergency Priming Line Figure Item MCP Getinge receives this sterile bag from a supplier that performs tests to verify sterility Incomplete sealing can potentially lead to a breach of the sterile barrier x D The Emergency Priming Line is first packaged in the above described sterile bag which is itself packaged in the sterile Intellipack tray with the other HLS Set components Thus even with the potentially compromised sterile bag the Emergency Priming Line is sterile when an HLS Set is first opened Therefore Emergency Priming Lines affected by this Field Safety Notice may only be used immediately after opening an HLS Set or else disposed of if not needed If an Emergency Priming Line is needed at a later time a new HLS Set needs to be opened to obtain a sterile Emergency Priming Line x D To gauge the scope of affected products MCP Getinge re tested all lots of the supplier provided packaging used for products that are still within shelf life This testing demonstrated that three supplier batches of sterile bags were affected | Inform all customers possessing the affected products promptly about this Field Action by sending the Field Safety Notice. • Depending on the feedback from the customer: o If a customer uses the Emergency Priming Line from a backup HLS Set: • Provide the customer with a replacement HLS Set or credit note o If a customer has one leftover affected product with no backup product: • Provide the customer with a replacement HLS Set or arrange the return of the affected product • Getinge is implementing additional corrective and preventive measures with the supplier in order to prevent the issue from happening again. • Verification testing for integrity of the sterile seal is now performed by MCP/Getinge in-house before the material is used in production. | 2024-03-17 14:47
Gyrus ACMI, Inc. | - | 6 | - | Olympus is conducting this field corrective action after learning internally that we may have supplied non compliant power cords for use with the Soltive Laser | Olympus representative will conduct an inspection of your Soltive Laser power cord and replace any non-compliant power cords at no charge to you. | 2024-03-18 05:04
Randox Laboratories Ltd | - | - | - | Software update | • Review the software release notes provided for further information. • Please complete the mandatory software upgrade. Please contact technical.services@randox.com or support@randoxtoxicology.com or support@randoxfood.com for the OneDrive link to the software update files. • Complete and return the response form 12187-QA technical.services@randox.com or support@randoxtoxicology.com or support@randoxfood.com within five working days. | 2024-03-20 11:25
Spineart SA | - | 0 | - | The entry point on the thread inside the screw head was not machined correctly | Please refer to the attachment for actions that should be taken. (Reference (FSN) FSCA-01-2024). | 2024-03-20 13:17
Canon Medical Systems Corporation | - | 2 | - | It has been found that some of the maximum SFG values displayed in System x D Information of the operation window and some of the maximum SFG values described in x D the safety manual are lower than the correct values | modified software will be installed and safety manuals will be replaced with manuals with correct values. | 2024-03-20 22:34
Abbott | - | - | - | Abbott s investigation determined that in certain instances routine x D manipulation of the pump or internal fluid pressures during implantation can result in a compressed x D sealing ring on one side leading to a leak path on the opposite side | The product is not being removed from the field and unused product does not need to be returned. Abbott considers the medical benefits to outweigh the risk of harm and recommends continued use of the HeartMate 3 TM LVAS per the Instructions For Use (IFU) and supplemental guidance provided below. In most cases, blood leakage from, or air entrainment into, the left ventricle or LVAD at any interface, including at the interface between the pump and the apical cuff, is addressed by adjusting the pump position. In the remaining cases, leakage or air entrainment is addressed using conventional strategies for resolving air leaks or surgical bleeding. If blood leak or air entrainment is suspected or observed, follow standard surgical processes and the existing IFU: • Residual air must be completely evacuated from the device blood chamber prior to initiating the LVAD support • Ensure that bleeding is assessed and ensure proper management of hemostasis before closing all wounds. • Use conventional strategies for resolving air leaks or surgical bleeding, including: adjusting the pump position, waiting for the natural tendency of blood to coagulate or upon reversal of anticoagulation, adding surgical materials, and exchanging the apical cuff, the pump, or both. • Always have a complete backup system (implant kit and external components) available on-site and in close proximity during the implantation procedure for use in the event of an emergency | 2024-03-21 12:18
Applied Medical Resources Corporation | - | - | - | The Epix Universal Clip Applier serves to provide ligation of tubular structure or vessels in laparoscopic and general surgical procedures The nonconformance may result in a clip not loading into the jaws of the device after the trigger is actuated The nonconformance does not impact the functionality of the clip when applied to a structure The clip itself will be properly formed and applied to the structure when the clip is loaded into the jaws of the device The potential hazard is unexpected medical intervention required to prevent bleeding or further aggravation of bleeding vessels in a patient when the clip applier is actuated on a vessel or structure with no clip in the jaws There were no reported events originating from Kingdom of Saudi Arabia associated with the nonconformance of a clip not loading into the jaws | Qur records indicate that you have received units from the affected lots. For recall effectiveness, we ask that you please complete the following actions: - Check your inventory for recalled product. - Complete the attached Customer Recall Notification Confirmation Form (Page 3) to acknowledge the recall and indicate if your facility is returning or has already used the lots listed above. ~ Please note that you must return the form even if you have no devices in inventory. - If no product is being returned, please indicate on the Customer Recall Notification Confirmation Form (Page 3). - Complete the attached Distributor Recall Notification Confirmation Form (Page 4) to acknowledge the recall and indicate if your facility is returning or has already used the lots listed above. - If no product is being returned, please indicate on the Distributor Recall Notification Confirmation Form (Page 4). - Please notify any facilities to which you distributed the affected product. Please also complete the Distributor Recall Notification Confirmation Report (Page 5). - Return the Recall Notification Confirmation Form to Applied Medical by emailing to Customer Satisfaction- MENAT@appliedmedical.com | 2024-03-21 13:23
Randox Laboratories Ltd | - | - | - | Randox Laboratories has identified that Inorganic Phosphate in Calibration Serum Level x D CAL is running with a negative bias on RX Series instruments compared to other x D methods We have reassigned the target values in the above lot numbers in line with our x D internal master calibrator lot Please refer to the table below for the updated calibrator x D targets You may experience a shift in Quality Control and patient sample recovery up to x D Please discard all copies of the calibrator IFU and download the updated IFUs from x D www randox com Quality Control targets are also being updated in line with the x D restandardisation and updated IFUs can be accessed via www randox com If further x D information is required please contact technical services randox com | * Discuss the contents of this notice with your Medical Director if you have used the RX Series targets for Inorganic Phosphorous in the aforementioned lots. *Complete and return the response form, 12187-QA to technical.services@randox.com within five working days. *Please discard all copies of the IFUs and download the latest versions from www.randox.com. | 2024-03-21 14:24
Abbott Molecular Division Inc. | - | 9 | - | PLEASE FIND ATTACHED FSN LETTER | PLEASE FIND ATTACHED FSN LETTER | 2024-03-21 14:32
Respironics, Inc. | - | 69 | - | The OmniLab Advanced OLA devices feature a Ventilator Inoperative alarm which occurs when the ventilator detects an internal error or a condition that may affect therapy This may manifest in the following ways x D The device may reboot intermittently for seconds stops providing therapy screen goes blank during the reboot and there is a single audible alert restarting therapy and returning to delivering therapy with same patient settings x D OR x D The device may reboot intermittently for seconds stops providing therapy screen goes blank during the reboot and there is a single audible alert restarting therapy and returning to delivering therapy but with factory default settings x D OR x D When there are three reboots within a hour period the device will enter a Ventilator Inoperative state therapy stopped audible and visual alarms present x D OR x D The device may enter a Ventilator Inoperative state without a reboot preceding this condition x D | Philips Respironics is currently investigating this issue. Philips Respironics will be in contact as soon as additional appropriate actions have been determined. As OmniLab Advanced + (OLA+) devices are included in the Sound Abatement Foam Field Action, the existing actions determined within the Sound Abatement Foam Remediation Update are applicable, they are as follows: • Continue to wait for additional remediation options to become available. | 2024-03-21 15:37
Respironics, Inc. | - | 122 | - | The affected products feature a Ventilator Inoperative alarm which occurs when the ventilator detects an internal error or a condition that may affect therapy This may manifest in the following ways x D x D The device may reboot intermittently for seconds stops providing therapy screen goes blank during the reboot and there is a single audible alert restarting therapy and returning to delivering therapy with same patient settings x D OR x D The device may reboot intermittently for seconds stops providing therapy screen goes blank during the reboot and there is a single audible alert restarting therapy and returning to delivering therapy but with factory default settings x D OR x D When there are three reboots within a hour period the device will enter a Ventilator Inoperative state therapy stopped audible and visual alarms present x D OR x D The device may enter a Ventilator Inoperative state without a reboot preceding this condition x D | Philips Respironics is currently investigating this issue. Philips Respironics will be in contact as soon as additional appropriate actions have been determined. | 2024-03-21 15:56
Boston Scientific Corporation | - | 455 | - | complaints related to an inability to advance the guidewire through the lumen of the device Please find more information in the attached Field Safety Notice | Recall | 2024-03-22 13:56
Parker Laboratories Inc | - | 1 | - | Parker has received nine complaints related to Redux Electrolyte Cr me Ref Lot Numbers A and A x D These customers reported that the gel has very low viscosity too thin A review of retained samples for these two lots confirmed the gel is too thin | Immediately examine your Inventory for the affected product listed Quarantie the un-used units of the above mentioned Lot numbers Return back the un-used units | 2024-03-24 15:25
Siemens Healthcare GmbH | - | 8 | - | If the coolant level in the cooling circuit drops below a certain level this may result in a situation in which the X ray tube is no longer sufficiently cooled and the system will display the message TUBE HOT have a break Several minutes later the system will block X ray to prevent further damages and displays the message NO XRAY TUBE TOO HOT x D | In the interim, customers are to follow the instructions in the attached FSN. Issues will be corrected by equipping the tube cooling unit with a water level sensor and the software will be updated accordingly. | 2024-03-25 12:09
Philips Healthcare (Suzhou) Co., Ltd. | - | 2 | - | The patient support floor plates for MR systems may not have been anchored to the floor according to Philips requirements causing the patient support to become unstable with the potential for tipping see Figure This movement can inadvertently lead to harm to the patient and or operator under certain circumstances x D x D Philips has not received any harm or injury associated with this issue as of February x D | Philips is providing this Field Service Notice (FSN) Letter which contains recommendations for continued use of the systems referenced in Section 4. A Philips representative will contact you to schedule time for an FSE to visit your site to perform an inspection regarding the patient support stability. (Reference FCO 78100581). | 2024-03-25 14:50
Medtronic, Inc. | - | 2610 | - | Medtronic received a complaint that prior to use of a DLP Intra Mammary Artery I M A Cannula Customer Facing Number CFN Finished Good FG lot number the customer identified that the sterile packaging was not sealed Seven pouched devices complaint unit plus additional units were returned and it was confirmed there were several un sealed areas with no adhesive transfer from the Tyvek onto the formed film During the investigation of this issue Device History Record s DHR s were reviewed for all Medtronic FG lots that were manufactured with the supplier Tyvek lot A that was used to package the confirmed affected Medtronic FG lot During this review a Tyvek sample in the DHR did not have adhesive coating for Medtronic FG lot number CFN Therefore the scope of confirmed affected FG lot numbers are and x D Per supplier investigation one specific Tyvek master roll did not have adhesive coating this master roll was tracked to a specific pallet of rolls and Medtronic was able to locate all Medtronic FG lot numbers manufactured with all rolls included in that pallet All Medtronic FG numbers manufactured with Tyvek rolls of that specific pallet are within the scope of this issue CFNs as those could potentially be impacted by this sterility issue The root cause will be further investigated and documented through CAPA x D | • Review your inventory for listed product. • Immediately identify and quarantine all unused listed product in your inventory. • Return unused listed product in your inventory to Medtronic. Your local Medtronic representative can assist you in the return of affected product as necessary. • Please complete and return the enclosed Customer Acknowledgment Form even if you do not have unused inventory. • Please share this notification with others in your organization as appropriate. If product listed above has been forwarded to another facility, please notify the facility of this Urgent Field Safety Notice. • Please maintain a copy of this communication in your records. | 2024-03-25 15:23
Coloplast A/S | - | 343 | - | Ureteral dilator is a sterile single use device intended to be used for dilation of the ureter during ureteroscopic procedure x D The device is packaged in a double pouch packaging x D Possible sterility issue was detected in Coloplast s facility on some Ureteral dilator products This issue on the Ureteral dilator packaging has been identified during testing plan Defect is not easily visible by the users All lot were put in quarantine for testing x D potentially non compliant devices were shipped to customers | A CAPA has been opened to investigate further and to set up corrective actions. Correction done immediately after awareness of the issue: - All defective batches in stock are in quarantine (on hold) - Control of all inventory | 2024-03-25 16:04
ResMed PTY LTD | - | 273 | - | In the latest software release R SX Astral ventilators will now monitor the ongoing x D performance of the supercapacitor If the software detects that the supercapacitor is degraded or x D malfunctioning one of three new System Faults will activate on the ventilator with an audible alarm x D and message on the User Interface These System Faults do not stop ventilation x D Returning the ventilator to a Service Centre to replace the circuit board will correct the System Fault x D detected x D It is mandatory that all affected Astral ventilators are upgraded to this latest software version by the x D end of March x D | In September 2023, ResMed updated the Astral service process to improve detection of this issue during service. ResMed will continue to replace circuit boards for ventilators that are identified in service to be affected by this issue. In March 2024, ResMed released SX544-0603 Astral Software application Release 6.2 for Astral 100/150 ventilators to improve detection of this issue during ventilator use. All ResMed service centres will upgrade the software of any Astral 100/150 ventilator to SX544-0603 Astral Software application Release 6.2, when returned for any service activity. | 2024-03-26 12:00
Laboratoires VIVACY | - | - | - | A syringe of STYLAGE XL Lidocaine is inserted in the packaging of STYLAGE M lidocaine x D The labelling of the box labelling of the blister and packaging items including needles G are those of STYLAGE M lidocaine The syringe is labelled STYLAGE XL lidocaine and does contain STYLAGE XL lidocaine gel | Action To Be Taken by the User: 1. Immediately check your inventory and put in quarantine all devices from batch 232281101. 2. Disseminate this notice to all relevant individuals in your organization. 3. Return the completed form of Appendix 1 to the sender (Laboratoires Vivacy France or your Distributor) which confirm the acknowledgement receipt and the quantity put in quarantine, even if no product identified. 4. After reception of Appendix 1 and in case you have products to return, Laboratoires VIVACY France or your Distributor will contact you to organize the return of products and proceed with the replacement. Action To Be Taken by the Distributor /Pharmacy: 1. Immediately check your inventory and put in quarantine all devices from batch 232281101. 2. Disseminate this notice to all relevant individuals in your organization. 3. Identify the users who have received the affected batch and send them the Field Safety Notice (FSN). Ask them to put in Quarantine affected products, to acknowledge receipt and send back affected products to your organization using Appendix 1. Please update Appendix 1 (section 4) with your contact details prior sending. 4. When you receive all replies from all customers, fill in the form in Appendix 2 and send it back to Laboratoires VIVACY. 5. After reception of the form and in case you have products to return, Laboratoires VIVACY will contact you to organize the return of products and proceed with replacement. | 2024-03-26 14:28
Smiths Medical | - | - | - | to reduce false alarms | software update | 2024-03-26 15:14
Philips | - | 135 | - | The UPS devices are backup batteries utilized by PIC iX central monitoring stations and their associated hardware components These devices are intended to provide power to the PIC iX system or components monitors networking devices servers etc in the event of power outages ensuring that the PIC iX system remains functional during power fluctuations x D It has been discovered that an internal component of the UPS devices produces less output energy than specified resulting in the UPS inability to power up as required which affects the power supply to PIC iX system These failures can occur during a power outage or when the UPS is only acting as a conduit for the sites main line power to the PIC iX system x D | A Philips representative will contact you to schedule a visit from a Philips Field Service Engineer who will replace the malfunctioning UPS. | 2024-03-27 09:56
Randox Laboratories Ltd | - | - | - | Randox Laboratories have realigned C and Haptoglobin in Liquid Protein Calibrators IT to x D reference material ERM DA k IFCC x D C has been reassigned in calibrators lots IT IT and IT IT packed into batches x D and C results for Quality Control material and x D patient samples is expected to recover higher by approximately using the updated calibrator x D targets x D Haptoglobin has been reassigned in calibrator lots T IT packed into batches and x D Haptoglobin results for Quality Control material and patient samples is expected to recover x D higher by approximately using the updated calibrator targets x D Updated calibrator targets have been listed below The updated Instructions For Use IFU are x D available on www randox com please discard the previous versions of the IFUs and download the x D latest versions | - Review results generated with the affected batches in line with the clinical profile of the patient. - Discuss the contents of this notice with your Medical Director. - Complete and return the response form 12187-QA to technical.services@randox.com within five working days. Transmission of Field Safety Notice: Send a copy of the FSN to all affected customers and to those who need to be aware within your organisation. | 2024-03-27 13:05
Smiths Medical | - | - | - | software updates | locate all affected pumps and ensure all users or potential users of these devices are immediately made aware of this notification and the completed mitigations. | 2024-03-27 14:56
Maquet Cardiopulmonary GmbH | - | 1 | - | During service measurements in course of Field Action on the CARDIOHELP i devices in the field it was identified that the measured patient leakage current exceeded the limit values of A specified in standard EN A Patient leakage currents of up to A were measured x D Investigations have shown that only Cardiohelp i devices with installed sensor panels with an old sensor bridge hardware version are affected The new hardware version is not affected by this failure x D The scope of the Cardiohelp i devices with the old sensor bridge hardware version extends from the first device built with the serial number to plus serial number and | On-site device modification/ inspection • Inform all customers possessing the affected products promptly about this Field Action by sending the Field Safety Notice for Customers. • Contact the customer immediately to arrange the performance of the Electrical Safety Test according to IEC 62353 and subsequent repair or the return of the CARDIOHELP-i to a Getinge representative | 2024-03-27 15:54
Fresenius Medical Care AG & Co. KGaA | - | 1238 | - | It has been determined through testing that non dioxin like polychlorinated biphenyl acids PCBAs could leach from the Stay Safe Catheter Extension which are manufactured using peroxide cross linked silicone tubing PCBA levels decrease with the Shorter Stay Safe Catheter Extensions and decrease over the treatment time x D To date Fresenius Medical Care has not received any complaints related to PCBAs | Until an updated silicone catheter extension material is available, it is recommended by Fresenius Medical Care that healthcare providers do the following: - Use only the shortest Stay Safe Catheter Extension length (e.g. 25 cm) for patients weighing less than 40kg. - Place the letter in a prominent position until the unaffected stock is available. Fresenius Medical Care is in the process of updating the catheter extension silicone to PCBA-free material. | 2024-03-27 23:24
Acutronic Medical Systems AG | - | - | - | Acutronic Medical Systems AG is pleased to make available Addendum to the fabianTM Instructions x D for Use IFU for SW V This Addendum provides users with additional information that supplements and x D clarifies the existing IFUs for the following three topics to ensure safe and proper operation of the ventilators x D x D fabianTM Flow Sensor Calibration Requirements x D Supported Exhalation Membrane Holders x D Setting up PRICO Predictive Intelligent Control of Oxygenation x D x D A brief summary of the supplemental information included in the attached IFU Addendum is shown below x D x D fabianTM Flow Sensor Calibration Requirements A reminder for users to install and calibrate a flow x D sensor in accordance with the IFU before using the fabian TM ventilators When operating in HFOV mode x D failure to install and calibrate the flow sensor can result in incorrect volume monitoring and false alarms x D In this situation the system may temporarily increase or decrease the Pmean and the amplitude over a x D period of seconds as a consequence of the patient disconnect alarm Fluctuations of settings and delays x D in appropriate treatment may lead to inadequate ventilation of patients that may result in x D Hyperventilation Hypoxia Hyperoxia Lung Injury x D x D Supported Exhalation Membrane holders Provide guidance for users to use only current version of the x D exhalation membrane holder REF that is identified in the IFU Changing the exhalation x D membrane holder type requires a new calibration to be performed x D x D Setting up PRICO Add guidance to users to set the O starting value before initiating PRICO As x D designed PRICO will be disabled under certain alarm conditions When PRICO is in use the O setting x D should be checked routinely during care to ensure it is appropriate for the patient s current condition x D When PRICO is switched off or disabled due to an alarm condition PRICO will restart at the current O x D setting To ensure the PRICO starts from an appropriate starting value it is recommended this is set x D before initiating PRICO x D x D Note The information in the Addendum to the fabian TM Instructions for Use for SW V will be available in x D the next revision of the IFUs | Actions to be taken by distributors / authorized technical service partners Provide Addendum to the fabian™ Instructions for Use for Software Version 5.2.2. Actions to be taken by the end users Please confirm receipt and review the contents of the notification package (includes this notification, the Distributor/End-User Response Form, and the Addendum to the fabian™ Instructions for Use for SW 5.2.2). Please complete the attached Distributor/End-User Response Form and return it to GMB-AMS-FSCAresponsecentre@vyaire.com We respectfully request to receive the fully completed and signed Distributor/End-User Response Form no later than March 23, 2024, or 30 days from receipt. | 2024-03-31 13:21
Fisher & Paykel Healthcare Limited | - | - | - | The reason for the voluntary limited recall relates to a speaker configuration in Airvo and myAirvo devices manufactured before August x D x D The speaker in the Airvo and myAirvo devices is intended to provide the user with auditory alerts and alarms under certain conditions In devices manufactured before August the speaker configuration may result in distorted intermittent or inaudible alarm sound levels x D x D This does not affect the therapy delivered by the Airvo and myAirvo device The device will otherwise perform as intended The visual alarm on the display will continue to function and notify the user of the alarm state However in the absence of an audible alarm if there is an interruption to therapy a patient may experience oxygen desaturation x D x D Beginning August a new speaker configuration from a different supplier was implemented into the manufacturing of Airvo and myAirvo devices x D x D Airvo and myAirvo devices manufactured on or after August are not subject to this recall x D | Actions required by you- Step 1 a. Identify any Affected Product in your inventory by checking the Reference (REF) and Serial Number (SN) on the product label underneath the base of the unit or the label on the box. b. Place the Affected Product in quarantine. Step 2 Complete Section A – Inspection of Stock on the Medical Device Recall Response Form and return as specified on the form. Step 3 Contact your F&P representative OR F&P Regional Office to arrange the collection of the Affected Product and to obtain replacement product. Actions for affected products you may have distributed- Step 1 a. Review your sales records and identify if any Affected Product has been distributed to your customers and complete Section B – Notification to Customers in the Medical Device Recall Response Form. b. Sign the completed Medical Device Recall Response Form and return it to your F&P Representative or F&P Regional Office. Step 2 If you identify that any of the Affected Product has been distributed to your customers, create a list of affected customers for tracking purposes using the Second Consignee Spreadsheet provided in the email. Step 3 a. Create a Medical Device Recall Letter and Medical Device Recall Response Form for each customer. b. Use the templates and edit only the text in red: • Second Consignee_Medical Device Recall_Initial Letter • Second Consignee_Medical Device Recall_Response Form Step 4 Send the Medical Device Recall Letter and Medical Device Recall Response Form to all affected customers within 5 business days of receiving this letter, using a courier service (with track and trace). Should a patient need to continue to use the device, the patient is instructed to contact their physician. Step 5 Update the relevant information on The Second Consignee Spreadsheet. Note: The Second Consignee Spreadsheet and all Response Forms must be kept and sent to your F&P Representative. Step 6 a. Where a customer fails to respond to the Medical Device Recall Letter within 5 business days of initial contact via letter, please follow up a minimum of three times via courier with a Reminder Letter once every subsequent 5 business days. b. Create the Reminder Letter using the Second Consignee_Medical Device Recall_Reminder Letter template. Enter the type of reminder letter (First, Second or Final) and the date on which you will send the Reminder Letter. c. Enter the date and summary of letters sent and contact attempts in The Second Consignee Spreadsheet. | 2024-04-01 09:57
Zimmer Inc | - | - | - | possible thread form issue for the locking holes resulting that the locking screws may not properly mate with the plate The improperly mated screw may not be recognizable by the user since the screw may not correctly lock | No specific patient monitoring instructions related to this Field Safety Corrective Action that are recommended beyond your existing follow-up schedule. quarantine all affected product. Zimmer Biomet representative will be removing the affected product(s) from \affected facilities facility. | 2024-04-02 14:36
SAUDI MAIS CO FOR MEDICAL PRODUCTS | - | 4800 | - | Based on hospital visit and investigation of retained samples we identified that the cause of change color of solution is may be due to poor storage and transportation conditions | All Hemocure with reference HC56 having these Lots 52207001- 52207002- 52207003- 52207005 - 52207006- 52207007- 52207008- 52207009- 52207010- 52207114- 52207115- 52207123 Saudi Mais Co for Medical Products will recalling the product and will replacing the remaining effected quantities with a new Lot | 2024-04-03 14:12
Siemens Healthcare Diagnostics Inc. | - | 201600 | - | Issue Description x D Siemens Healthineers has confirmed the potential for falsely elevated Chol LDLC and Trig results on the Atellica CH and Atellica CI analyzers when the previous result in the cuvette was Iron This issue can impact calibrator quality control QC and patient results x D See Appendix A for Additional Data in Tables x D All future lots are impacted until further notice x D The Atellica CH HDL Cholesterol HDLC and Atellica CH Iron Iron assays are not impacted by this issue x D x D Kindly review the attached UFSN | Customer Actions • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Refer to short-term mitigation steps in “Appendix B” while Siemens works towards transitioning customers back to the Iron_2 assay. • Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. • Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Appendix B Customers with Multiple Atellica Analyzers Segregate Iron3 to a single analyzer and perform Chol_2, Trig_2, and LDLC testing on a different analyzer(s). For customers utilizing automation systems, please consult the appropriate instructions to perform the steps below. 1. On the system(s) that will run Iron3 assay: • Allow the system to complete testing and go into standby. • Once in standby, disable the Chol_2, LDLC, and Trig_2 assays by following instructions in the Atellica Online Help "About Enabling and Disabling Assays on the System." • Resume processing all assays except Chol_2, LDLC, and Trig_2. 2. On the system(s) that will run Chol_2, LDLC, and Trig_2 assay: • Allow the system to complete testing and go into standby. • Once in standby, disable the Iron3 assay by following instructions in the Atellica Online Help "About Enabling and Disabling Assays on the System." • For Atellica CH: o Perform 221 Chol_2 tests to eliminate the Iron3 reagent impact prior to testing with Chol_2, LDLC, or Trig_2. • For Atellica CI: o Perform 130 Chol_2 tests to eliminate the Iron3 reagent impact prior to testing with Chol_2, LDLC, or Trig_2. • Resume processing all assays except Iron3. 3. You may request 1 replacement kit of Chol_2 reagent to be used in the steps above, via the attached FIELD CORRECTION EFFECTIVENESS CHECK response form. Customers with a Single Atellica Analyzer Stop use of Iron3 and identify alternative methods for Iron testing. Prior to continuing routine processing on the Atellica system for assays besides Iron3, follow the steps below: 1. When the system is in standby, disable the Iron3 assay by following instructions in the Atellica Online Help "About Enabling and Disabling Assays on the System." 2. Remove Iron3 reagent prior to testing with Chol_2, LDLC, or Trig_2. • For Atellica CH: o Perform 221 Chol_2 tests to eliminate the Iron3 reagent impact prior to testing with Chol_2, LDLC, or Trig_2. • For Atellica CI: o Perform 130 Chol_2 tests to eliminate the Iron3 reagent impact prior to testing with Chol_2, LDLC, or Trig_2. 3. Resume all testing except for Iron3. 4. You may request 1 replacement kit of Chol_2 reagent to be used in the steps above, via the attached FIELD CORRECTION EFFECTIVENESS CHECK response form. Kindly review the attached UFSN | 2024-04-04 09:28
Abbott | - | 179 | - | Abbott is informing customers of the potential for electromagnetic interference EMI to cause an inadvertent mode change in x D a subset of Aveir VR LSP V devices manufactured with firmware version This issue is corrected through a x D firmware upgrade x D x D There have been zero reports of permanent harm to patients due to this issue with two devices replaced due to early x D detection of Recommended Replacement Time RRT see Risk to Health below If present the mode change is detected x D during a Merlin programmer interrogation session as the Aveir VR device may potentially present in Emergency VVI EVVI x D or MRI VOO mode As communication with the implanted Aveir VR pacemaker requires a Merlin programmer remote x D monitoring is not currently available the issue will be detected at a scheduled in clinic follow up unless patient symptoms x D prompt an earlier evaluation x D x D The issue may cause an Aveir VR device to enter either EVVI or MRI mode The parameters for EVVI mode are VVI pacing x D at V ms and bpm and MRI mode is VOO mode at V ms and bpm Compared to nominal settings x D the x D increased pacing output and rate of each mode may reduce longevity | Upgraded Merlin™ PCS 3650 programmer software facilitates the download of Aveir device firmware version 19.12.00 through an automatic prompt to the user during an in-clinic interrogation. All device settings and therapies remain active during the firmware download. Zero (0) devices with firmware 19.12.00 have experienced the reported mode change issue. This programmer software and upgraded device firmware are available to clinics starting April 2024 | 2024-04-07 15:01
SAUDI MAIS CO FOR MEDICAL PRODUCTS | - | 100612 | - | Leakage is due to some variation in product dimensions | all 3way stopcock with reference number 401-0010 having the lot: 220717 and 230815 shall be returned to the manufacturer and we will replacing the effected quantities with a good one | 2024-04-07 15:20
Hubei MEIBAO Biotechnology Co, Ltd. | - | 371520 | - | CT lines of the device are not displayed | recall the problem batch | 2024-04-14 09:18
GE Healthcare Finland Oy | - | 10 | - | Patient received too much medication due to a confusing dosing details in the Administration Recording window in Centricity High Acuity Critical Care CHA CC product The administration was based on a mixture order of two products with primary ingredient and the dosing by was defined by the main preparation of the order x D | Customer letter to be sent to affected customers. Product change to be deployed to correct impacted software versions. | 2024-04-15 10:24
Covidien llc | - | 402 | - | One lot of Tri Staple Black Reinforced Intelligent Reload has the potential for a damaged cartridge Potential damage to the cartridge was caused by an incorrect assembly station equipment adjustment If the damage is present poor staple formation and or incomplete staple line may result | • Immediately identify and quarantine all unused Tri-Staple™ 2.0 Black Reinforced Intelligent Reload, Model SIGTRSB60AXT, from the specified lot mentioned above. • Return all unused product in your inventory, corresponding to the specified lot mentioned above, to Medtronic as indicated in the enclosed Customer Acknowledgement Form. ▪ Please complete the enclosed Customer Acknowledgement Form, even if you have no product to return. ▪ Pass on this information to all those who need to be aware within your organization or to any organization where the potentially affected devices have been transferred. ▪ Please keep a copy of this letter in your file. | 2024-04-16 16:11
Philips Healthcare (Suzhou) Co., Ltd. | - | 15 | - | What the problem is and under what circumstances it can occur x D x D Philips has identified a hardware issue with a metal mounting box on the rotating scanner on rotor heat change box located within the Incisive CT CT CT system This component s structural integrity may become compromised resulting in potential contact with other gantry components located within the Incisive CT CT CT system during rotation Other components could be damaged due to the contact with the detached component x D x D If this issue occurs a loud noise will be emitted and the system will shut down x D The cover may become displaced if the following sequence occurs x D Damaged component s come loose x D The damaged component s contact other components in the system and cause them to break x D A broken component makes contact with the top right gantry cover x D x D The displaced cover could create a small gap allowing a fragment of a damaged component to be expelled at a low velocity The maximum weight of a fragment that could be expelled is g ounces x D x D This issue was discovered internally by Philips As of Mar Philips has not received any customer complaints related to this issue x D | Philips will contact you to schedule a time for a Field Service Engineer (FSE) to visit your site and install a solution to address the issue (reference FCO72800819 and FCO72800820). | 2024-04-17 15:10
Smith Medical | - | 230 | - | This failure mode can lead to inadequate ventilation for the patient and complete dislodgement of the x D tracheostomy tube Hypoxia underdose cardiopulmonary collapse bradycardia hypotension x D respiratory arrest or asphyxia can potentially result from the partial or complete detachment of the flange x D To date Smiths Medical has received five reports of serious injury and zero deaths potentially x D related to this issue | 1) Check all inventory locations within your institution for the affected catalog numbers and lot numbers listed in the notification and discontinue use. Discard all affected products following your institution’s process for discarding. If discarding is not immediately possible at your facility, then the product should be quarantined until disposal. 2) Share this notification with all potential users of the device to ensure they are aware of this notification and proposed mitigations. If the devices are used at another location, please ensure this communication is delivered there. 3) Complete and return the attached Customer Response Form to EMEA-FSN@icumed.com within 10 days of receipt to acknowledge your understanding of this notification, even if you do not have any affected product. | 2024-04-18 11:24
QUOTIENT DIAGNOSTICS | - | 79 | - | to instability of the kits lots | immediate discontinue the use of the remaining affected product beyond a 6 months shelf life. | 2024-04-18 13:51
Boston Scientific | - | - | - | some patients experience undesired sensation and or transient worsening of movement disorder symptoms secondary to a transient loss of stimulation during IPG charging for patients with rechargeable DBS IPGs | Boston Scientific field representatives can install firmware update on implanted IPGs for patients who have experienced undesirable sensations during IPG charging due to this device reset behavior | 2024-04-19 10:21
Philips Medical Systems (Cleveland), Inc. | - | 7 | - | What the problem is and under what circumstances it can occur x D x D Philips has identified multiple software issues affecting Philips IQon iCT Ingenuity and Brilliance CT systems Detailed descriptions and advice to customers pertaining to these issues are provided in Appendix A x D x D Philips has not received any reports of injury or serious harm associated with these issues | Philips will contact you to schedule a time for a Philips Field Service Engineer (FSE) to visit your site and install the software update to resolve the issues listed in Appendix A. | 2024-04-19 11:23
KARL STORZ ENDOSKOPE | - | 132 | - | It was found that there is insufficient evidence to show that the reprocessing method of the products was adequately validated This issue affects all lot numbers of the referenced KARL STORZ article numbers | 1. Immediately quarantine and discontinue use of associated part numbers listed. 2. Pass on this urgent field safety notice to all users of the products listed above and all other persons who need to be aware within your organization. 3. If you have or may have distributed the devices listed, please identify and promptly notify those recipients, or provide KARL STORZ a list of customers who received/may have received the products listed. 4. Return the filled feedback form by Fax or E-Mail to the indicated contact within 15 calendar days from the date of receipt. 5. Get in touch with your KARL STORZ representative to return affected products. 6. Please report any incidents related to this issue to the manufacturer, dealer or local representative and, if applicable, to the national competent authority, as this is important feedback. | 2024-04-21 10:54
Maquet GmbH | - | 1 | - | We have identified that an issue might prevent the devices listed above from performing as intended Therefore the device could break apart during surgery We are currently not aware of any incidents where this issue has occurred in the field | Getinge will initiate an immediate field action of all affected device units. You will be contacted by your Getinge sales or service representative to plan for the update of your device. | 2024-04-21 16:51
Olympus Medical Systems Corp. | - | 70 | - | The triangle tip of the KD L and KD L Triangle Tip Electrosurgical Knives breaking off during use The Triangle Tip Electrosurgical Knives are single use and are designed to be used with Olympus endoscopes and electrosurgical units The KD L knife is intended to cut is issuing high frequency current within the upper digestive tract The KD L knife is intended to cut and coagulate issue using high frequency current and flushing devices for submucosal injection within the digestive tract | Distribute the FSN | 2024-04-21 17:33
Medtronic Navigation, Inc. | - | 19 | - | An anomaly was identified related to the loss of text and numerical characters in the banner s of the software application This anomaly impacts the cranial procedures within StealthStation S Application versions and The software anomaly causes numbers or letters to be missing from the displayed text and to be replaced by a space instead All other navigation information on the screen is accurate including active navigation images The symptoms of the missing character anomaly can appear in two different large text fields in the software x D x D The Distance to Target banner x D The characters can be missing when the screen is not split into multiple views at once and only guidance view is visible i e up view In this instance the anomaly is expected to be transient as the information updates as the instrument moves during live navigation x D The Tip Stop Point banner x D This text is only visible in biopsy procedures after a surgical plan within the software is set and the trajectory is locked During the placement and navigation of the biopsy needle to the target location the navigation system accurately tracks the needle location with respect to the surgical plan x D x D Based on the missing character s that can impact the tip stop point or distance to target this anomaly could occur in the following procedures Cranial Resection Cranial Biopsy Catheter Placement and Electrode and Probe Placement x D The missing character s can vary in the displayed text any numerical value or letter in the Tip Stop Point banner or the Distance to Target text can be missing It is unlikely that both the Distance to Target and Tip Stop Point errors would occur at the same time as the Distance to Target text is only missing with the up guidance view x D | • Please review this information with all physician users. If you have any questions related to this issue, please contact your Medtronic Representative • Please complete and return the customer acknowledgement form enclosed with this letter acknowledging receipt of this information. • This notice needs to be passed on to those who need to be aware within your organization or to any organization where the potentially affected devices have been transferred. Please maintain a copy of this notice in your records. | 2024-04-21 18:28
Leica Biosystems Nussloch GmBH | - | 4 | - | As part of our post market surveillance we have identified an issue related to poorly processed or damaged x D biopsy tissue specimens on the HistoCore PEGASUS and HistoCore PEGASUS Plus systems The problem is x D associated with reagent levels exceeding the maximum fill level marks on reagent bottles or in the paraffin tanks x D The overfilling of reagents by users excess reagent carried from the basket during initial loading or excessive x D reagent in the biopsy pads wraps can lead to reagent levels surpassing the maximum fill level mark As a result x D excess reagent may flow into other bottles through the air manifold during a processing protocol causing crosscontamination | At the time you receive this notification, you should perform the following actions immediately: BIO SYSTEMS -Check reagent status in reagent bottles and paraffin status in paraffin bath. In case of any signs of contamination (such as muddiness, turbidity, or liquid separation) are observed, you should stop using the instrument for tissue processing and call for service. -When no cross contamination was observed, continue with the following actions: -In order to ensure defined condition, please check reagent/paraffin level. If they are over the maximum level mark, you should pour out the excess reagent/paraffin, and empty the condensation bottle. -To secure further processing, reagent and wax bath levels shall be always checked in terms of not exceeding maximum fill level. As long as maximum fill level is between the minimum and maximum level, your instrument is fine, and you can continue processing. -Please always kindly follow the recommendations below. 1) Before running the protocol, inspect the reagent level in reagent bottles and paraffin level in paraffin baths. Make sure that all processing reagents and paraffin level is between the Min and Max. and regularly empty the condensation bottle (weekly). (Refer to IFU 2.2.3 Operating the instrument; 9.3 Cleaning and maintenance schedule) 2) The reagent bottle and paraffin bath should be filled between the Min and Max line as indicated in the Instructions for Use. (Reagent level: Refer to IFU 4.4 Basic instrument/hardware 4.4.1 Retorts) 3) Formalin should be dripped out from the basket before adding it into the retort to avoid large carryover volume of formalin into the reagent bottle. 4) Please check if the current carrier material has large volume of carryover. If yes, please change to an appropriate carrier material. 5) Please do not clean the Mold or metal lid in the retort during the cleaning protocol. As a further countermeasure, a Leica Service Engineer will arrange a visit at your facility to replace the current air manifold on your HistoCore Pegasus and HistoCore Pegasus Plus. | 2024-04-22 08:54
Ortho Clinical Diagnostics | - | 596 | - | confirmed an issue involving VITROS Immunodiagnostic Products Free T Reagent Pack where some lots x D listed below may produce positively biased results | Discontinue distributing and discard all remaining sales units of the above listed lots of VITROS Free T3 Reagent Pack (and associated calibrators) from your inventory. • Send the enclosed customer letter and customer Confirmation of Receipt form to all customers who were shipped the above listed lots of VITROS Free T3 Reagent Pack from your facility. • Complete and return the enclosed distributor Confirmation of Receipt form no later than April 26, 2024. Upon receipt of your completed form, QuidelOrtho will provide credit for, or replacement of, discarded product. • Please forward this notification if the affected product was distributed outside of your facility. | 2024-04-22 09:05
Randox Laboratories | - | - | - | Randox Laboratories can confirm the analyte NEFA in the above listed lots and packed into the above listed batches has increased over the shelf life of the product and requires the targets and ranges to be reassigned Please see below table for updated targets and ranges x D x D The value sheets have been updated and are available on www randox com please discard the previous version and download the updated version | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2024-04-22 09:30
Randox Laboratories Ltd | - | - | - | Randox Laboratories have identified reagent cartridge packs with a compromised vacuum x D seal in the above listed batches This may impact the stability and performance of the x D reagents Root cause under investigation | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Review stock of affected batches. Send an image of any affected cartridges to technical.services@randox.com Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2024-04-22 10:09
Applied medical | - | - | - | Applied Medical is conducting a voluntary recall of Voyant EA Electrosurgical Generator from specified lot numbers due to the identification of an electrical board component from a non approved manufacturer thereby making it non conforming to our UL approved Safety Certification Report list of critical components This non conformance presents no risk to the patient or user safety and does not impact the use of compatible Voyant devices | Due to our commitment to providing only the highest quality products, Applied Medical has decided to recall the impacted EA020 units and resolve the non-conformance by replacing each generator with one that contains ‎عطapproved UL critical component. Please note that Applied Medical is taking all ‎needed field action to determine the affected quantities in the market and proceed in retrieving them. For product return and regulatory questions, please contact Mariane ‎Kfoury, Senior Operations Manager on CustomerSatisfaction-MENAT@a¦ liedmedical.com | 2024-04-22 11:09
Olympus Medical Systems Corp. | - | 70 | - | Olympus has become aware of an increased trend from both repairs and customer complaints of the UHI stopping CO gas supply with the front panel LED turning off Based on the analysis of customer complaints this issue was determined to be associated with a Control board or CR board pressure sensor circuit failure x D In an effort to maximize patient safety and mitigate any potential risk to patient health Olympus will replace the CR board of UHI devices which were manufactured more than years ago Olympus will be contacting customers to schedule this replacement requirement x D | Action 1. You should continue to quarantine the product until: For units 5 years and older: Your unit receives both the software update and a CR Board replacement. For units under 5 years of age: Your unit receives the software update. 2. Indicate in the reply form if your facility continues using any of the UHI-4 devices. 3. Olympus will contact you based on device age and parts availability to schedule a repair of the CR Board. Olympus will be prioritizing customers who have continued to use the UHI-4. 4. Olympus will contact you in late summer of 2024 regarding the software update to address the over insufflation. | 2024-04-22 11:19
Technopath Manufacturing Ltd | - | 691 | - | Please find attached FSN letter | Please find attached FSN letter | 2024-04-23 13:28
FUJIFILM Healthcare Americas Corporation | - | 1 | - | FUJIFILM has become aware that measurements on a Secondary Capture D image that does x D not have pixel spacing in the DICOM header when combined with a Breast Tomo Series results x D in incorrect measurements | FUJIFILM is providing Synapse PACS Version 7.3.1 to resolve this issue. | 2024-04-24 14:04
Spineart SA | - | 152 | - | The entry point on the thread inside the screw head was not machined correctly | Please refer to the attachment for actions that should be taken. (Reference (FSN) FSCA-01-2024). | 2024-04-25 13:24
Philips Medical Systems Nederland B.V. | - | 68 | - | What the problem is and under what circumstances it can occur x D x D Philips has identified an issue where the g MDU global Mains Distribution Unit L terminal connection may become loose creating a hotspot see figure that may cause smoke fire to alarm in the hospital s technical room The g MDU located in the technical room is the single entry point for the hospital electricity supply and distributes the electricity toward the various cabinets and components of the MR Scanner x D x D If the connection failure occurs the user may observe the following x D Smoke and or fire alarm in the examination room x D Smoke and or fire in the hallway or technical room x D Power being cut from the system x D Philips has received five complaints of burnt g MDU terminal connections and smoke burning smell in the technical room associated with the issue There was no report of injury or serious harm | Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to inspect the g-MDU connections in the technical room and apply the proper torque to the connection if necessary. (reference FCO78100582). Philips plans to start implementing corrections in July 2024. | 2024-04-25 14:18
Randox Laboratories Ltd | - | - | During routine quality checks, Randox has identified that the foam gasket on the Evidence MultiSTAT chip heater assembly may not perform as expected. Foam that ‎كا امهfunctioning as expected may cause light to leak into the Charged Coupled Device (CCD) Camera during imaging. You may continue to use the analyser. As a precautionary measure, Randox will review images from your Evidence MultiSTAT to assess analyser performance, and a site visit by a Service Engineer will be arranged to replace the affected chip heater assembly. | During routine quality checks Randox has identified that the foam gasket on the Evidence x D MultiSTAT chip heater assembly may not perform as expected Foam that functioning x D as expected may cause light to leak into the Charged Coupled Device CCD Camera during x D imaging You may continue to use the analyser As a precautionary measure Randox will x D review images from your Evidence MultiSTAT to assess analyser performance and a site visit x D by a Service Engineer will be arranged to replace the affected chip heater assembly x D | - Export support files for the last 10 samples run and share with Randox using the link to be provided - Discuss the contents of this notice with your Medical Director. - Complete the return the response form 12187-QA to technical.services@randox.com within five working days. e Affected chip heater assembly replacement to be arranged by Randox Service Engineer | 2024-04-28 09:48
Medartis AG | - | 15 | - | One lot of mm screws was placed in packaging for mm screws x D x D The outer packaging of the medical device indicates a different size to the one supplied in the box The error in the packaging used does not affect the quality of the product within the packaging | Inspect stock immediately and quarantine the incorrectly packaged stock on hand. Any sets with products from incorrect packaging are to be checked to ensure the correct length of the screws in the 22 mm slots; quarantine any 16 mm screws from the 22 mm screw slots. Replacement of the products incorrectly packaged will be arranged as soon as possible after the products have been returned. Recommendations for patients who have already received an incorrectly packaged implant: Post-operative routine examination is sufficient. No additional follow-up is needed. | 2024-04-28 15:47
Getinge | - | 35 | - | There is a discoloration corrosion inside the vaporizer have been observed under normal operating conditions Analysis has revealed chemical degradation of Sevoflurane resulting in formation of hydrogen fluoride within the vaporizer This issue has only been observed when using low water content Sevoflurane Piramal Sevoflurane and Sevoflurane Baxter This issue has only been observed on vaporizers of specific batches please refer the Appendix table in the FSN Kindly refer to the FSN for more detailed description on the issue and potential hazards | Do not use Piramal Sevoflurane or Sevoflurane Baxter or other low water content Sevoflurane with: 6682282, Vaporizer Sevoflurane Maquet filling (SN 17003 - 23454) Do not use AbbVie Sevorane®/Ultane® or Maruishi Sevofrane with: 6682282, Vaporizer Sevoflurane Maquet filling (SN 17003 - 23454) if it has ever been used with Piramal Sevoflurane and/or Sevoflurane Baxter or other low water content Sevoflurane, even if currently used with AbbVie Sevorane®/Ultane® or Maruishi Sevofrane. Do not use 6682285, Vaporizer Sevoflurane Quik-Fil (SN 17111 - 23424): • If it has ever been used with Piramal Sevoflurane and/or Sevoflurane Baxter or other low water content Sevoflurane, even if currently used with AbbVie Sevorane®/Ultane® or Maruishi Sevofrane. | 2024-04-28 18:54
Stryker | - | 299 | - | The issue could prevent the device from delivering instructional voice prompts x D to the user during use of the device The device has visual instructional icons still x D present and is functional but if the issue is not identified by the customer prior x D to use it could potentially lead to no therapy or a delay in therapy In addition x D there may be risk of shock to the user due to the absence of the stand clear x D voice prompt There has been one reported serious incident to date in x D which the device failed to deliver audio prompts Serious incidents or quality x D problems experienced with the use of this product may be reported to the x D Stryker Representative | The company is notifying all customers that have HeartSine Samaritan PAD & Omron HDF-3500 devices within the identified range of potentially affected devices | 2024-04-29 10:30
GE Healthcare Finland Oy | - | 3079 | - | GE HealthCare has become aware of the potential for loss of monitoring on CARESCAPE Canvas CARESCAPE Canvas Smart Display and certain versions of CARESCAPE B and CARESCAPE B monitors if following the replacement of the CPU timekeeper battery a mains power loss occurs during active monitoring In the unlikely event that this occurs it can result in the displays or monitors not turning on after the mains power is restored This can result in a delay in recognizing changes in patient condition until a replacement monitor is connected x D There have been no injuries reported as a result of this issue | Customer letter (UMDC/FSN) and updated Service Manuals & Supplemental Information Manuals (SIM) are available at Customer Documentation Portal. Discontinue usage and destroy and/or delete all affected versions of the Service Manuals and Supplemental Information Manual. FMI kit containing updated manuals on USB device to be delivered as requested. | 2024-04-29 11:02
Welch Allyn, Inc. | - | - | - | Baxter is issuing a Correction for the power cords used with the Welch Allyn Connex ProBP Digital Blood x D Pressure Device and Welch Allyn Spot Vision Screener Baxter received reports of an issue related to the x D construction of the power cord not meeting the insulation rating per country specific requirements and x D international electrical standards | Advice on actions to be taken by the distributor and the user 1. Inspect the condition of the power cords. If fraying or other damage is observed, users should discard the power cord immediately 2. Healthcare providers may continue to use the affected power cords after they are inspected for damage. 3. Healthcare providers should regularly inspect the power cords for fraying or other damage. 4. Once Baxter has replacement power cords, a follow-up notification will be sent with additional instructions on how to request replacement power cords. 5. Acknowledge receipt of the Field Safety Notice (Using the Customer Reply Form). 6. Forward a copy of this letter to other facilities or departments within their institutions to ensure that those locations are aware of this information. 7. Distributors are asked to notify their customers of this communication in accordance with their procedures. | 2024-04-29 12:23
Draeger Medical Systems, Inc. | - | 293 | - | M does not have type CF floating isolation causing higher patient leakage current at certain ports on Infinity M | Phase 1 The attached Safety Notice is being issued to all affected customers until 2nd May 2024 and track responses. Phase 2 Start installing new induction plate of M500 using TSB rework kit after availability of the kit and rework should be finished within 12 months. | 2024-04-29 13:30
Draeger Medical, Inc. | - | - | تم فتح حالة برقم SA-12-03-24-319 | x D SA | A case was opened under reference number: SA-12-03-24-319 | 2024-04-29 13:32
Synovis Surgical Innovations | - | 1246 | - | On March Baxter implemented labeling changes on CE marked Peri Guard Repair Patch and Supple Peri Guard Repair Patch products including removal of the indications for abdominal wall defect and hernia diaphragmatic femoral incisional inguinal lumbar paracolostomy scrotal and umbilical repair from the Instruction for Use IFU Baxter has since received seven complaints reporting infection abscess from the same clinic in Italy Six of them were determined to be correlated to off label use of Peri Guard Repair Patch multiple product codes in abdominal surgery occurring at one customer in Italy x D The indications on the labeling changed from x D For use as a prosthesis for pericardial closure and soft tissue deficiencies which include defects of the abdominal and thoracic wall hernias diaphragmatic femoral incisional inguinal lumbar paracolostomy scrotal and umbilical and intracardiac and great vessel repair x D To x D For use as a prosthesis for pericardial closure and soft tissue deficiencies which include defects of the thoracic wall and intracardiac and great vessel repair x D x D As these changes in indications may not have been taken into account by all product user on February Italian Ministry of Health Directorate General of Medical Devices and Pharmaceutical Services asked Baxter to inform customers of the changed intended use limitation of the use to ensure the correct use of the devices in the market | Description and justification of the action (corrective/preventive) The change was not driven by any known safety concerns or to reduce a potential for patient harm. However, there is a lack of clinical data supporting the safety and effectiveness of the Peri-Guard Repair Patch and Supple Peri-Guard Repair Patch for abdominal wall and hernia defect repair. As such, these indications are no longer approved for CE-Marked Peri-Guard Repair Patch and Supple Peri-Guard Repair Patch and should be considered “off-label” in the European Union. Advice on actions to be taken by the distributor and the user • Clinicians may continue to use the Peri-Guard Repair Patch and Supple Peri-Guard Repair Patch products listed above however, clinicians should be aware of the recent removal of the product abdominal wall defect and hernia repair indications from the IFUs. • Acknowledge receipt of the Field Safety Notice | 2024-05-01 15:58
Smith Medical | - | 25 | - | The dimensional changes made to the Cardinal Health Monoject syringes | Customer Acknowledgement | 2024-05-02 09:51
Smith Medical | - | 25 | - | The dimensional changes made to the Cardinal Health Monoject syringes | Customer Acknowledgement | 2024-05-02 10:01
Philips | - | 1 | - | Philips Hemodynamic Application provides the functionality of hemodynamic calculations including physiological measurements iFR and FFR For an iFR FFR measurement the Philips Hemodynamic Application has to synchronize pressure wave data received from two devices the IntelliVue X and the FM PIM in order to compare those data x D x D The aorta blood pressure Pa or AO is measured via a catheter connected to the IntelliVue X patient monitor while the internal blood pressure Pd is measured with a pressure wire connected to a separate transducer FM PIM Both the IntelliVue X and FM PIM send the pressure measurements to the Philips Hemodynamic Application x D x D Before an iFR FFR measurement is performed the waveforms of the Pa and Pd pressures need to be normalized see Section of the Instruction for Use After normalization the waveforms of the Pa and Pd pressures as acquired by the IntelliVue X and FM PIM should overlap meaning that the diastolic and systolic points waveform peaks and valleys should be at the same time position and the average pressures of both measurements should be identical x D x D Philips has become aware that this synchronization may not be correct which causes the wave data from one device to be delayed with respect to the wave data from the other device This time shift can occur in two different forms being x D A delay of the Pd wave with respect to the Pa wave that is increasing over time drift The drift decreases in time towards zero eventually leading to a constant shift between the Pa wave and Pd wave see Figure below x D A constant delay of the Pd wave with respect to the Pa wave of exactly one or more cardiac cycles without any drift x D x D Note that the issue is in certain cases corrected by the system If the cumulative delay of the Pd wave with respect to the Pa wave becomes larger than two seconds the Philips Hemodynamic Application will reset all calculations From that moment on the initial error is not present anymore and the drift will not occur anymore x D | Philips will install a software update to correct this issue. A Philips representative will contact you to schedule an onsite visit to perform the upgrade (reference FCO72200542). | 2024-05-02 15:27
Sleepnet Co | - | 522 | - | Do not use this mask if you or anyone in close physical contact with your mask has an active medical implant or metallic implant that will interact with magnets x D | Notify hospitals and clients of the safety alert. | 2024-05-06 10:16
Verathon Medical (canada) ULC | - | 33 | - | Verathon has received complaints from customers about a loss of image when using a Core x D or Core FHD monitor in conjunction with unique combinations of the GlideScope Core m x D QuickConnect cable and Spectrum QC Single use Video Laryngoscopes No injuries x D have been reported Verathon has implemented a correction in the enclosed software update Core x D v and Core FHD v to prevent a loss of image | Please take the following actions to install the software correction via the USB. 1. Plug your GlideScope Core 15 or Core 15 FHD into the power supply and connect to a hospital-grade power outlet. 2. Use the enclosed USB drive with the corrective software upgrade and related instructions to upgrade your GlideScope Core 15 or Core 15 FHD monitor(s) (approximately 5-7 minutes for the upgrade). 3. Verify the software version on the Administrative Settings tab. 4. Contact Verathon Customer Care to contact your Territory Manager to perform the upgrade at your facility. 5. Complete the Correction Response Form in this packet and return it to Verathon via email to CSNotifications@verathon.com | 2024-05-06 14:54
I-CODI CO. LTD | - | 205000 | - | The label is missing information such as the period of use and recommended temperature of storage | ensure the missing information such as ( the period of use and recommended temperature of storage) is available and clearly mentioned on the outer package label of all affected product | 2024-05-06 19:06
Maquet GmbH & Co KG | - | 3 | - | Under certain conditions we have identified that an issue might prevent the device from performing as intended Batteries in above mentioned devices may not be loaded completely by the charger unit The required charging voltage is not provided during the charging period | Replacement of all potentially affected Charging Units by a Getinge service technician. | 2024-05-07 14:54
Becton, Dickinson and Company BD Biosciences | - | 172 | - | BD observed an increasing trend of complaints for unusual staining pattern nonspecific aggregates high background double positive signal and or antibodies contaminated with substance for the affected product codes in Table The investigation detected a damaged component in affected product codes that may affect integrity of stained cells causing increase in cell debris i e background as shown on SSC vs CD positive leukocyte flow cytometry plots resulting in atypical increase of double positive events on CD CD vs CD CD and or CD CD CD vs CD CD plots as shown in Figure | Clinical User Actions 1. It is recommended that the clinical laboratory staff do not use BD Multitest™ 6-Color TBNK affected lots listed above and/ or replace with the product lots that are manufactured using different non-compromised raw materials. 2. BD recommends discussing the content of this letter with your laboratory management regarding the need to review previous test results associated with these affected lots. BD Actions: BD is currently investigating the root cause and will implement appropriate corrective actions to prevent recurrence of this issue. Customer Actions: • Cease use of any unused affected lots of BD Multitest™ 6-color TBNK. • Identify and quarantine all unused affected lots of BD Multitest™ 6-color TBNK. • Make a note of the lot numbers and destroy all unused affected units. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 31st May 2024. • Circulate this notice to all those who need to be aware within your organization or to any organization where the potentially affected products have been transferred. • If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: • Cease distribution. • Identify, quarantine, making a note of the lot numbers then destroy all unused affected lots of BD Multitest™ 6-color TBNK. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 31st May 2024. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. | 2024-05-07 15:04
Sleepnet Co. | - | - | - | Do not use this mask if you or anyone in close physical contact with your mask has an active medical implant or metallic implant that will interact with magnets x D | Notify hospitals and clients of the safety alert. | 2024-05-08 09:28
Illumina, Inc. | - | - | - | While conducting routine cybersecurity analysis Illumina s Product Security Team identified an uncontrolled product security risk associated with the communication protocol between Illumina s products and the customer provided network storage described above servers and or computers Illumina determined that if secure protocols are not used an unauthorized actor who has already gained privileged access to the customer network could intercept and exploit these communications | A Field Safety Notice (FSN) communication is planned to be distributed to all impacted customers in all the affected countries on May 8, 2024. This FSN informs the users of the availability of a Cybersecurity Guidance to be downloaded, whichincludes recommendations on how customers can secure their internal network’s communication with Illumina’s products. It also informs them that failure to follow these instructions or implement network security best practices to protect their systems could leave their organization exposed to cybersecurity risks. This FSN also requests users to complete the Verification Form after carrying out all the steps in the instructions provided on their specific product(s) identified as affected and to return it to Illumina Technical Support. Alternatively, the users can also provide the requested confirmation through email. For new devices, Illumina will ensure this Cybersecurity Guidance is provided to all customers as part of the labelling materials. Illumina is providing Cybersecurity Guidance that includes recommendations on how customers can secure their internal network’s communication with Illumina’s product. Failure to follow these instructions or implement network security best practices to protect your systems could leave your organization exposed to the risks described above. The pertinent local and international regulatory bodies, including the Competent Authorities, are being notified of this issue. | 2024-05-08 10:59
Philips | - | 20 | - | Philips has identified that oil may leak from the HA FlexTrack Trolley when pressing the hydraulic pedal resulting in a potential slipping falling hazard x D x D Philips has received no reports of adverse events associated with this issue as of October x D | Based on Philips investigation, it has been identified that the oil leakage has occurred after servicing. As the leakage may not be immediately visible during servicing activities, a Philips representative will contact you to schedule a time for a Field Service Engineer (FSE) to visit your site to perform the final correction using updated instructions (reference FCO78100595). Additionally, Philips is in process of updating the service manual instructions to prevent reoccurrence of the issue. | 2024-05-08 14:15
Q-med AB | - | 3 | - | Not provided | Not provided | 2024-05-08 15:46
Q-Med AB | - | 1 | - | Not provided | Not provided | 2024-05-08 15:53
Abbott | - | 17 | - | Abbott has received customer complaints for the System Monitor due to atypical screen display behaviors All x D System Monitors can potentially show atypical screen display behaviors Restarting the System Monitor should x D resolve these issues in most cases There have not been any serious adverse health consequences associated x D with these complaints | Please distribute this notice to those who need to be aware within your institution and forward to any institution where potentially affected devices have been transferred. Abbott is in the process of notifying the applicable regulatory agencies about this issue. Please report any adverse reactions or quality problems experienced with the use of these products to your local Abbott representative. | 2024-05-08 17:25
Megadyne Medical Product, Inc. | - | 879 | - | Megadyne has received reports of patient burn injuries in procedures where the Mega Soft Patient Return Electrodes were used A root cause investigation on the reports included testing which showed a combination of factors when present together may result in potential for thermal injuries x D The combination of these conditions may be more likely when the pad is used with infants and small children Because the pediatric pad is designed for patients between to pounds which would be predominantly patients under the age of the decision was made to discontinue and recall the v pediatric pad product | Customers are required to immediately return all MEGADYNE™ MEGA SOFT™ Pediatric Patient Return Electrodes subject to this recall that are in inventory. FA: 2364349 | 2024-05-08 22:41
Buttterfly Network, Inc. | - | - | - | The problem is likely due to insufficient mobile device memory space on certain iOS models | Product Defect Correction, Customers are to upgrade their Butterfly iQ app through the Apply App Store to software version 2.33.1 and re-execute the diagnostic health check. | 2024-05-09 09:42
Dongyang Tianhong Daily Necessities Co; LTD | - | 285 | بناءا على طلب هيئة الغذاء والدواء ان المنتج لا يحمل رخصة تسويق طبية , وانما يحمل رخصة تسويق كمنتج تجميلي | - | The sold quantity has been fully recalled, and an attached image confirms that the customer’s inventory balance for the item is zero. | 2024-05-09 13:33
Blue Belt Technologies, Inc. | REAL INTELLIGENCE CORI Starter Kit | 6 | - | Blue Belt Technologies Inc a wholly owned subsidiary of Smith Nephew Inc collectively referred to herein as Smith Nephew has initiated a field action to voluntarily correct the CORI Surgical System with Real Intelligence Software due to increased Marker Registration Error MRE which may cause tracker arrays to flicker on screen of the CORI Surgical System during use | 1. Ensure that the contents of this Field Safety Notice are read and understood by those within your organisation who may use Real Intelligence CORI. 2. Please see attached IFU Addendum for troubleshooting tips and notes to enhance the consistency of visibility of the markers and tracking arrays. This IFU Addendum can be found at https://ifu.smith-nephew.com/ by searching IFU number 1000264245. 3. Please complete the Customer Response form and email or fax it to your national Smith+Nephew agency/distributor. 4. Please maintain awareness of this notice and resulting action for an appropriate period to ensure the effectiveness of the corrective action. 5. A Smith & Nephew Robotics Representative will be deployed to your site to implement a new software version once it is available. | 2024-05-09 16:53
Arrow International LLC | Intra-Aortic Balloon Catheter Kit | 89 | - | Teleflex is initiating this voluntary FSCA for the above mentioned products due to reports indicating an infrequent condition that when not identified and corrected promptly could result in serious health consequences x D | "1-Read the FSN carefully. 2-Place a copy of this notification with all affected product and ensure that the notification is included with future distributions of affected product. 3-Product is not being removed; you may continue to use the products in scope of this advisory notice per the additional instructions, warnings, and cautions contained herein." | 2024-05-11 21:44
Intersurgical Limited | ANAESTHETIC breathing system | 2 | - | We received a report where circuits were supplied without the hole in the end of the reservoir bag tail which means the circuit is unable to be used MHRA Ref It appears the reservoir bags with closed tail had been accidentally mixed with the open tail bags and subsequently assembled in error All systems were found before use We have inspected further products still in the UK stock and have identified further small quantities in other similar products Lots with the same error | Intersurgical has received a report from a customer where 3 circuits were supplied without the hole in the end of the reservoir bag tail which means the circuit is unable to be used (MHRA Ref: 2024/003/027/601/031). It appears the reservoir bags with 'closed tail' had been accidentally mixed with the 'open tail' bags and subsequently assembled in error. All 3 systems where found before use. We have inspected further products still in the UK stock and have identified further small quantities in other similar products/Lots with the same error. We would expect the problem to be identified during set-up and pre-use tests in accordance with the IFU, as was the case in the report we received. These are fast moving products, where we have not received any other similar reports for more than 5 years and so we do not expect this to be a widespread problem. However, to minimize any further risk we intend to distribute the attached FSN and Customer Reply Form to all countries that have received the potentially affected products in the last 5 years (as per product expiry period), to warn them of the possible error and to check the reservoir bag tail is open before use. | 2024-05-12 12:19
Intersurgical Limited | PAEDIATRIC MAPLESON F INFANT T-PIECE WITH 0.5L OPEN TAIL BAG AND ELBOW | 2 | - | We received a report where circuits were supplied without the hole in the end of the reservoir bag tail which means the circuit is unable to be used MHRA Ref It appears the reservoir bags with closed tail had been accidentally mixed with the open tail bags and subsequently assembled in error All systems were found before use We have inspected further products still in the UK stock and have identified further small quantities in other similar products Lots with the same error | Intersurgical has received a report from a customer where 3 circuits were supplied without the hole in the end of the reservoir bag tail which means the circuit is unable to be used (MHRA Ref: 2024/003/027/601/031). It appears the reservoir bags with 'closed tail' had been accidentally mixed with the 'open tail' bags and subsequently assembled in error. All 3 systems where found before use. We have inspected further products still in the UK stock and have identified further small quantities in other similar products/Lots with the same error. We would expect the problem to be identified during set-up and pre-use tests in accordance with the IFU, as was the case in the report we received. These are fast moving products, where we have not received any other similar reports for more than 5 years and so we do not expect this to be a widespread problem. However, to minimize any further risk we intend to distribute the attached FSN and Customer Reply Form to all countries that have received the potentially affected products in the last 5 years (as per product expiry period), to warn them of the possible error and to check the reservoir bag tail is open before use. | 2024-05-12 12:19
Ethicon LLC | PDS PLUS | 313 | - | Ethicon identified a manufacturing issue on a specific packaging machine that resulted in a hole in the primary packaging of a small percentage of VICRYL VICRYL Plus PDS PDS Plus MONOCRYL and MONOCRYL Plus sutures manufactured between January and March The occurrence of this defect is rare with an estimated rate of of product presenting the condition of product is not impacted by this defect When present the hole is always and only on the first package in the horizontal box of quantity and it occurs in the same location on the bottom side foil cavity of the first package towards the peelable flaps | Medical device recall (removal) FA: 2364579 | 2024-05-12 20:36
Ethicon LLC | PDS II | 313 | - | Ethicon identified a manufacturing issue on a specific packaging machine that resulted in a hole in the primary packaging of a small percentage of VICRYL VICRYL Plus PDS PDS Plus MONOCRYL and MONOCRYL Plus sutures manufactured between January and March The occurrence of this defect is rare with an estimated rate of of product presenting the condition of product is not impacted by this defect When present the hole is always and only on the first package in the horizontal box of quantity and it occurs in the same location on the bottom side foil cavity of the first package towards the peelable flaps | Medical device recall (removal) FA: 2364579 | 2024-05-12 20:36
Philips Medical Systems DMC GmbH | MobileDiagnost wDR,MobileDiagnost wDR R2,MobileDiagnost wDR Release 2 | 6 | - | Philips has been made aware by supplier Sedecal of a potential problem related to interference between the steel cable and metal plate in the carriage of MobileDiagnost wDR systems telescopic arm The potential interference is caused by the shape of the cable s crimped terminals see Figure This issue may result in failure of the anti fall system which would prevent the telescopic arm from falling if the steel cable breaks If the steel cable breaks and the anti fall system does not activate the telescopic arm may fall on operators patients or service personnel x D x D Philips has not received any reports of this issue or harm as of April x D | A Philips representative will contact you to schedule a time for a Field Service Engineer (FSE) to visit your site to perform an inspection and if needed, install a new cable. (Reference FCO71200237). | 2024-05-14 14:10
Shifaa United for Medical Device | DMD SODIUM HYPOCHLORITE | 804 | - | Sell Without get MDMA | Contact with all customer to recall products, stop production till getting MDMA for product | 2024-05-14 15:01
Medtronic, Inc. | 1x4 Pocket Adaptor for Deep Brian Stimulation | 25 | - | Description of Issue x D In specific scenarios that include a replacement INS with a Pocket Adaptor and the use of a REPLACEMENT workflow also known internally as a cloning feature available in the Clinician Programmer Application A CP the Deep Brain Stimulation DBS CP and the Patient Programmer Application A PP incorrectly show Full Body Magnetic Resonance Imaging MRI eligibility instead of the correct Head Only MRI eligibility x D The issue with the CP and PP showing the incorrect MRI eligibility arises when an Implantable Neurostimulator INS Activa SC Model is replaced with an INS device with a different header design Model B or B With that replacement a Pocket Adaptor is used to connect the existing implanted extensions Model A or to the replacement INS During that replacement the HCP may use the cloning feature available in the CP to transfer data fields previously set for the existing INS and replicate them in the replacement INS One of those data fields is the portion of INS memory that is used by the CP and PP applications to determine MRI eligibility It is during the use of that cloning feature that this portion of the INS instrument memory portion that is used by the CP and PP applications to determine MRI eligibility does not get updated to account for the presence of a Pocket Adaptor and continues to indicate that there is no Pocket Adaptor implanted If the portion of INS instrument memory that is read by the CP and PP applications to determine MRI eligibility incorrectly indicates that the patient has no Pocket Adaptor then the displayed eligibility in the CP and PP applications incorrectly shows Full Body MRI eligibility rather than the correct Head Only MRI eligibility x D | • Complete and return the Customer Acknowledgement Form enclosed with this letter acknowledging receipt of this information. • Pass on this notice to all those who need to be aware within your organization and to other organizations on which this action has an impact. • Please keep a copy of this letter in your file. • Medtronic has provided an Optional Patient Letter template to facilitate your discussions with patients (attached). | 2024-05-15 16:59
Medtronic, Inc. | 2x4 Pocket Adaptor for Deep Brian Stimulation | 25 | - | Description of Issue x D In specific scenarios that include a replacement INS with a Pocket Adaptor and the use of a REPLACEMENT workflow also known internally as a cloning feature available in the Clinician Programmer Application A CP the Deep Brain Stimulation DBS CP and the Patient Programmer Application A PP incorrectly show Full Body Magnetic Resonance Imaging MRI eligibility instead of the correct Head Only MRI eligibility x D The issue with the CP and PP showing the incorrect MRI eligibility arises when an Implantable Neurostimulator INS Activa SC Model is replaced with an INS device with a different header design Model B or B With that replacement a Pocket Adaptor is used to connect the existing implanted extensions Model A or to the replacement INS During that replacement the HCP may use the cloning feature available in the CP to transfer data fields previously set for the existing INS and replicate them in the replacement INS One of those data fields is the portion of INS memory that is used by the CP and PP applications to determine MRI eligibility It is during the use of that cloning feature that this portion of the INS instrument memory portion that is used by the CP and PP applications to determine MRI eligibility does not get updated to account for the presence of a Pocket Adaptor and continues to indicate that there is no Pocket Adaptor implanted If the portion of INS instrument memory that is read by the CP and PP applications to determine MRI eligibility incorrectly indicates that the patient has no Pocket Adaptor then the displayed eligibility in the CP and PP applications incorrectly shows Full Body MRI eligibility rather than the correct Head Only MRI eligibility x D | • Complete and return the Customer Acknowledgement Form enclosed with this letter acknowledging receipt of this information. • Pass on this notice to all those who need to be aware within your organization and to other organizations on which this action has an impact. • Please keep a copy of this letter in your file. • Medtronic has provided an Optional Patient Letter template to facilitate your discussions with patients (attached). | 2024-05-15 16:59
Medtronic, Inc. | Vanta™ AdaptiveStim™ | 2 | - | Description of the Issue x D Medtronic has determined that Vanta may be susceptible to not communicating with the clinician programmer and or the patient programmer system HH Handset and TM Communicator Although the INS can continue to deliver therapy to the patient within programmed parameters the inability to communicate can put the patient at risk In some instances if there is no communication with the INS during the implant procedure troubleshooting may delay the surgical procedure and may require a replacement INS In other instances if there is no communication with the INS post procedure it might not be possible to adjust therapy and may cause the patient to experience inadequate therapy i e return of underlying disease symptoms In rare instances if communication cannot be restored to the INS via a recommended mitigation an unanticipated surgical intervention may be needed to explant and replace the INS x D | • At the patient’s initial programming session, ensure the patient programmer system is paired with the patient’s INS. • If a patient’s INS cannot communicate with the clinician programmer and/or patient programmer system, and it is not possible to make therapy adjustments, contact your Medtronic Representative. | 2024-05-15 17:24
Atrium Medical Corporation | Advanta VXT Vascular Graft | 1056 | - | The root cause of the issue was determined to be a piece of manufacturing equipment that was not operating as intended resulting in insufficient crimping of the Slider GDS components | - If you have any affected grafts of the Product REF numbers and Expirations Dates 27Feb2027 on Pages 1 and 2 of the FSN, please return the affected product. - If there is a medical need to use the device and there is no alternative device available, then the graft may be used by removing the Slider GDS component as described in the notice. | 2024-05-16 11:22
Atrium Medical Corporation | Flixene Vascular Graft | 1056 | - | The root cause of the issue was determined to be a piece of manufacturing equipment that was not operating as intended resulting in insufficient crimping of the Slider GDS components | - If you have any affected grafts of the Product REF numbers and Expirations Dates 27Feb2027 on Pages 1 and 2 of the FSN, please return the affected product. - If there is a medical need to use the device and there is no alternative device available, then the graft may be used by removing the Slider GDS component as described in the notice. | 2024-05-16 11:22
Cepheid | Xpert Carba R | 200 | - | se positive patient test results can result in unnecessary contact precautions leading to non x infectious adverse events and side effects from inappropriate antibiotic treatment | Please dispose of any GXCARBARP-CE-10 kits from batch 1001208183 lot 10808 in your inventory and Cepheid will provide replacement kits with the quantity requested. The replacement should ship within 5 business days following receipt of the completed response form. | 2024-05-16 12:28
Medtronic, Inc. | Ascenda™ Intrathecal Catheter | 440 | - | Medtronic has developed a revised design and manufacturing process for the Ascenda catheter The intent of the design changes is to reduce the risk of a specific possible pathway for tissue in growth inside the catheter interface with the catheter connector x D x D Design Changes x D The Ascenda design change focuses on improvement of the catheter connection seal quality functionality to reduce the probability of tissue growth into the Leak Path that could potentially cause an occlusion of the drug path within the connector The design updates to the Ascenda implement the same design of tube inner diameter ID and cone angle as that of the SC catheter to improve seal quality x D | • Complete the enclosed Customer Acknowledgement Form acknowledging that you have received this information. • Share this notice with all those who need to be aware of this design update within or outside your organization or to any organization where the potentially affected product has been transferred or distributed and maintain a copy of this notice in your records. | 2024-05-19 17:36
EIT Emerging Implant Technologies GmbH | EIT PLIF | 6 | - | The subject product is being recalled due to improper labeling on some parts in the lot the label on the outer carton of some of the subject products does not match the actual device in the package x D x D NOTE The implants are laser etched with the correct part number and correct lot number | Please Take the Following Steps: 1. Examine your inventory immediately to determine if you have the subject products and quarantine them immediately. DO NOT USE THE SUBJECT PRODUCTS. See attachment 1 to aid in identifying the subject products. 2. Contact your DePuy Synthes Sales Consultant or contact the customer support services at (Hanan Alzamel, HAlzamel@its.jnj.com) to coordinate the return/credits of the subject products. 3. Review, complete, sign, and return the attached Business Response Form (page 3 of this letter) to (Hanan Alzamel, HAlzamel@its.jnj.com) within three (3) business days of receipt of this notification. 4. Please complete the attached Business Response Form even if you do not have the subject products on hand. 5. Forward this notice to anyone in your facility that needs to be informed (e.g., those who manage, transport, store, stock, or use the subject products). 6. If any of the subject products have been forwarded to another facility, contact that facility and provide them with this notice. 7. Post a copy of this notice in a visible area for awareness and keep a copy for your records. FA: 2364175 | 2024-05-20 15:05
Acutronic Medical Systems AG | Fabian HFO Classic Ventilator | 28 | - | As part of FSCA SW V changed the criteria for detecting a patient circuit disconnect during HFO ventilation with the intent to reduce the potential for patient safety related risks x D While ventilating a patient in HFO mode with pressure amplitude Pamp x mean pressure Pmean and with the breathing circuit properly connected if the breath by breath leak monitor exceeds the ventilator triggers the disconnect alarm When this condition occurs the displayed volume and leak monitors reset to zero and the ventilator reduces the HFOV amplitude This is a protection mechanism to prevent lung injury upon reconnection following a true disconnection During this time if the patient is not disconnected this results in a small increase in Pmean and potentially large reduction in amplitude until the alarm is cleared x D | Acutronic has determined the root cause of the failure and will be providing a SW update. The update will occur in two stages. The first update to SW V5.2.3 will address the issues that were resolved with SW V5.2.2 except the ETT disconnect detection issue (issue 2). The subsequent update will address the ETT disconnect detection issue and associated alarm. Acutronic expects SW V5.2.3 to be available by end of Q3 calendar year 2024. Additional information regarding the availability of the subsequent revision to address the ETT disconnect detection issue will be communicated at release of SW V5.2.3. Acutronic will send the FSCA package, containing this FSN in English and in national languages, the SW V5.2.1 IFU with Addendum, the Distributor Response Form, and the End-User Response Form. Acutronic will make available SW V5.2.1 for fabian FTP. Acutronic will collect and follow-up on all response forms and the execution and completion of this FSCA. | 2024-05-20 17:17
Acutronic Medical Systems AG | fabian HFOi Ventilator | 28 | - | As part of FSCA SW V changed the criteria for detecting a patient circuit disconnect during HFO ventilation with the intent to reduce the potential for patient safety related risks x D While ventilating a patient in HFO mode with pressure amplitude Pamp x mean pressure Pmean and with the breathing circuit properly connected if the breath by breath leak monitor exceeds the ventilator triggers the disconnect alarm When this condition occurs the displayed volume and leak monitors reset to zero and the ventilator reduces the HFOV amplitude This is a protection mechanism to prevent lung injury upon reconnection following a true disconnection During this time if the patient is not disconnected this results in a small increase in Pmean and potentially large reduction in amplitude until the alarm is cleared x D | Acutronic has determined the root cause of the failure and will be providing a SW update. The update will occur in two stages. The first update to SW V5.2.3 will address the issues that were resolved with SW V5.2.2 except the ETT disconnect detection issue (issue 2). The subsequent update will address the ETT disconnect detection issue and associated alarm. Acutronic expects SW V5.2.3 to be available by end of Q3 calendar year 2024. Additional information regarding the availability of the subsequent revision to address the ETT disconnect detection issue will be communicated at release of SW V5.2.3. Acutronic will send the FSCA package, containing this FSN in English and in national languages, the SW V5.2.1 IFU with Addendum, the Distributor Response Form, and the End-User Response Form. Acutronic will make available SW V5.2.1 for fabian FTP. Acutronic will collect and follow-up on all response forms and the execution and completion of this FSCA. | 2024-05-20 17:17
Vyaire Medical GmbH | Twin Tube (V-707327) | 22 | - | The purpose of this communication is to inform you that Vyaire Medical GmbH is conducting a voluntary field action for the Twin Tube V and Twin Tube Probenschlauch due to the potential of the nozzle separating during patient use | Coordinate with Distributors/End-Users upon receipt of the fully completed and signed Distributor/End-User Response Form to return confirmed affected devices and exchange as required. ACTION TO BE TAKEN BY THE DISTRIBUTOR/END-USER: • Confirm receipt and thoroughly review the contents of the Customer Notification Package (includes this notification and the Distributor/End-User Response Form). • If potentially affected devices have been transferred to another location or organization, please forward the complete Customer Notification Package to the respective parties. • Inspect current inventory on-hand as described above under “How to Identify Affected Devices”. A 100% physical inventory inspection should immediately be performed to identify and isolate potentially affected devices. • Perform the Pull-out Test on all current inventory on-hand as described in Appendix A, “Pull-out Test”, to verify the adhesive is adequate prior to use. • Fully complete the attached Distributor/End-User Response Form and return it to GMB-EMEA-FSCA-RDX-INTL@Vyaire.com. The email subject line should be labeled “Response Form: FSCA-24-002-FSN-1". • We respectfully request the completed and signed Distributor/End-User Response Form to be returned within 30 days of receipt. | 2024-05-21 08:36
Vyaire Medical GmbH | Twin Tube Probenschlauch (707004) | 22 | - | The purpose of this communication is to inform you that Vyaire Medical GmbH is conducting a voluntary field action for the Twin Tube V and Twin Tube Probenschlauch due to the potential of the nozzle separating during patient use | Coordinate with Distributors/End-Users upon receipt of the fully completed and signed Distributor/End-User Response Form to return confirmed affected devices and exchange as required. ACTION TO BE TAKEN BY THE DISTRIBUTOR/END-USER: • Confirm receipt and thoroughly review the contents of the Customer Notification Package (includes this notification and the Distributor/End-User Response Form). • If potentially affected devices have been transferred to another location or organization, please forward the complete Customer Notification Package to the respective parties. • Inspect current inventory on-hand as described above under “How to Identify Affected Devices”. A 100% physical inventory inspection should immediately be performed to identify and isolate potentially affected devices. • Perform the Pull-out Test on all current inventory on-hand as described in Appendix A, “Pull-out Test”, to verify the adhesive is adequate prior to use. • Fully complete the attached Distributor/End-User Response Form and return it to GMB-EMEA-FSCA-RDX-INTL@Vyaire.com. The email subject line should be labeled “Response Form: FSCA-24-002-FSN-1". • We respectfully request the completed and signed Distributor/End-User Response Form to be returned within 30 days of receipt. | 2024-05-21 08:36
ICU Medical | Infusion Systems | 2860 | - | runtime may decrease earlier in the battery lifecycle than expected | Replace batteries | 2024-05-21 15:12
Agilent Technologies Singapore (International) Pte Ltd. | Dako CoverStainer | 0 | - | It has come to our attention that the assembly supporting the weight of the lower front door in some cases may x D cause the plexiglass to break at the upper left or upper right corners | Actions to be taken by the customer 1. Confirm that you have received this information by signing and returning enclosed acknowledgement form 2. Monitor your device and contact the service organization in case you discover a problem with the plexiglass on the lower front door. • In case the plexiglass of the lower front door is broken, immediately cease operation of the instrument • Do not remove the lower front door assembly, and do not run the instrument without the cover Actions to be taken by Agilent Agilent is currently improving the design of the assembly to correct the issue. When the new design becomes available, the Agilent Service Organization will update your device at the next service visit. | 2024-05-22 14:03
Maquet Critical Care AB | VITROS Immunodiagnostic Products Anti-HBs Reagent Pack | 13 | - | This issue may lead to positively biased patient and QC results due to the reduced calibrator signals in Calibrator Levels and x D x D The table below demonstrates the bias observed for results generated during our investigation of the affected lots Please note that as this issue is related to calibrator signal reduction over the shelf life of the product QuidelOrtho estimates an increase in the observed bias over time x D Lot Number Age From Manufacture Date x D In Weeks Average Bias Observed x D x D x D x D Positively biased QC results may be outside the expected QC limits leading to QC failure and potentially causing a delay in patient testing However this delay is unlikely to impact patient management due to the clinical utility of Anti HBs as described in the Intended Use in the Instructions For Use IFU of this product x D x D Falsely elevated patient results may breach the clinical thresholds causing Negative results to report as Borderline and Borderline results to report as Positive Refer to the VITROS Anti HBs Reagent Pack and Calibrators IFU for guidance on the interpretation of Borderline results x D x D A false positive result may lead to an erroneous assessment of the patient s immunity to Hepatitis B potentially resulting in the patient not receiving the Hepatitis B vaccine Hepatitis B Immune Globulin or appropriate counseling which may increase the patient s risk of contracting the Hepatitis B virus x D x D QuidelOrtho recommends a review of previous results between mIU mL generated using the associated lot s of VITROS Anti HBs Reagent Pack where only Anti HBs testing was performed and was positive for example post vaccination testing or where only x D Anti HBs was positive in the triple panel test Anti HBs HBsAg anti HBc Please review the Questions and Answers section at the end of this notification for more information x D x D As of May QuidelOrtho has received complaints relating to this issue with no reports of adverse events | • Discontinue using, render unusable, and discard the affected lots of VITROS Anti-HBs Calibrators and associated Reagent Pack remaining in your inventory. • Complete and return the enclosed Confirmation of Receipt form no later than Month DD, 2024. Upon receipt of your completed Confirmation of Receipt form, QuidelOrtho will provide credit or replacement for your discarded inventory. • Save this notification with your User Documentation or post this notification by each VITROS ECi/ECiQ/3600/5600/XT 7600 System until the issue has been resolved. • Please forward this notification if the affected product was distributed outside of your facility. • If your laboratory has experienced the issue with this product and you have not already done so, please report the occurrence to your local Global Services Organization. | 2024-05-23 13:52
Maquet Critical Care AB | VITROS Immunodiagnostic Products Anti-HBs Calibrators | 13 | - | This issue may lead to positively biased patient and QC results due to the reduced calibrator signals in Calibrator Levels and x D x D The table below demonstrates the bias observed for results generated during our investigation of the affected lots Please note that as this issue is related to calibrator signal reduction over the shelf life of the product QuidelOrtho estimates an increase in the observed bias over time x D Lot Number Age From Manufacture Date x D In Weeks Average Bias Observed x D x D x D x D Positively biased QC results may be outside the expected QC limits leading to QC failure and potentially causing a delay in patient testing However this delay is unlikely to impact patient management due to the clinical utility of Anti HBs as described in the Intended Use in the Instructions For Use IFU of this product x D x D Falsely elevated patient results may breach the clinical thresholds causing Negative results to report as Borderline and Borderline results to report as Positive Refer to the VITROS Anti HBs Reagent Pack and Calibrators IFU for guidance on the interpretation of Borderline results x D x D A false positive result may lead to an erroneous assessment of the patient s immunity to Hepatitis B potentially resulting in the patient not receiving the Hepatitis B vaccine Hepatitis B Immune Globulin or appropriate counseling which may increase the patient s risk of contracting the Hepatitis B virus x D x D QuidelOrtho recommends a review of previous results between mIU mL generated using the associated lot s of VITROS Anti HBs Reagent Pack where only Anti HBs testing was performed and was positive for example post vaccination testing or where only x D Anti HBs was positive in the triple panel test Anti HBs HBsAg anti HBc Please review the Questions and Answers section at the end of this notification for more information x D x D As of May QuidelOrtho has received complaints relating to this issue with no reports of adverse events | • Discontinue using, render unusable, and discard the affected lots of VITROS Anti-HBs Calibrators and associated Reagent Pack remaining in your inventory. • Complete and return the enclosed Confirmation of Receipt form no later than Month DD, 2024. Upon receipt of your completed Confirmation of Receipt form, QuidelOrtho will provide credit or replacement for your discarded inventory. • Save this notification with your User Documentation or post this notification by each VITROS ECi/ECiQ/3600/5600/XT 7600 System until the issue has been resolved. • Please forward this notification if the affected product was distributed outside of your facility. • If your laboratory has experienced the issue with this product and you have not already done so, please report the occurrence to your local Global Services Organization. | 2024-05-23 13:52
Philips | Fetal Monitor Transducers and Accessories /Avalon Ultrasound Transducer | 249 | - | It was found that the latest version of the wired Avalon Ultrasound Transducers may report inaccurate Fetal Heart Rate FHR measurements when monitoring multiples twins or triplets In situations where the physiological signal echo from the fetal heart is absent or very weak e g in early weeks of pregnancy there is a tendency of the wired Ultrasound transducers to interfere with each other and subsequently to produce an artificial FHR mostly at approximately bpm The use cases affected by this issue are twin or triplet monitoring with wired Ultrasound Transducers including at least one transducer with software version L The population at risk is women with multiple gestation twins triplets x D x D Note Wireless transducers and wired transducers with a different software revision than L are not affected by this issue x D | Philips is currently working on a solution. A Philips representative will contact you once the solution is available. | 2024-05-23 16:29
Beckman Coulter Inc. | IgM | 58 | - | Beckman Coulter has determined that the LIH Influence Check settings for IgM Lipemic Interference on the AU AU DxC AU and DxC AU analyzers incorrectly states which is equivalent to mg dL of Intralipid x D The correct LIH Influence Check setting for IgM Lipemic Interference for the AU AU DxC AU and DxC AU analyzers is which is equivalent to mg dL of Intralipid x D The LIH Influence Check setting for IgM Lipemic Interference on the AU is not affected x D The Lipemia interference specification in the IgM Instructions For Use IFU IFU Code BAOSR X BLOSR X and BLOSR X EU is correctly stated | Update the LIH Influence Check settings for Lipemia on AU480, AU680, DxC 500 AU and DxC 700 AU analyzers if they are enabled for IgM. To determine if the LIH Influence Check is enabled for IgM and to update your Lipemic setting, perform the following action: AU480 / AU680:  Select Menu > Parameters > Specific test parameters > General  Select the IgM test name and select ‘Edit’  Update the LIH Influence Check for IgM Lipemic interference to +++ DxC 700 AU:  Select Menu List > Configuration Parameters > Specific Test Parameters > Test Volume and Methods > General.  Select the IgM test name and select ‘Edit’  Update the LIH Influence Check for IgM Lipemic interference to +++ DxC 500 AU:  Select Menu > System Configuration > Test Menu  Select the IgM test name and select ‘Chemistry Details’  Select ‘Customer Configurable Options’  Update the LIH Influence Check for IgM Lipemic interference to +++  Select System Configuration > Activate Draft Beckman Coulter recommends sharing the content of this letter with your laboratory and/or Medical Director to determine if a review of previous patient test results should be conducted.  Discontinuance or disposal of this product is not necessary.  Retain a copy of this letter as it serves as current labeling.  Per the IFU, avoid highly lipemic samples when using the IgM assay | 2024-05-26 12:52
Abbott | 2nd Generation CentriMag Primary Consol | 27 | - | Support System Internal testing showed that due to a component change the CentriMag Console may x D not conform to the IEC standard If subjected to a power surge above kV and up to kV x D the CentriMag Console may shut down completely without alarm A power surge may be caused by x D including but not limited to the following lightning strike unregulated power switching or faulty x D equipment Pumping would cease and the patient would have to be switched to a backup | f the CentriMag Console shuts down without alarms or alerts present, follow Abbott Instruction for Use (IFU) for Console and Motor malfunction as described in the Section “Emergency / Troubleshooting”, and switch to a backup CentriMag Console which must be prepared as required backup components in the immediate vicinity. Note: When the backup CentriMag console is put into use, the previously used console can be restarted and set up as a backup console. Per the IFU and current practice, a full CentriMag Console backup system is always required to be available in the immediate vicinity of each patient whenever the CentriMag system is used Abbott performed additional testing and confirmed that the ability to pump and alarm is not affected in backup systems. As stated in the IFU these backup systems should have a battery charge sufficient for at least one hour of operation, be connected to wall power, and be powered off. Abbott is in the process of implementing an additional in-process test by Q3 of 2024, which will ensure | 2024-05-26 17:07
RPC Diagnostic Systems | DS-EIA-ANTI-TOXO-M (capture)” | 503 | - | error was found in the design of the experiment | inform the users that the correct value is shuold be not less than 1.0 instead than 1.8 so the user can update the assay SOP and accepted QC to the new value. either they are using the a manual or Automated "Third party" analyzer | 2024-05-26 22:33
MAQUET Cardiovascular, LLC | VASOVIEW HEMOPRO Endoscopic Vessel Harvesting System | 1826 | - | The VasoView HemoPro HemoPro Endoscopic Vessel Harvesting EVH Systems VH and VH W are indicated for use in minimally invasive surgery allowing access for vessel harvesting and is primarily indicated for patients undergoing endoscopic surgery for arterial bypass The harvesting tool of the VH VH and VH devices was not able to provide cautery or continued to provide cautery when not intended during use as per the complaints received No adverse events have been reported to date in association with this issue | Getinge will notify each facility using the VasoView HemoPro VH-4000 and VH-3000-W Endoscopic Harvesting System regarding these fluid ingress failures and potential risks to users and patients if this situation were to occur. Getinge is currently investigating this issue further to determine root cause and will notify customers in the event additional action needs to be taken to correct this issue. Additionally, the current VasoView HemoPro VH-4000 and VH-3000-W Instructions for Use (IFU) will be updated to include the additional safety information related to fluid ingress. | 2024-05-27 18:25
cadwell industries inc | Cascade IOMAX | 3 | - | An assembly error was discovered on the Cortical Module PN A Cortical Module x D insulator that is used to provide an electric barrier between circuit boards for TCS LCS and other x D internal parts was found to be installed incorrectly The insulator provides clearance between safety x D critical components and incorrect installation does not meet the minimum creepage air gap required x D to meet IEC x D IOMAX Cortical Modules manufactured between and may have been shipped x D with the insulator installed incorrectly | Cease use of the device immediately and return all identified devices to Cadwell Industries, and replace it Inc. | 2024-05-28 12:19
cadwell industries inc | Cascade IOMAX | 3 | - | An assembly error was discovered on the Cortical Module PN A Cortical Module x D insulator that is used to provide an electric barrier between circuit boards for TCS LCS and other x D internal parts was found to be installed incorrectly The insulator provides clearance between safety x D critical components and incorrect installation does not meet the minimum creepage air gap required x D to meet IEC x D IOMAX Cortical Modules manufactured between and may have been shipped x D with the insulator installed incorrectly | Cease use of the device immediately and return all identified devices to Cadwell Industries, Inc. | 2024-05-28 12:22
Fresenius Medical Care AG & Co. KGaA | multi (Version MultiFiltratePRO) | 28 | - | A software issue was identified in multiFiltratePRO software versions and This issue affects the initial functional test of the machine known as the T Test which occurs prior to setting up the device for treatment x D | Customers are advised to continue operating the above-affected devices according to the guidelines in the attachment. Fresenius Medical Care will provide stickers to be placed on the machine monitor. These stickers will serve as a reminder for displaying important message in the attachment, please contact your Fresenius Medical Care representative to obtain these stickers and ensure they are visibly attached to each affected device. The manufacturer is in the process of developing a software update to rectify this error. However, until the update is available and installed, please adhere carefully to the instructions provided in the attachment. | 2024-05-29 12:17
SACO | VAGINAL EXAMINATION GEL | 159400 | - | The medical device does not have Medical Device Marketing Authorization Therefore it can t ensure its safety efficacy and performance x D | Stop using and return all affected products to the manufacturer. | 2024-05-29 13:08
Implantcast GmbH | MUTARS® RS extension piece | 42 | - | Safety related changes to the instruction for use and surgical technique | Corrective action done by informing the user | 2024-05-29 19:00
Implantcast GmbH | MUTARS® RS stem | 42 | - | Safety related changes to the instruction for use and surgical technique | Corrective action done by informing the user | 2024-05-29 19:00
Medtronic MiniMed | Ambulatory Insulin Infusion pumps | 2645 | - | Kindly refer to the letter | Replace the battery cap on your previous case pump by following the steps below. 1. Gather your previous case insulin pump and all battery caps in your possession. a. Only replace the battery cap on your previous case pump. If you have a new case pump, do not replace the cap on it. 2. Before you begin, make sure you have a new battery available; do not remove the battery cap unless you have a new battery available. 3. Look at the bottom and top of each battery cap and identify which battery cap(s) is the new cap (model number ACC-1529) for your previous case pump and which is the previous battery cap(s) (model number ACC-1527). 4. Discard all previous battery caps, model number ACC-1527 for the previous case pump. There is no need to return anything to Medtronic. 5. Install the new battery cap (model number ACC-1529) onto your previous case pump. 6. The battery cap provided is for your primary previous case insulin pump. If you are using more than one previous case pump impacted by this issue, please contact Helpline | 2024-05-30 02:04
ZOLL Medical Corporation | 731 Series Ventilator | 21 | - | The documented Patient Safety Ventilator MRI information was inadvertently omitted from the manuals x D | - Locate all old versions of affected Operator’s Guides and/or Quick Reference Guides. - Request for replacement manuals. - Alert all users of the ZOLL Ventilator - with MRI Compatibility, to ensure they are aware of the safe distance requirements and are trained to these manual updates, refer to the attachment for more information. - Dispose of obsolete guides. | 2024-05-30 07:48
GE Healthcare Finland Oy | Centricity High Acuity Anesthesia 5.X | 5 | - | GE HealthCare has become aware of an issue where certain medication and fluid administration details are not shown in the Drugs and Fluids frame within the CHA A Trend View when using TCI Target Controlled Infusion pumps The issue can occur when all the following conditions are present x D A medication infusion initially being administered in a non TCI mode is restarted in TCI mode via the same pump x D The Infusion Mode for the medication is not adjusted to TCI within CHA A x D AND x D The utilized pump does not communicate TCI units of measure x D x D If all of the above three conditions occur details of the TCI dosing as well as administrations of other drugs and fluids following restarting the infusion in TCI mode will not appear in the Trend View This can result in potential incorrect medication administration x D x D x D | The CHA system will be upgraded with the corresponding patch level. | 2024-05-30 16:52
Beckman Coulter Inc. | Access HBc Ab | 109 | - | Beckman Coulter has determined that the Access HBc Ab reagent kit lot number listed in the table above may contain an unusual amount of paramagnetic particle PMP clumping at the top of Well | Discontinue using the Access HBc Ab reagent kit lot listed above. Discard all remaining reagent packs from this lot, following your local laboratory protocols.  Beckman Coulter recommends sharing the content of this letter with your laboratory and/or Medical Director regarding the need to review previous patient test results.  Contact your local Beckman Coulter representative to order replacement product for this reagent kit lot. | 2024-06-02 14:20
One Lambda, Inc. | FlowPRA® HLA Class I & II Screening Tests | 3 | - | The FlowPRA HLA Class I x D Screening Test Catalog ID FL Lot and FlowPRA HLA Class I and II Screening Test Catalog ID x D FL Lot may demonstrate a lower percentage of positive reactions with low levels of A or A x D antibody specificities A lower percentage of positive reactions can lead to potential false negatives for x D HLA Class I antibodies in samples with low panel reactive antibody levels In addition lot specific user x D document Document ID TDX OLI DMR LOT revision has incorrect reference values for the x D distribution of certain HLA Class I antigen specificities | Refer to updated lot specific user documents (Document ID TDX-OLI-DMR-LOT-8536, revision 02). - Review test results – review samples that were assigned negative for HLA Class I Antibody using the impacted product. - Alert other organizations to which incorrect test results from the use of the devices have been sent. - Complete the attached Acknowledgement Form and return to One Lambda, Inc. - Complete all action items within one month upon receipt. | 2024-06-02 15:19
Philips Healthcare (Suzhou) Co., Ltd. | Multiva 1.5T | 2 | - | Philips has identified an issue where the g MDU global Mains Distribution Unit L terminal connection may become loose creating a hotspot see figure that may cause smoke fire to alarm in the hospital s technical room The g MDU located in the technical room is the single entry point for the hospital electricity supply and distributes the electricity toward the various cabinets and components of the MR Scanner x D x D If the connection failure occurs the user may observe the following x D Smoke and or fire alarm in the examination room x D Smoke and or fire in the hallway or technical room x D Power being cut from the system x D Philips has not received any complaints of burnt g MDU terminal connections and smoke burning smell in the technical room associated with the issue from Multiva T systems There was no report of injury or serious harm x D | Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to inspect the g-MDU connections in the technical room and apply the proper torque to the connection if necessary. (reference FCO78100583). Philips plans to start implementing corrections in July 2024. Please be assured that maintaining a high level of safety and quality is our highest priority. If you need any further information or support concerning this issue, please contact your local Philips representative. | 2024-06-03 12:55
Ortho Clinical Diagnostics Inc. | Vitros Chemistry Products Calibrator Kit 29 | 8 | - | maybe Impact to Results | • Discontinue using, render unusable, and discard any remaining inventory of VITROS Chemistry Products Calibrator Kit 29, Lot 2953 by the correct expiration date, 24-Jul-2024. • If you have previously obtained a Certificate of Analysis, then download a new one from ORTHO PLUS on quidelortho.com. • To help assure the product is not supported beyond its expiry date, upload ADD DRV 6286 or higher before 24-Jul-2024. Once DRV 6286 or higher has been uploaded the product will no longer be supported and you must use an alternate lot. • Complete the enclosed Confirmation of Receipt form no later than Month XX, 2024. Upon receipt of your completed Confirmation of Receipt form, QuidelOrtho will provide credit or replacement of your existing inventory of VITROS Chemistry Products Calibrator Kit 29, Lot 2953. • Save this notification with your User Documentation or post this notification by each VITROS 5,1 FS/4600/5600/XT 7600 System until the issue has been resolved. • Please forward this notification if the product was distributed outside of your facility. | 2024-06-04 09:24
DiaSorin Inc. | LIAISON Testosterone Kit | 16 | - | DiaSorin has become aware that five kits from the affected lot contain an empty conjugate vial There is no risk to patient health The analyzer would alert the user that no liquid was detected in the conjugate vial and stop the assay routine No patient results would be reported If the conjugate vial is filled and all assay validation criteria pass you may confidently report patient results | Please check your inventory of the affected lot of LIAISON® Testosterone kits and dispose of any with an empty conjugate vial (part number 25493, lot number 206443). The conjugate is a large vial attached to the integral with a white seal. | 2024-06-04 09:42
Becton Dickinson and Company | BD BBL™ Staphyloslide™ Latex Test | 200 | - | Description of the problem x D x D BD previously provided a field action notification regarding BD BBL Sensi Disc to inform users to cease use for Haemophilus testing Since then BD has conducted a root cause investigation and has determined that AST testing can be reinstated with additional instructions x D x D Laboratories should perform and QC results should be acceptable for each day that Haemophilus spp testing is performed If zone diameter measurements are difficult to interpret another testing method should be used for Haemophilus spp x D x D BD has updated the instructions for use IFU to include the statement above The updated IFU is available at the following link https eifu bd com x D x D Important note x D x D Users who have received BD BBL Sensi Disc labelled with the sticker stating This product should not be used for the semi quantitative in vitro susceptibility testing of Haemophilus influenzae should not use the product for Haemophilus testing x D x D If your product does not contain the sticker then the product can be used for AST testing in line with the updated instructions for use x D | Actions for Clinical Users • Users should review and follow the guidance provided in this letter and in the updated IFU. • It is not necessary to review previous test results and no additional clinical actions are recommended. BD Actions: BD has implemented appropriate corrective and preventative measures to prevent recurrence. Customer Actions: • Review the information in Appendix 1 to determine if the BD BBL™ Sensi-Disc™ in your possession are impacted. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 28th June 2024. • Circulate this notice to all those who need to be aware within your organisation or to any organisation where the potentially affected product has been transferred. • If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: • Review the information in Appendix 1 and determine if the BD BBL™ Sensi-Disc™ in your possession are impacted. • Identify the facilities where you have distributed affected product and notify them immediately of this notice.  Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 28th June 2024. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. | 2024-06-05 11:07
Mirion Technologies (Capintec), Inc. | Captus 4000eThyroid Uptake System | 3 | - | Potential unexpected detachment of the collimator from Captus e arm x D | A hardware replacement component will be available soon which will correct the problem. In the interim, follow the preventative actions listed in the attachment. | 2024-06-06 11:00
Invivo Corporation | Sense XL Torso Coil | 4 | - | Philips has identified an issue in the SENSE XL Torso T and T coils which may result in localized heating during the scan resulting in potential harm to patients x D x D Philips has received complaints of coil heating events including reports of patient harm See section associated with this issue as of April x D | Philips is developing field corrections which include: 1. Warning label for coil and pad: Per your response form Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to apply a warning label to the coil and pad starting in January 2025. (Ref FCO78100607) 2. Software (SW) Update: Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to upgrade the SW as they are released. • Instruction For Use (IFU) Update: Including additional warnings and illustrations. • User interface: Updated with additional information and workflow support (cool-down periods) to enable safe usage of the coil during an examination. | 2024-06-06 13:23
Qiagen GmbH | QIAcube Connect MDx instrument | 4 | - | The heater shaker module does not perform heating if the temperature is set below C x D for static incubation steps when running these protocols | Product Defect Correction | 2024-06-06 13:42
Mako Surgical Corp. | RESTORIS MCK Patella Component | 1 | - | Stryker has discovered specific to the applications listed in Table an increase in the Software Error SE error code when a Mako System shutdown or a Mako System restart is not performed prior to switching between applications i e TKA to THA | Our records indicate that you may have received the affected product(s). It is Stryker’s responsibility as the manufacturer to ensure that customers who may have received these affected products also receive this important communication. We therefore request that you read this notice carefully and complete the following actions: 1. Please inform users of this Urgent Medical Device Correction and forward this notice to all individuals who need to be made aware. 2. Maintain awareness of this notice internally until all required actions have been completed within your facility. 3. Utilize the above risk mitigation steps to prevent the occurrence of software error #3 or to clear software error #3 if it is generated by the Mako Robotic System. 4. Inform Stryker if any of the subject devices have been distributed to other organizations. a. Please provide contact details so that Stryker can inform the recipients appropriately. b. If you are a Distributor, note that you are responsible for notifying your affected customers. 5. Please inform Stryker of any serious incidents concerning the use of the subject devices. a. Please comply with any local laws or regulations concerning the notification of serious incidents to your National Competent Authority. 6. Complete the attached customer response form. It may be that you no longer have any physical inventory on site. Completing this form will allow us to update our records and will also negate the need for us to send any further unnecessary communications on this matter. Therefore, please complete it even if you no longer have any of the subject devices in your physical inventory. 7. Return the completed form to your nominated Stryker Representative (indicated below) for this PFA. | 2024-06-06 14:42
SCHILLER Medical SAS | DEFIGARD HD-7 | 47 | - | Correction of software faults x D An automatic self test time stamp fault on can set off an alarm on the device without there being a technical fault x D In rare cases there are reports that when the device is switched on for manual defibrillation the defibrillator was in synchronized defibrillation mode instead of in direct mode x D | Customers/users are to follow the instructions provided by Schiller Medical in the attachment. Schiller Medical is releasing a new software version that corrects these faults. | 2024-06-09 11:48
SAHHA MEDICAL FACTORY COMPANY | SACO LUBRICANT GEL | 534700 | - | The medical device does not have Medical Device Marketing Authorization Therefore it can t ensure its safety efficacy and performance x D | Stop using and return all affected products to the manufacturer. | 2024-06-09 14:00
Smiths Medical ASD, Inc. | Portex ® Bivona® Silicone Tracheostomy Tube ADULT TRACHEOSTOMY TUBE | 2600 | - | Smiths Medical has identified that the securement flange of specific lots of the Bivona Neonatal Pediatric x D and Adult Tracheostomy products may tear because of a manufacturing defect See picture below for an x D example of a torn flange | Smiths Medical is sending this notification to all Bivona® customers who received products from Smiths Medical listed in Attachment 1_Affected Product. Smiths Medical will provide replacement product(s) or credit to affected customers upon receipt of a completed response form to certify product destruction. When using the device, all instructions, including warning and cautions contained in the Instructions for Use Documentation must be followed with heightened awareness. Please complete the following actions listed below 1) Check all inventory locations within your institution for the affected catalog numbers and lot numbers listed in the notification and discontinue use. Discard all affected products following your institution’s process for discarding. If discarding is not immediately possible at your facility, then the product should be quarantined until disposal. 2) Share this notification with all potential users of the device to ensure they are aware of this notification and proposed mitigations. If the devices are used at another location, please ensure this communication is delivered there. 3) Complete and return the attached Customer Response Form to EMEA-FSN@icumed.com within 10 days of receipt to acknowledge your understanding of this notification. | 2024-06-10 14:37
Philips Medizin Systeme Boeblingen GmbH | IntelliVue X3 | 2 | - | The problem and under what circumstances it can occur x D The Philips IntelliVue X Multi Measurement Module acquires multiple physiological patient signals displays measurement values waves trends generates physiological and technical alarms provides data recording and supports patient data management The device offers a monitoring solution optimized for the surgical cardiac medical and neonatal care environments IntelliVue X devices can be located in the patient s vicinity at the bedside or can be used in mobile applications during patient transport inside hospitals x D x D The IntelliVue Multi Measurement Module X with Option C has a hemo specific profile which is used when IntelliVue X is connected to Philips Hemodynamic Application One of the settings in hemo profile is the calibration setting of Invasive Blood Pressure x D This calibration setting has been inadvertently switched to OFF in the hemo profile with software version P preventing the user of the Philips Interventional Hemodynamic Application to be able to change the calibration factor of the reusable Invasive Blood Pressure transducer or to initiate a calibration procedure x D | Actions planned by Philips to correct the problem A Philips representative will contact you to schedule a visit from a Philips Field Service Engineer who will update the configuration profile on your IntelliVue X3 to correct this issue. | 2024-06-10 15:22
SECURED MEDICAL DIRECTION UK CO., LTD | SMD Disposable Infusion Set | 28675400 | - | The end user confused the use of the device between the generic IV set Gravity Feed and the specialized type x D Pump driven IV Set | The authorized representative in KSA will inform the affected end users of the safety alert and instructions on proper use, including that it is unsuitable for use with pump devices. The label will be updated to include the required statement, 'Gravity Feed Only,’ explaining the proper use. Then, the updated label will be submitted to the Medical Devices Marketing Authorization database. | 2024-06-10 16:30
Hill-Rom, Inc. | Centrella max mattress | 290 | - | The top cover of affected mattresses may delaminate x D x D | If bleach is occasionally used (for example as a disinfection against Clostridium difficile), remove any disinfectant residue prior to and after the use of bleach with a new or clean cloth soaked in tap water. Ensure all staff are trained in the protocol for the use of bleach during daily cleaning and as part of reprocessing between patients, by not using bleach as your primary cleaner/disinfectant. Identify if the top cover is damaged by examining for signs of abrasions, cuts, tears, or fluid ingress. If top cover damage is identified, discontinue the use of that mattress. If damage is due to the use of bleach, a replacement will be arranged by Hill Rom representative. Baxter will be updating the IFU to include additional instructions and warnings regarding the use of cleaning agents that include bleach. | 2024-06-10 16:53
Abbott | ™ DBS System (Model Numbers 6660, 6661, 6662, 6663) | 38 | - | product labeling EOS refers to IPG battery reaching end of life The shorter duration is due to IPG reaching ERI threshold x D later than expected it does not involve premature IPG battery depletion and estimates of overall battery life are not x D impacted To date there have been no reports of permanent harm to patients resulting from this issue x D From February through April there have been a total of four complaints received out of more than | his issue will continue to affect your patients for the foreseeable future. Abbott is working on an update to align the product labeling with actual device performance. In the meantime, the following recommendations are being provided: • When an ERI notification appears, use the following considerations to make a determination with your patients on timing of IPG replacement: o The duration between ERI threshold and EOS may be 45-55% shorter than estimates provided in the IPG Clinician’s Manual. (For example, if the expected duration between ERI threshold and EOS were calculated to be six (6) months using the IPG Clinician’s Manual, the actual expected duration would be approximately three (3) months.) o Battery life remaining from ERI to EOS is based on programmed stimulation parameters and patient usage and is unique to each patient. o Patients with higher energy programmed settings may have shorter duration between ERI and EOS than patients with lower energy programmed settings. • Recommended strategies for preventing lapse in therapy: o Schedule a replacement following ERI appearance using the guidance provided in this letter to determine the timing of IPG replacement. o Please inform your patients about this medical device correction letter | 2024-06-10 20:23
Abbott | Proclaim™ DRG Neurostimulation System | 62 | - | The purpose of the ERI is to provide advance warning of upcoming EOS so that surgery to replace the IPG can be x D scheduled The ERI notification first appears when the Patient Controller PC or Clinician Programmer CP connects with x D the IPG once the device has reached the ERI voltage threshold volts as the IPG battery is approaching EOS Once the x D ERI threshold is reached the following notification will appear on the CP and PC applications each time the IPG is x D connected prior to EOS including during Neurosphere Virtual Clinic sessions | The IPG Clinician’s Manual provides estimates of remaining battery life from the time at which the ERI threshold is reached to EOS based on programmed stimulation settings. The actual duration between the IPG reaching ERI threshold and EOS may be 45-55% shorter than the estimates provided. Estimates of overall device longevity are not impacted. The Proclaim IPGs will continue to safely deliver therapy from the time of first ERI appearance until the devices reach EOS. Patient Management Recommendations This issue will continue to affect your patients for the foreseeable future. Abbott is working on an update to align the product labeling with actual device performance. In the meantime, the following recommendations are being provided: • When an ERI notification appears, use the following considerations to make a determination with your patients on timing of IPG replacement: o The duration between ERI threshold and EOS may be 45-55% shorter than estimates provided in the IPG Clinician’s Manual. (For example, if the expected duration between ERI threshold and EOS were calculated to be six (6) months using the IPG Clinician’s Manual, the actual expected duration would be approximately three (3) months.) o Battery life remaining from ERI to EOS is based on programmed stimulation parameters and patient usage and is unique to each patient. o Patients with higher energy programmed settings may have shorter duration between ERI and EOS than patients with lower energy programmed settings. • Recommended strategies for preventing lapse in therapy: o Schedule a replacement following ERI appearance using the guidance provided in this letter to determine the timing of IPG replacement. o Please inform your patients about this medical device correction lette | 2024-06-10 20:31
Ortho Clinical Diagnostics Inc. | VITROS Chemistry Products FS Diluent Pack 3 | 11 | - | QuidelOrtho received complaints related to lower than expected results when testing patient samples known to have elevated CRP concentrations Our investigation determined that systems processing samples which are above the measuring range when processed neat greater than mg L or mg dL may generate results with a potential negative bias of up to approximately when diluted using Specialty Diluent from VITROS FS Diluent Pack Lot Please note that samples which were not diluted using Lot are unaffected by this issue | • Discontinue using, render unusable, and discard your remaining inventory of VITROS FS Diluent Pack 3, Lot 01-1645, including packs currently loaded on your VITROS System(s). • Complete and return the enclosed Confirmation of Receipt form no later than Month DD, 2024. Upon receipt of your completed Confirmation of Receipt form, QuidelOrtho will credit or replace your discarded inventory. • Save this notification with your User Documentation or post this notification by each VITROS 5,1 FS/4600/5600/XT 7600 System until the issue has been resolved. • Please forward this notification if the affected product was distributed outside of your facility. • If your laboratory has experienced the issue with this product and you have not already done so, please report the occurrence to your local Global Services Organization. | 2024-06-11 10:14
Canon Medical Systems Corporation | Vantage Elan | 4 | - | In the attached report Evaluation Report | [Actions that have been taken] For sites using the applicable systems, in order to eliminate the risk of power cable disconnection, emergency inspections have been conducted to confirm that the terminal block securing screws are correctly tightened. In addition, a notice of the actions to take if abnormalities occur has been issued. | 2024-06-12 13:13
Siemens Healthcare GmbH | ARTIS icono biplane | 3 | - | In very rare situations the first x ray release after patient registration might be performed with a wrong copper refiltration After an acquisition or significant changes in the projection or angulation the issue will automatically be resolved x D | Until the corrective action is implemented, the manufacturer recommend taking any action, which lead to a change of the prefiltration, prior to the first x-ray release after a patient registration. This could be the first acquisition, the change of the organ program or changed water values due to steeper angulation. A software update for the affected systems is planned to correct the issue and ensure the correct function of the prefiltration. As soon as the software update is available, Siemens service organization will get in contact with you for an appointment to perform the corrective action. | 2024-06-12 13:27
SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | 20 | - | Mindray has identified a potential issue associated with BeneHeart D Automated External Defibrillator It was found that BeneHeart D can support MR MR MR MR electrode pads and some customers may purchase more than one model of electrode pads | Advice on action to be taken by the Hospital administrator: 1. Please pass the FSN to anyone in your organization who needs to know this information, and to organizations or individuals in to whom the affected products may have been sold or transferred 2. You can continue to use the device when you receive this field safety notice (FSN) and the updated user manual. 3. Our local Mindray Service Representatives or cooperative distributors will contact you as soon as possible to provide the new version of the product manuals of the affected units. Advice on action to be taken by the distributor: 1. Please pass the FSN to anyone in your facility who needs to know this information, and to organizations or individuals to whom the affected products may have been distributed. 2. If any BeneHeart D1 Automated External Defibrillators in your facility is on the affected list, please sell these devices to customers with the updated manual. Mindray's Service Representatives will contact you to provide the new version of the product manual of the affected units. | 2024-06-13 09:48
Philips | Trilogy Evo International | 249 | - | The design requirements of the Trilogy Evo devices do not align with the requirements of the ISO standards The standards specify that the maximum delay from obstruction to alarm shall be no more than breath cycles or seconds whichever is greater However the current obstruction alarm delay is seconds which is a delay of seconds greater than required by the standards | Philips Respironics will release a software update to resolve this issue, aligning Obstruction Alarm trigger conditions with standard requirements. A separate notification will be provided when a software solution is available. | 2024-06-14 14:47
Philips | Trilogy Evo O2, International | 249 | - | The design requirements of the Trilogy Evo devices do not align with the requirements of the ISO standards The standards specify that the maximum delay from obstruction to alarm shall be no more than breath cycles or seconds whichever is greater However the current obstruction alarm delay is seconds which is a delay of seconds greater than required by the standards | Philips Respironics will release a software update to resolve this issue, aligning Obstruction Alarm trigger conditions with standard requirements. A separate notification will be provided when a software solution is available. | 2024-06-14 14:47
Philips | Trilogy EV300 INTL | 249 | - | The design requirements of the Trilogy Evo devices do not align with the requirements of the ISO standards The standards specify that the maximum delay from obstruction to alarm shall be no more than breath cycles or seconds whichever is greater However the current obstruction alarm delay is seconds which is a delay of seconds greater than required by the standards | Philips Respironics will release a software update to resolve this issue, aligning Obstruction Alarm trigger conditions with standard requirements. A separate notification will be provided when a software solution is available. | 2024-06-14 14:47
Biocartis | Idylla EGFR Mutation Test | 1 | - | If there will be a single positive S I positive call in the past months that have an EGFR total cq higher than then they should revert with this information and retest in another cartridge or in another patient x D x D x D | Change the IFU | 2024-06-20 16:03
Thermo Fisher Scientific | Cholesterol | 72 | - | It has been identified that the response level for Cholesterol catalog number lot WB x D and WA has been reduced The decreased response level has resulted in decreased linearity x D and creates a risk of falsely decreased patient results | 1. Please be aware that the above mentioned (Table 1) Thermo Fisher Scientific products are affected. 2. As appropriate, contact your Medical Professional for evaluation of further action. 3. Please stop using the affected Cholesterol lots WB28 and WA26 and discard the remaining stock of the affected Cholesterol lot. 4. Please contact your Thermo Fisher Scientific representative for free of charge replacement of discarded products within the scope of this Field Safety Corrective Action through your normal ordering channel. 5. Retain a copy of this letter for your laboratory records if appropriate. 6. Fill out the RESPONSE FORM and return it within 5 days of the date of the letter to your Thermo Fisher Scientific representative as instructed in the form. 7. Please contact your local Thermo Fisher Scientific representative for further information, if needed. | 2024-06-23 10:13
Thermo Fisher Scientific | CHOLESTEROL | 72 | - | It has been identified that the response level for Cholesterol catalog number lot WB and WA has been reduced The decreased response level has resulted in decreased linearity and creates a risk of falsely decreased patient results | 1. Please be aware that the above mentioned (Table 1) Thermo Fisher Scientific products are affected. 2. As appropriate, contact your Medical Professional for evaluation of further action. 3. Please stop using the affected Cholesterol lots WB28 and WA26 and discard the remaining stock of the affected Cholesterol lot. 4. Please contact your Thermo Fisher Scientific representative for free of charge replacement of discarded products within the scope of this Field Safety Corrective Action through your normal ordering channel. 5. Retain a copy of this letter for your laboratory records if appropriate. 6. Fill out the RESPONSE FORM and return it within 5 days of the date of the letter to your Thermo Fisher Scientific representative as instructed in the form. 7. Please contact your local Thermo Fisher Scientific representative for further information, if needed. | 2024-06-23 10:40
Medtronic Inc | ICM LNQ22 LINQ II | 43 | - | In August evidence of corrosion on the E electrode feedthrough was observed on a returned device with a reported complaint of oversensing noise The cause of the corrosion was due to moisture ingress under one electrode which is in contact with tissue and up through the via connecting the electrode to the interior of the ICM Oversensing noise can lead to a lack of reliable cardiac monitoring data and subsequently a delay in the treatment or diagnosis of the patient s condition and or need for a system revision As of Apr there have been returned devices from Bonder with confirmed corrosion of the feedthrough via No reports of harm have been received other than system revision e g repositioning explant x D x D Investigation of the issue identified a visual signature on the electrode visible only via acoustic microscopy which indicated delamination of Ti foil from the Ti sputter Penetrating dye testing confirmed a moisture pathway between the Ti foil and the Ti sputter It was determined the potential for creating this moisture pathway was related to the ZLSRBND referred to as Bonder laser bonder MTC s oldest laser bonder which was decommissioned on Sep PHO was initiated to control remaining products that had not been sold while the investigation continued x D x D During the investigation into the issue CareLink data from this population was reviewed to understand if the noise due to via corrosion was visible to or perceptible by a clinician Through that evaluation a tipping point pattern was seen in the daily ECG and fast ventricular tachycardia FVT metrics that showed a distinct transition from normal noise to major noise that could potentially impede a clinician s ability to read the patient s heart rhythm Further investigation into CareLink was conducted to identify other Bonder devices that exhibited a similar tipping point pattern to potentially indicate corrosion within devices still implanted Through that review a total of devices including confirmed via corrosion devices were identified leading to a projected failure rate of over the year LINQ II device lifetime across the population of the Bonder devices The overall patient safety risk associated with this issue was assessed as Medium Zone x D In December evidence of corrosion was again observed but this time on both the E and E electrode feedthroughs of two returned devices from the ZLSRBND referred to as Bonder laser bonder At the time the enabling mechanism of the corrosion could not be determined for these events through analysis In January additional events with oversensing allegations and via corrosion were identified There have been complaint devices from Bonder with returned product analysis confirming via corrosion as of this assessment At this time no reports of harm have been received other than system revision e g repositioning explant x D x D Due to the patient safety risk associated with the issue and knowledge of the enabling mechanism FA LINQ II Insertable Cardiac Monitoring Systems LNQ with Potential for Amplified Noise was approved with communications to a subset of customers in the field on Nov In addition to communications to affected consignees periodic searches for amplified noise patterns of the Bonder device population were initiated As of this assessment a total of devices from that population have been identified via CareLink to exhibit a tipping point pattern and was suspected to have via corrosion x D x D In December evidence of corrosion was again observed but this time on both the E and E electrode feedthroughs of two returned devices from the ZLSRBND referred to as Bonder laser bonder At the time the enabling mechanism of the corrosion could not be determined for these events through analysis In January additional events with oversensing allegations and via corrosion were identified There have been complaint devices from Bonder with returned product analysis confirming via corrosion as of this assessment At this time no reports of harm have been received other than system revision e g repositioning explant x D x D Further investigation into the performance of Bonder identified other signals indicative of a systemic risk of corrosion It was found that the bonder was operating in a region where the bonds were very dark and creating damage in the sapphire which could increase the risk of moisture gas ingress Studies into peel strength of the Ti foil have found that strength increases as the bond energy increases to a peak and then decreases past the peak as bond energy creates damage in the sapphire In addition preliminary CareLink data was reviewed to identify devices with high numbers of FVT and rejected FVT episodes which can potentially be indicative of via corrosion resulting in noise based on past experience with Bonder devices Relative to the other bonders Bonder exhibited an elevated rate of devices with noise particularly for devices bonded in January Due to these signals production was halted formal holds initiated and PHO was initiated to control remaining products that had not been sold while the investigation continued Consistent with the prior investigation into Bonder CareLink data from the Bonder population was reviewed and identified devices from Bonder that exhibited a tipping point in their device metrics x D Leveraging the current CareLink data across Bonder and Bonder projected failure rates at and months across the combined population were determined to be and respectively x D Due to the similar nature of clinical impact for both populations of devices the scope of the existing field action FA was extended to include the Bonder population The previous patient management recommendations provided for Bonder devices are also appropriate for Bonder devices x D x D | • Please review your inventory for devices with serial numbers listed in Table 1 of the FSN. Identify, quarantine, and return non-implanted devices. Your local Medtronic Representative can assist you as necessary. • Encourage enrollment in and regular transmissions to CareLink. o Medtronic will apply recurring algorithmic searches on CareLink for the noise pattern and notify the clinician if present. The information used to identify this pattern is not visible to the clinician through CareLink. No further action is required for patients regularly transmitting to CareLink. • For patients not followed in CareLink: o Consider whether enrolling in CareLink is an option, per HRS/EHRA/APHRS/LAHRS guidance. Ongoing CareLink monitoring will reduce the potential for episodes caused by amplified noise to overwrite true episodes before they are reviewed. o If noise interferes with the ability to assess the patient’s rhythm or reason for monitoring, contact your Medtronic Representative. • If the ICM is no longer in use, no further action is necessary. | 2024-06-24 08:14
Integra Lifesciences | CUSA® Excel C2600 23KHz Handpiece | - | - | through an internal investigation of customer complaints it was identified that there is a potential for the CUSA Excel KHz Handpiece housing to crack In order to mitigate this issue the devices will be reworked through our service and repair process for the units manufactured prior to November x D x D A replacement housing was introduced in November to fix the crack housing as such any new parts manufactured after November do not have the potential of cracked housing issue | Actions to be Taken by Distributor: 1. Please review and understand the information provided in this letter. 2. Review your inventory and determine if you have any of the serial numbers that are impacted. All serial numbers for the CUSA Excel 23KHz straight handpiece start with i.HA – indicating the product family followed by ii.Letter – indicating the month of manufacture where A = January; B = February; C = March etc noting that the letter I is not used. iii.Two digits e.g 24 indicating the year of manufacture with 24 representing 2024. iv.The next set of 3 digits are random serial numbers unique to each handpiece. v.The last 2 digits represent the model number. vi.IE indicates the location of manufacture. Example: Serial number HAL2300103IE represents a product that is manufactured in November 2023 and is acceptable for continued use. 3. If your handpiece was manufactured prior to HAL23XXXXXXIE (affected products) and you have it in your warehouse, you should take the following actions: a. Follow “actions to be taken by customers” (below) b. Check the box “I do have affected product” in the enclosed reply form. c. Record on the reply form the total quantity of affected products and serial numbers that you have. 4. If you do not have affected products in your warehouse, check the box, “I do not have affected products”. 5. Please check your customer traceability records for shipments of affected products. 6. Forward a copy of the enclosed Field Safety Notice to any of your customers that have purchased the affected products and ask them to follow the instructions under “Actions to be Taken by Customer” below. 7. The handpiece(s) should be returned to Integra S&R for service, email: TCS.Distributors@integralife.com, and there will be free replacement of the handpiece housing. | 2024-06-24 09:50
Löwenstein Medical SE & Co KG, Arzbacher Straße 80, 56130 Bad Ems, Germany | leoni plus | - | - | During ventilation with HFO switched on and the recruitment function active there is a possibility that a failsafe T x D is triggered This stops ventilation and the patient is not ventilated x D The recruitment function is active when a value is selected for FreqRec x D If it says OFF the function is deactivated x D This behavior can be triggered if the user starts ventilation with the recruitment function active The behavior only x D occurs when a user takes active action We are aware of one case in which this led to additional action by a user x D | Action to be taken by the operator/user: • Use the OFF (FreqRec) setting for the recruitment function. • Forward this safety message to all persons in your organization who use the device! • Inform you service provider or technician he must deactivate the recruitment function. • Confirm receipt of the letter using the form in Appendix A. Corrective actions to be taken by the sales and service partner: • Forward this safety information to operators, customers. • Deactivate the recruitment function in service mode. Preventive action to be taken by the manufacturer: • The malfunction will be fixed in a future version and will be offered at the end of 2024. | 2024-06-24 15:20
Philips | BiPAP A40 EFL International | 84 | - | The High Internal Oxygen alarm is intended to detect oxygen accumulation within the ventilator that may occur while supplemental oxygen is being provided to a patient Philips Respironics has identified issues in this oxygen s sensor manufacturing process that can cause the sensor to malfunction inaccurately reporting elevated oxygen levels to the device when elevated levels are not present When this happens it can cause the device to incorrectly detect elevated oxygen concentration even when supplemental oxygen is not connected to the device The device will continue to provide therapy while the alarm is addressed in accordance with the User Manual This issue can manifest itself in the following forms x D The device continuously raises the High Internal Oxygen alarm while supplemental oxygen is connected x D The device continuously raises the High Internal Oxygen alarm while supplemental oxygen is not connected | The User Manual instructs to disconnect the oxygen supply when the “High Internal Oxygen” alarm occurs. This may not clear the alarm if the sensor malfunctions. • All alarms must be responded to following the instructions provided by the device User Manual. • Any patient that is dependent on their device should have suitable backup therapies available in the event of a device malfunction. This includes situations where an alarm cannot be cleared. • Remove the patient from the device and switch to an alternative therapy device if you experience a “High Internal Oxygen” alarm that cannot be cleared. • Contact your device distributor/healthcare provider if you experience an alarm or malfunction that cannot be resolved following the device User Manual. Actions for Physicians: • Refer to Appendix A: Guidance for physicians/healthcare professionals related to FSN • Complete the response form attached if this came directly to you from Philips Respironics Actions for Patients and Users: • Follow these steps if “High Internal Oxygen” alarm occurs: For facility-based clinicians, if the High Oxygen Alarm occurs, immediately remove the patient from the device and connect them to an alternate source of ventilation. For home-based patients, if the High Oxygen Alarm occurs, immediately remove the device and connect to an alternative device, if available. Contact your home care equipment provider for service and/or an alternative device. • Refer to Appendix B: Guidance for end users/home care providers related to FSN. As an optional action, you may perform a “hard reboot” (forced device restart) that may temporarily restore device function. The details and instructions for performing this hard reboot are contained in Appendix C, shown below. tained in Appendix C, shown below. o After a hard reboot is performed, please confirm that the ventilator settings are correct. Philips Respironics is currently investigating this issue and will implement appropriate actions to prevent recurrence. | 2024-06-24 16:49
Philips | BiPAP A40 Pro International | 84 | - | The High Internal Oxygen alarm is intended to detect oxygen accumulation within the ventilator that may occur while supplemental oxygen is being provided to a patient Philips Respironics has identified issues in this oxygen s sensor manufacturing process that can cause the sensor to malfunction inaccurately reporting elevated oxygen levels to the device when elevated levels are not present When this happens it can cause the device to incorrectly detect elevated oxygen concentration even when supplemental oxygen is not connected to the device The device will continue to provide therapy while the alarm is addressed in accordance with the User Manual This issue can manifest itself in the following forms x D The device continuously raises the High Internal Oxygen alarm while supplemental oxygen is connected x D The device continuously raises the High Internal Oxygen alarm while supplemental oxygen is not connected | The User Manual instructs to disconnect the oxygen supply when the “High Internal Oxygen” alarm occurs. This may not clear the alarm if the sensor malfunctions. • All alarms must be responded to following the instructions provided by the device User Manual. • Any patient that is dependent on their device should have suitable backup therapies available in the event of a device malfunction. This includes situations where an alarm cannot be cleared. • Remove the patient from the device and switch to an alternative therapy device if you experience a “High Internal Oxygen” alarm that cannot be cleared. • Contact your device distributor/healthcare provider if you experience an alarm or malfunction that cannot be resolved following the device User Manual. Actions for Physicians: • Refer to Appendix A: Guidance for physicians/healthcare professionals related to FSN • Complete the response form attached if this came directly to you from Philips Respironics Actions for Patients and Users: • Follow these steps if “High Internal Oxygen” alarm occurs: For facility-based clinicians, if the High Oxygen Alarm occurs, immediately remove the patient from the device and connect them to an alternate source of ventilation. For home-based patients, if the High Oxygen Alarm occurs, immediately remove the device and connect to an alternative device, if available. Contact your home care equipment provider for service and/or an alternative device. • Refer to Appendix B: Guidance for end users/home care providers related to FSN. As an optional action, you may perform a “hard reboot” (forced device restart) that may temporarily restore device function. The details and instructions for performing this hard reboot are contained in Appendix C, shown below. tained in Appendix C, shown below. o After a hard reboot is performed, please confirm that the ventilator settings are correct. Philips Respironics is currently investigating this issue and will implement appropriate actions to prevent recurrence. | 2024-06-24 16:49
Olympus | HF units ESG-410 | - | - | Olympus is conducting this field correction after identifying an increase in complaints related to the ESG generator not powering on Olympus investigation has determined that this phenomenon is caused by a failure of electrical components on the generator s power supply However if the device has successfully powered on it will work within specification To address this issue Olympus will be replacing the current power supply with a new power supply part that is not susceptible to this issue | in the attached UFSN. | 2024-06-25 10:26
Olympus | HF units ESG-410 | 3 | - | Olympus is conducting this field correction after identifying an increase in complaints related to the ESG generator not powering on Olympus investigation has determined that this phenomenon is caused by a failure of electrical components on the generator s power supply However if the device has successfully powered on it will work within specification To address this issue Olympus will be replacing the current power supply with a new power supply part that is not susceptible to this issue x D Olympus has identified forty seven complaints for the ESG generator not powering on that may have resulted from this issue No serious injuries or reportable malfunctions have been reported in relation to this issue | in the attached UFSN | 2024-06-25 10:40
B. Braun Melsungen AG | Spaceplus Infusomat® | 141 | - | It has been identified that a wave spring installed in Spaceplus Infusomat infusion pumps may have reduced spring force in a certain batch due to deformation in some cases Consequently the occlusion pressure of the peristaltic may be less than specified x D | Please identify affected pumps and adjust the electronic pressure cut off to ≤ P5 (IFU section 9.11.2) until the pump has been repaired. Alternatively, take the affected device out of operation if an operation with the mentioned limitation is not feasible. You will be contacted by the local B. Braun representative to arrange for replacement of the wave spring. | 2024-06-25 12:19
Stryker Neurovascular | Synchro2 Soft | 12 | - | Stryker Neurovascular has distributed multiple lots of Trevo Retrievers Trevo x D Microcatheters FlowGate DAC and Synchro devices that did not meet the minimum x D sample size required for endotoxin testing during manufacturing processes and are x D out of compliance with American National Standard Bacterial endotoxins Test x D methods routine monitoring and alternatives to batch testing AAMI ST standard x D These lots of Stryker Neurovascular devices may exceed the acceptable limits of x D endotoxin levels | Stryker Neurovascular is notifying all impacted customers to return devices within scope of this field action and to perform the actions listed below. Replacement units are available. Stryker Neurovascular is implementing process changes to prevent distribution of devices not meeting the minimum sample size required for endotoxin testing during manufacturing. | 2024-06-25 12:19
Stryker Neurovascular | Trevo® XP ProVue Retriever | 12 | - | Stryker Neurovascular has distributed multiple lots of Trevo Retrievers Trevo x D Microcatheters FlowGate DAC and Synchro devices that did not meet the minimum x D sample size required for endotoxin testing during manufacturing processes and are x D out of compliance with American National Standard Bacterial endotoxins Test x D methods routine monitoring and alternatives to batch testing AAMI ST standard x D These lots of Stryker Neurovascular devices may exceed the acceptable limits of x D endotoxin levels | Stryker Neurovascular is notifying all impacted customers to return devices within scope of this field action and to perform the actions listed below. Replacement units are available. Stryker Neurovascular is implementing process changes to prevent distribution of devices not meeting the minimum sample size required for endotoxin testing during manufacturing. | 2024-06-25 12:19
Abbott GmbH. & Co. KG | Alinity hq Analyzer | 155 | - | The following issues has been identified when using the above Alinity hq Analyzer x D Issue When processing a CBC DIFF sample that immediately follows a CBC DIFF RETIC sample diluent may backflow into the RBC dilution cup x D Issue An overestimation of the basophil BASO count may occur on some samples when cell events are incorrectly counted as BASO during sample analysis This may result in falsely increased BASO counts and B BASO x D | Please follow necessary actions to be taken by customer that provided in the attachment. You will be contacted by Abbott representative for scheduling an upgrade of your Alinity hq Analyzer to install software version 5.8 once its available. | 2024-06-25 12:41
Löwenstein Medical SE & Co KG, Arzbacher Straße 80, 56130 Bad Ems, Germany | leoni plus ( HVO and transport ) | - | - | During ventilation with HFO switched on and the recruitment function active there is a possibility that a failsafe T x D is triggered This stops ventilation and the patient is not ventilated x D The recruitment function is active when a value is selected for FreqRec x D If it says OFF the function is deactivated x D This behavior can be triggered if the user starts ventilation with the recruitment function active The behavior only x D occurs when a user takes active action We are aware of one case in which this led to additional action by a user | Action to be taken by the operator/user: • Use the OFF (FreqRec) setting for the recruitment function. • Forward this safety message to all persons in your organization who use the device! • Inform you service provider or technician he must deactivate the recruitment function. • Confirm receipt of the letter using the form in Appendix A. Corrective actions to be taken by the sales and service partner: • Forward this safety information to operators, customers. • Deactivate the recruitment function in service mode. Preventive action to be taken by the manufacturer: • The malfunction will be fixed in a future version and will be offered at the end of 2024. | 2024-06-25 13:41
Shenzhen Mindray Bio-Medical Electronics Co., Ltd. | Automated External Defibrillator | 1 | - | Incorrect function configurations | Advise on action to be taken by the CUSTOMER / USER: 1. Please pass the Field Safety Notice (FSN) to all those who need to be aware within your organization or to any organization where the affected BeneHeart C2 Automated External Defibrillators have been delivered. 2. After you receive this Field Safety Notice, please confirm that if you have any BeneHeart C2 Automated External Defibrillators on the affected list, please discontinue use of the device and contact us. Mindray service team will repair the problem as soon as possible. 3. Return the Field Safety Notice Acknowledgement Form to Mindray confirming that the Field Safety Notice has been read and understood. Advise on action to be taken by the distributor: 1. Please pass the Field Safety Notice (FSN) to all those who need to be aware within your organization or to any organization where the affected BeneHeart C2 Automated External Defibrillators have been delivered. 2. If your organisation has any BeneHeart C2 Automated External Defibrillators on the affected list, please do not sell or install these devices to your customers. Mindray's service team will contact you as soon as possible to provide software and upgrade instructions. If you have any questions regarding the software upgrade process, please contact our local Mindray service team. 3. Return the Field Safety Notice Acknowledgement Form to Mindray confirming that the Field Safety Notice has been read and understood. | 2024-06-25 14:46
Baxter Medical Systems GmbH + Co. KG | iLED 7 Ceiling Single | 137 | - | Many factors may increase the risk of thermal injuries to the skin organs and tissue when using LED overhead surgical lights The risk increases when multiple LED overhead surgical lights are overlapped and set at a high intensity during long exposure times Thermal injuries may include skin burns and drying of tissue that may lead to scarring infection and or internal organ damage Baxter has received reports of serious injury associated with this issue x D x D Baxter will be updating the IFU to include various situations that can result in thermal skin injury as well as highlighting residual risks due to surgical overhead light usage An updated IFU will be provided once available Additionally Baxter is currently working on a product resolution and will provide additional information to customers when available x D | 1. Operators may continue to use the iLED 7 surgical light systems by following the current instructions for use (IFU) cautions and warnings while considering the additional information below : • The lowest possible illumination level suitable for the procedure to be utilized, especially in certain neurological or intestinal procedures on delicate, thin, dry, or abnormal tissue. • Overlapping light fields that are set at high intensity (80% or higher) is to be avoided. - If the light intensity of one light-head is set at 80% or higher the second light-head, if overlapping, should be set at 50% or lower. - If the light intensity of two or more light heads is set at 80% or higher, to minimize risk should not overlap. • If a very high-intensity setting is temporarily required, operator is to reduce intensity as soon as the need passes. • Operator is to ensure the Adaptive Light Control (ALC) Plus sensor is activated and indicated on the wall or mobile control panel. If activation of the ALC Plus is not possible, operator should contact a technical service representative for further instructions. • If the ALC Plus is switched off, the light-head needs to be positioned at 100 centimeters (39.37 inches) from the surgical field. At other distances the illuminance and irradiance limits can be exceeded, which may result in thermal skin injuries. 2. Additional safety information reviewing the potential risks of using an overhead surgical light, including educational safety content is available through Pfiedler education (a division of AORN). 3. Customers are asked to acknowledge receipt of the Field Safety Notice (using the Customer Reply Form) 4. This information is to be provided to all users of the iLED 7 surgical light. If this product is to be distributed to other facilities or departments, customers are asked to forward a copy of this communication to them in accordance with their customary procedures. | 2024-06-25 15:19
Baxter Medical Systems GmbH + Co. KG | iLED 7 Mobile | 137 | - | Many factors may increase the risk of thermal injuries to the skin organs and tissue when using LED overhead surgical lights The risk increases when multiple LED overhead surgical lights are overlapped and set at a high intensity during long exposure times Thermal injuries may include skin burns and drying of tissue that may lead to scarring infection and or internal organ damage Baxter has received reports of serious injury associated with this issue x D x D Baxter will be updating the IFU to include various situations that can result in thermal skin injury as well as highlighting residual risks due to surgical overhead light usage An updated IFU will be provided once available Additionally Baxter is currently working on a product resolution and will provide additional information to customers when available x D | 1. Operators may continue to use the iLED 7 surgical light systems by following the current instructions for use (IFU) cautions and warnings while considering the additional information below : • The lowest possible illumination level suitable for the procedure to be utilized, especially in certain neurological or intestinal procedures on delicate, thin, dry, or abnormal tissue. • Overlapping light fields that are set at high intensity (80% or higher) is to be avoided. - If the light intensity of one light-head is set at 80% or higher the second light-head, if overlapping, should be set at 50% or lower. - If the light intensity of two or more light heads is set at 80% or higher, to minimize risk should not overlap. • If a very high-intensity setting is temporarily required, operator is to reduce intensity as soon as the need passes. • Operator is to ensure the Adaptive Light Control (ALC) Plus sensor is activated and indicated on the wall or mobile control panel. If activation of the ALC Plus is not possible, operator should contact a technical service representative for further instructions. • If the ALC Plus is switched off, the light-head needs to be positioned at 100 centimeters (39.37 inches) from the surgical field. At other distances the illuminance and irradiance limits can be exceeded, which may result in thermal skin injuries. 2. Additional safety information reviewing the potential risks of using an overhead surgical light, including educational safety content is available through Pfiedler education (a division of AORN). 3. Customers are asked to acknowledge receipt of the Field Safety Notice (using the Customer Reply Form) 4. This information is to be provided to all users of the iLED 7 surgical light. If this product is to be distributed to other facilities or departments, customers are asked to forward a copy of this communication to them in accordance with their customary procedures. | 2024-06-25 15:19
Baxter Medical Systems GmbH + Co. KG | iLED 7 Pendant | 137 | - | Many factors may increase the risk of thermal injuries to the skin organs and tissue when using LED overhead surgical lights The risk increases when multiple LED overhead surgical lights are overlapped and set at a high intensity during long exposure times Thermal injuries may include skin burns and drying of tissue that may lead to scarring infection and or internal organ damage Baxter has received reports of serious injury associated with this issue x D x D Baxter will be updating the IFU to include various situations that can result in thermal skin injury as well as highlighting residual risks due to surgical overhead light usage An updated IFU will be provided once available Additionally Baxter is currently working on a product resolution and will provide additional information to customers when available x D | 1. Operators may continue to use the iLED 7 surgical light systems by following the current instructions for use (IFU) cautions and warnings while considering the additional information below : • The lowest possible illumination level suitable for the procedure to be utilized, especially in certain neurological or intestinal procedures on delicate, thin, dry, or abnormal tissue. • Overlapping light fields that are set at high intensity (80% or higher) is to be avoided. - If the light intensity of one light-head is set at 80% or higher the second light-head, if overlapping, should be set at 50% or lower. - If the light intensity of two or more light heads is set at 80% or higher, to minimize risk should not overlap. • If a very high-intensity setting is temporarily required, operator is to reduce intensity as soon as the need passes. • Operator is to ensure the Adaptive Light Control (ALC) Plus sensor is activated and indicated on the wall or mobile control panel. If activation of the ALC Plus is not possible, operator should contact a technical service representative for further instructions. • If the ALC Plus is switched off, the light-head needs to be positioned at 100 centimeters (39.37 inches) from the surgical field. At other distances the illuminance and irradiance limits can be exceeded, which may result in thermal skin injuries. 2. Additional safety information reviewing the potential risks of using an overhead surgical light, including educational safety content is available through Pfiedler education (a division of AORN). 3. Customers are asked to acknowledge receipt of the Field Safety Notice (using the Customer Reply Form) 4. This information is to be provided to all users of the iLED 7 surgical light. If this product is to be distributed to other facilities or departments, customers are asked to forward a copy of this communication to them in accordance with their customary procedures. | 2024-06-25 15:19
Baxter Medical Systems GmbH + Co. KG | iLED 7 Ceiling Duo | 137 | - | Many factors may increase the risk of thermal injuries to the skin organs and tissue when using LED overhead surgical lights The risk increases when multiple LED overhead surgical lights are overlapped and set at a high intensity during long exposure times Thermal injuries may include skin burns and drying of tissue that may lead to scarring infection and or internal organ damage Baxter has received reports of serious injury associated with this issue x D x D Baxter will be updating the IFU to include various situations that can result in thermal skin injury as well as highlighting residual risks due to surgical overhead light usage An updated IFU will be provided once available Additionally Baxter is currently working on a product resolution and will provide additional information to customers when available x D | 1. Operators may continue to use the iLED 7 surgical light systems by following the current instructions for use (IFU) cautions and warnings while considering the additional information below : • The lowest possible illumination level suitable for the procedure to be utilized, especially in certain neurological or intestinal procedures on delicate, thin, dry, or abnormal tissue. • Overlapping light fields that are set at high intensity (80% or higher) is to be avoided. - If the light intensity of one light-head is set at 80% or higher the second light-head, if overlapping, should be set at 50% or lower. - If the light intensity of two or more light heads is set at 80% or higher, to minimize risk should not overlap. • If a very high-intensity setting is temporarily required, operator is to reduce intensity as soon as the need passes. • Operator is to ensure the Adaptive Light Control (ALC) Plus sensor is activated and indicated on the wall or mobile control panel. If activation of the ALC Plus is not possible, operator should contact a technical service representative for further instructions. • If the ALC Plus is switched off, the light-head needs to be positioned at 100 centimeters (39.37 inches) from the surgical field. At other distances the illuminance and irradiance limits can be exceeded, which may result in thermal skin injuries. 2. Additional safety information reviewing the potential risks of using an overhead surgical light, including educational safety content is available through Pfiedler education (a division of AORN). 3. Customers are asked to acknowledge receipt of the Field Safety Notice (using the Customer Reply Form) 4. This information is to be provided to all users of the iLED 7 surgical light. If this product is to be distributed to other facilities or departments, customers are asked to forward a copy of this communication to them in accordance with their customary procedures. | 2024-06-25 15:19
Baxter Medical Systems GmbH + Co. KG | iLED 7 Ceiling Trio | 137 | - | Many factors may increase the risk of thermal injuries to the skin organs and tissue when using LED overhead surgical lights The risk increases when multiple LED overhead surgical lights are overlapped and set at a high intensity during long exposure times Thermal injuries may include skin burns and drying of tissue that may lead to scarring infection and or internal organ damage Baxter has received reports of serious injury associated with this issue x D x D Baxter will be updating the IFU to include various situations that can result in thermal skin injury as well as highlighting residual risks due to surgical overhead light usage An updated IFU will be provided once available Additionally Baxter is currently working on a product resolution and will provide additional information to customers when available x D | 1. Operators may continue to use the iLED 7 surgical light systems by following the current instructions for use (IFU) cautions and warnings while considering the additional information below : • The lowest possible illumination level suitable for the procedure to be utilized, especially in certain neurological or intestinal procedures on delicate, thin, dry, or abnormal tissue. • Overlapping light fields that are set at high intensity (80% or higher) is to be avoided. - If the light intensity of one light-head is set at 80% or higher the second light-head, if overlapping, should be set at 50% or lower. - If the light intensity of two or more light heads is set at 80% or higher, to minimize risk should not overlap. • If a very high-intensity setting is temporarily required, operator is to reduce intensity as soon as the need passes. • Operator is to ensure the Adaptive Light Control (ALC) Plus sensor is activated and indicated on the wall or mobile control panel. If activation of the ALC Plus is not possible, operator should contact a technical service representative for further instructions. • If the ALC Plus is switched off, the light-head needs to be positioned at 100 centimeters (39.37 inches) from the surgical field. At other distances the illuminance and irradiance limits can be exceeded, which may result in thermal skin injuries. 2. Additional safety information reviewing the potential risks of using an overhead surgical light, including educational safety content is available through Pfiedler education (a division of AORN). 3. Customers are asked to acknowledge receipt of the Field Safety Notice (using the Customer Reply Form) 4. This information is to be provided to all users of the iLED 7 surgical light. If this product is to be distributed to other facilities or departments, customers are asked to forward a copy of this communication to them in accordance with their customary procedures. | 2024-06-25 15:19
Baxter Medical Systems GmbH + Co. KG | iLED 7 Ceiling Quad | 137 | - | Many factors may increase the risk of thermal injuries to the skin organs and tissue when using LED overhead surgical lights The risk increases when multiple LED overhead surgical lights are overlapped and set at a high intensity during long exposure times Thermal injuries may include skin burns and drying of tissue that may lead to scarring infection and or internal organ damage Baxter has received reports of serious injury associated with this issue x D x D Baxter will be updating the IFU to include various situations that can result in thermal skin injury as well as highlighting residual risks due to surgical overhead light usage An updated IFU will be provided once available Additionally Baxter is currently working on a product resolution and will provide additional information to customers when available x D | 1. Operators may continue to use the iLED 7 surgical light systems by following the current instructions for use (IFU) cautions and warnings while considering the additional information below : • The lowest possible illumination level suitable for the procedure to be utilized, especially in certain neurological or intestinal procedures on delicate, thin, dry, or abnormal tissue. • Overlapping light fields that are set at high intensity (80% or higher) is to be avoided. - If the light intensity of one light-head is set at 80% or higher the second light-head, if overlapping, should be set at 50% or lower. - If the light intensity of two or more light heads is set at 80% or higher, to minimize risk should not overlap. • If a very high-intensity setting is temporarily required, operator is to reduce intensity as soon as the need passes. • Operator is to ensure the Adaptive Light Control (ALC) Plus sensor is activated and indicated on the wall or mobile control panel. If activation of the ALC Plus is not possible, operator should contact a technical service representative for further instructions. • If the ALC Plus is switched off, the light-head needs to be positioned at 100 centimeters (39.37 inches) from the surgical field. At other distances the illuminance and irradiance limits can be exceeded, which may result in thermal skin injuries. 2. Additional safety information reviewing the potential risks of using an overhead surgical light, including educational safety content is available through Pfiedler education (a division of AORN). 3. Customers are asked to acknowledge receipt of the Field Safety Notice (using the Customer Reply Form) 4. This information is to be provided to all users of the iLED 7 surgical light. If this product is to be distributed to other facilities or departments, customers are asked to forward a copy of this communication to them in accordance with their customary procedures. | 2024-06-25 15:19
GE Healthcare (Tianjin) Company Limited | Signa Prime | 121 | - | GE HealthCare has become aware that for certain MR systems the system predicted B RMS value can exceed the B RMS user prescribed limit when scanning in Low SAR Mode and the following two conditions are met x D D FSE T FLAIR OR T FLAIR Propeller imaging sequence is selected x D AND x D the Optimized T FLAIR sequence option is turned off OR not included in the MR configuration x D If this occurs it can result in overheating of an MR conditional implant x D There have been no injuries reported to GE HealthCare as a result of this issue | All affected customers will receive a Field Safety Notice and all affected products will be corrected via a software upgrade. | 2024-06-25 17:36
GE Healthcare (Tianjin) Company Limited | Signa Artist | 121 | - | GE HealthCare has become aware that for certain MR systems the system predicted B RMS value can exceed the B RMS user prescribed limit when scanning in Low SAR Mode and the following two conditions are met x D D FSE T FLAIR OR T FLAIR Propeller imaging sequence is selected x D AND x D the Optimized T FLAIR sequence option is turned off OR not included in the MR configuration x D If this occurs it can result in overheating of an MR conditional implant x D There have been no injuries reported to GE HealthCare as a result of this issue | All affected customers will receive a Field Safety Notice and all affected products will be corrected via a software upgrade. | 2024-06-25 17:36
GE Healthcare (Tianjin) Company Limited | Signa Voyager | 121 | - | GE HealthCare has become aware that for certain MR systems the system predicted B RMS value can exceed the B RMS user prescribed limit when scanning in Low SAR Mode and the following two conditions are met x D D FSE T FLAIR OR T FLAIR Propeller imaging sequence is selected x D AND x D the Optimized T FLAIR sequence option is turned off OR not included in the MR configuration x D If this occurs it can result in overheating of an MR conditional implant x D There have been no injuries reported to GE HealthCare as a result of this issue | All affected customers will receive a Field Safety Notice and all affected products will be corrected via a software upgrade. | 2024-06-25 17:36
GE Healthcare (Tianjin) Company Limited | Signa Artist Evo | 121 | - | GE HealthCare has become aware that for certain MR systems the system predicted B RMS value can exceed the B RMS user prescribed limit when scanning in Low SAR Mode and the following two conditions are met x D D FSE T FLAIR OR T FLAIR Propeller imaging sequence is selected x D AND x D the Optimized T FLAIR sequence option is turned off OR not included in the MR configuration x D If this occurs it can result in overheating of an MR conditional implant x D There have been no injuries reported to GE HealthCare as a result of this issue | All affected customers will receive a Field Safety Notice and all affected products will be corrected via a software upgrade. | 2024-06-25 17:36
GE Healthcare (Tianjin) Company Limited | Signa Creator | 121 | - | GE HealthCare has become aware that for certain MR systems the system predicted B RMS value can exceed the B RMS user prescribed limit when scanning in Low SAR Mode and the following two conditions are met x D D FSE T FLAIR OR T FLAIR Propeller imaging sequence is selected x D AND x D the Optimized T FLAIR sequence option is turned off OR not included in the MR configuration x D If this occurs it can result in overheating of an MR conditional implant x D There have been no injuries reported to GE HealthCare as a result of this issue | All affected customers will receive a Field Safety Notice and all affected products will be corrected via a software upgrade. | 2024-06-25 17:36
GE Healthcare (Tianjin) Company Limited | Signa Explorer | 121 | - | GE HealthCare has become aware that for certain MR systems the system predicted B RMS value can exceed the B RMS user prescribed limit when scanning in Low SAR Mode and the following two conditions are met x D D FSE T FLAIR OR T FLAIR Propeller imaging sequence is selected x D AND x D the Optimized T FLAIR sequence option is turned off OR not included in the MR configuration x D If this occurs it can result in overheating of an MR conditional implant x D There have been no injuries reported to GE HealthCare as a result of this issue | All affected customers will receive a Field Safety Notice and all affected products will be corrected via a software upgrade. | 2024-06-25 17:36
GE Medical Systems, LLC | Signa Hero | 121 | - | GE HealthCare has become aware that for certain MR systems the system predicted B RMS value can exceed the B RMS user prescribed limit when scanning in Low SAR Mode and the following two conditions are met x D D FSE T FLAIR OR T FLAIR Propeller imaging sequence is selected x D AND x D the Optimized T FLAIR sequence option is turned off OR not included in the MR configuration x D If this occurs it can result in overheating of an MR conditional implant x D There have been no injuries reported to GE HealthCare as a result of this issue | All affected customers will receive a Field Safety Notice and all affected products will be corrected via a software upgrade. | 2024-06-25 17:36
GE Medical Systems, LLC | Signa Pioneer | 121 | - | GE HealthCare has become aware that for certain MR systems the system predicted B RMS value can exceed the B RMS user prescribed limit when scanning in Low SAR Mode and the following two conditions are met x D D FSE T FLAIR OR T FLAIR Propeller imaging sequence is selected x D AND x D the Optimized T FLAIR sequence option is turned off OR not included in the MR configuration x D If this occurs it can result in overheating of an MR conditional implant x D There have been no injuries reported to GE HealthCare as a result of this issue | All affected customers will receive a Field Safety Notice and all affected products will be corrected via a software upgrade. | 2024-06-25 17:36
GE Medical Systems, LLC | Signa Architect | 121 | - | GE HealthCare has become aware that for certain MR systems the system predicted B RMS value can exceed the B RMS user prescribed limit when scanning in Low SAR Mode and the following two conditions are met x D D FSE T FLAIR OR T FLAIR Propeller imaging sequence is selected x D AND x D the Optimized T FLAIR sequence option is turned off OR not included in the MR configuration x D If this occurs it can result in overheating of an MR conditional implant x D There have been no injuries reported to GE HealthCare as a result of this issue | All affected customers will receive a Field Safety Notice and all affected products will be corrected via a software upgrade. | 2024-06-25 17:36
GE Medical Systems, LLC | Signa Premier | 121 | - | GE HealthCare has become aware that for certain MR systems the system predicted B RMS value can exceed the B RMS user prescribed limit when scanning in Low SAR Mode and the following two conditions are met x D D FSE T FLAIR OR T FLAIR Propeller imaging sequence is selected x D AND x D the Optimized T FLAIR sequence option is turned off OR not included in the MR configuration x D If this occurs it can result in overheating of an MR conditional implant x D There have been no injuries reported to GE HealthCare as a result of this issue | All affected customers will receive a Field Safety Notice and all affected products will be corrected via a software upgrade. | 2024-06-25 17:36
GE Medical Systems, LLC | SIGNA PET/MR | 121 | - | GE HealthCare has become aware that for certain MR systems the system predicted B RMS value can exceed the B RMS user prescribed limit when scanning in Low SAR Mode and the following two conditions are met x D D FSE T FLAIR OR T FLAIR Propeller imaging sequence is selected x D AND x D the Optimized T FLAIR sequence option is turned off OR not included in the MR configuration x D If this occurs it can result in overheating of an MR conditional implant x D There have been no injuries reported to GE HealthCare as a result of this issue | All affected customers will receive a Field Safety Notice and all affected products will be corrected via a software upgrade. | 2024-06-25 17:36
Stryker Neurovascular | Trevo® XP ProVue Retriever | 117 | - | Stryker Neurovascular has distributed multiple lots of Trevo Retrievers Trevo x D Microcatheters FlowGate DAC and Synchro devices that did not meet the minimum x D sample size required for endotoxin testing during manufacturing processes and are x D out of compliance with American National Standard Bacterial endotoxins Test x D methods routine monitoring and alternatives to batch testing AAMI ST standard x D These lots of Stryker Neurovascular devices may exceed the acceptable limits of x D endotoxin levels | Stryker Neurovascular is notifying all impacted customers to return devices within scope of this field action and to perform the actions listed below. Replacement units are available. Stryker Neurovascular is implementing process changes to prevent distribution of devices not meeting the minimum sample size required for endotoxin testing during manufacturing | 2024-06-26 13:56
Stryker Neurovascular | Synchro2 Soft | 2458 | - | Stryker Neurovascular has observed an increased frequency of PTFE coating x D damage occurring on the Synchro Guidewires that may be caused by the practice x D of backloading the guidewire through the optional introducer accessory This issue x D is limited to certain lots of the Synchro Guidewire listed in the Product Affected x D section above that contain an older version of the introducer accessory Due to x D variation in the manufacturing process of the supplier of the introducer accessory x D certain lots of introducers have sharper than intended edges that can peel off the x D PTFE coating when physicians use a technique known as backloading Users with x D impacted product in their inventory are cautioned not to use this backloading x D technique Refer to Appendix A for more information on the backloading x D technique | Stryker Neurovascular improved the optional introducer accessory packaged with the Synchro Guidewires. The improvement was made to reduce the likelihood of the introducer causing damage to the PTFE coating on the guidewire. This communication is for users who have product in their inventory which was manufactured prior to the improvement as listed in the “Product Affected” section above. | 2024-06-26 14:26
Megadyne Medical Product, Inc. | Mega Soft Patient Return Electrode Pad | 2626 | - | Megadyne Medical Products Inc Megadyne has received reports of patient burns identified after surgical procedures in which Mega Soft pads were used including product codes and Product codes and were approved for use in patients over lbs however to mitigate the potential risk to health in the population of children under years of age Megadyne is taking this corrective action to now limit use to patients age years or older x D Users should continue to follow the current Mega Soft Instructions for Use IFU except for this new limitation in population of intended use We will notify customers if we identify any additional actions that may help to ensure safe use of the products | 1. Share this notification update with all users of Mega Soft and Mega 2000 pads. 2. Confirm that personnel using the Mega Soft and Mega 2000 pads understand the intended use is changing to patients aged 12 years and older. 3. Post a copy of this communication to remind staff not to use the Mega Soft and Mega 2000 pads on patients under 12 years old. Although the current Mega Soft and Mega 2000 have printing of > 25 lbs (>11.3kg) on the pads, they should only be used for patients aged 12 years and older, and over labeling of the pad is not required. 4. If any subject product has been forwarded to another facility, contact that facility to share this information. Please share a copy of this notification when communicating. 5. Complete the Business Reply Form (BRF) Attachment 1 confirming receipt of this notice and fax or email it to [HAlzamel@ITS.JNJ.com] within three (3) business days. 6. As a reminder, it is important to follow proper cleaning, placement and setup steps for the Mega Soft and Mega 2000 pads. Failure to follow the Mega Soft and Mega 2000 pad IFU may contribute to patient burns. 7. If you need additional copies of this communication or have questions about returning the BRF, please contact HAlzamel@ITS.JNJ.com] and reference Event # 2370372. As with any medical device, adverse reactions or quality problems experienced with the use of this product should be reported to your Sales Representative, directly to Megadyne, or your National Health Authority. If medical engagement is requested, please have the Healthcare Provider submit the request using the Medical Information Request website: https://www.jnjmedtech.com/mir FA: 2370372 | 2024-06-26 18:09
Randox Laboratories Ltd | Rx Imola | - | - | The RX Imola parameters have been updated in Parameter version v The rerun settings Sample x D Rerun High Prozone and Rerun Low and dilutions used for Auto reruns have been disabled Randox x D are in the process of generating and updating existing data to comply with In vitro Diagnostic x D Regulations IVDR and will provide further updates on auto rerun settings in due course For samples x D above the assay range a manual dilution should be performed and results multiplied by the dilution x D factor | Action to be taken: Download and update parameters, or manually enter updates to avoid losing any previously obtained data, for assays performed on your Rx Imola. Screenshots are available upon request for reference, please contact technical.services@randox.com for more information. Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. If you have any questions or concerns please contact Randox Technical Services. | 2024-06-27 15:46
Covidien llc | Puritan Bennett™ 560 Ventilator - Puritan Bennett™ 520 Ventilator | 500 | - | Medtronic investigated the potential patient impact of exposure to Polyester Polyurethane PE PUR foam used for sound abatement in the PB Series Ventilators According to published literature the PE PUR foam may degrade over time releasing particulates and or Volatile Organic Compounds VOCs which can enter the gas pathway Degradation may be accelerated following exposure to elevated temperature and humidity conditions Patient exposure was assessed taking into account clinical use presence of ventilator bias flow and less than lifetime year exposure and was validated by external toxicology reports x D The toxicological risk assessment evaluated the patient daily dose of VOCs and compared this to defined tolerable exposures for those VOC compounds Per the reanalysis of VOC data this comparison expressed as a margin of safety MOS was found to be acceptable in both the infant paediatric and adult populations for all VOCs carcinogenic and non carcinogenic endpoints in devices with new and field aged foam That is with a MOS in the patient population the daily exposure to VOC compounds is less than the total amount that is considered to be without appreciable risk of harm For VOCs associated with a potential carcinogenicity hazard the MOS assessment was performed assuming a cancer risk level of which indicates increased cancer risk of at most x D The PB series of ventilators are labelled for years of service life Patients may discontinue use of this device prior to the year labeled service life if they no longer require ventilation are switched to a different manufacturer s device or succumb to their underlying illness before years of use has elapsed Some patients requiring chronic ventilation may receive ventilatory support for more than years these patients will be transitioned to an alternate PB series ventilator over the course of their illness therefore it is conceivable that some patients may cumulatively receive ventilation from a PB series ventilator for a period of time longer than years The vast majority of patients require less than years of home mechanical ventilation although there is evidence that a small minority require longer periods No patient has experienced more than years of use for any PB ventilator as the device was released to market in x D | Please complete and return the enclosed Customer Acknowledgement Form as directed to confirm the receipt and understanding of this information. • Please provide this notice to those who need to be aware within your organization or to any organization or person where the product has been transferred or distributed. • Any single PB500 series ventilator should only be used for the 10-year labelled service life. If a patient requires ventilatory support for a longer period of time, cumulative sequential use of more than one PB500 series ventilator should be limited to 14 years. After 14 years of cumulative use, patients should be transitioned to an alternate device. • When approaching 14 years of cumulative use of a PB500 series ventilator, it is recommended that patients or their caregivers discuss transitioning to an alternate ventilator with the prescribing physician. | 2024-06-27 18:39
Terumo BCT, Inc. | Spectra Optia Apheresis System | 118 | - | The purpose of this letter is to remind Spectra Optia system users of a potential safety hazard when using a blood warmer attached to a Spectra Optia tubing set and to reinforce the actions required to mitigate this risk If the blood warmer is not primed before use air could be delivered to a patient x D Many Spectra Optia procedures include the option to use a blood warmer The Spectra Optia operator s manual instructs operators how to connect and prime a blood warmer when one is used The operator s manual also includes warnings about the risk of delivering air to a patient if a blood warmer is not primed These warnings appear throughout the manual in the General procedural warnings section in the Preface of the operator s manual and in the instructions for priming the lines before connecting the patient for each procedure for which a blood warmer is an option x D If an operator did not observe this warning and or did not follow the instructions in the operator s manual this safety hazard could occur x D REASON FOR THE SAFETY ALERT x D Terumo Blood and Cell Technologies has received complaints where air was returned to a patient due to an unprimed blood warmer A blood warmer that is not primed and is connected to the Spectra Optia tubing set return line could cause a risk x D A blood warmer can be attached to the tubing set either at the beginning of a procedure or during a procedure x D Kindly review the attached VOLUNTARY MEDICAL DEVICE SAFETY ALERT | ACTIONS REQUIRED BY HEALTHCARE PROVIDERS AND DISTRIBUTORS 1. Distribute this notification to all Spectra Optia system users within your organization. 2. Continue to use your Spectra Optia system(s) in accordance with the operator’s manual and the operator training materials. 3. When you configure the use of a blood warmer for the tubing set return line, the instructions to connect the blood warmer are displayed on the screen before you connect the patient. You must prime the blood warmer tubing set before you connect the patient. Please refer to Spectra Optia system operator’s manual, “Configuring the Use of a Blood Warmer.” 4. In the “Selecting Procedure Options” section of the operator’s manual, follow the updated instructions to indicate the use of a blood warmer during the procedure. You must prime the blood warmer tubing set before you connect the patient. 5. IMPORTANT: Complete the attached acknowledgement and fax or email the acknowledgement to Terumo Blood and Cell Technologies by August 31, 2024. Your return of the acknowledgement is critical so that we can confirm that you have received the Safety Alert. Kindly review the attached VOLUNTARY MEDICAL DEVICE SAFETY ALERT | 2024-06-30 08:06
Siemens Healthcare GmbH | Luminos dRF Max | 10 | - | The mounting screw on the support of the display ceiling wall suspension may loosen or be missing causing the support arm with displays to lower downwards with continued use | The corrective action will involve check and replacement of the attachment of the support arm for the display ceiling/wall suspension by a service technician in an onsite visit. In instances where the mounting screw is broken. | 2024-06-30 09:56
Siemens Healthcare GmbH | Luminos Agile Max | 10 | - | The mounting screw on the support of the display ceiling wall suspension may loosen or be missing causing the support arm with displays to lower downwards with continued use | The corrective action will involve check and replacement of the attachment of the support arm for the display ceiling/wall suspension by a service technician in an onsite visit. In instances where the mounting screw is broken. | 2024-06-30 09:56
Olympous | TFL Laser Footswitch – Wireless | 9 | - | FY OSTA Soltive Wireless Footswitch is being initiated because of complaint C where a hour delay resulted from a Soltive Wireless Footswitch TFL AFSWL unpairing from the SOLTIVE laser system console during a procedure The surgeon had to de scrub mid procedure to try to get the wireless footswitch to pair without success | The FCA Review Board decided to execute a Field Corrective Action consisting of a Customer Letter informing all Soltive Laser customers to have a wired footswitch available during procedures. Olympus will provide a wired footswitch to customers that do not already have one. In the interim, until a wired footswitch is provided to customers with only wireless footswitches, Olympus will recommend that customers ensure the wireless footswitch is working correctly prior to clinical use. | 2024-06-30 13:34
Zimmer Inc. | CPT® Hip System Femoral Stem | 27 | - | Increased risk of postoperative periprosthetic femoral fracture PFF While PFF is a generally recognized risk associated with all total hip arthroplasty the current instructions for use IFU for the CPT Hip System Femoral Stem Neck Taper do not identify post operative bone fracture as a risk | 1- The IFU is in the process of being updated to reflect the risk of PFF, with an anticipated release by the end of August 2024. 2- Phase out the CPT Hip System Femoral Stem 12/14 Neck Taper, with a target of December 2024. | 2024-06-30 18:32
Ventana Medical Systems, Inc. (Roche) | VENTANA anti-CD10 (SP67) Rabbit Monoclonal Primary Antibody | 149 | - | Risk of False Positive results with the above affected product due to High Background x D | - Please immediately discontinue the use of and discard any inventory of all affected lots mentioned above. - Customers must review all positive results generated with the affected lots, and follow their standard laboratory operating procedures to investigate any suspected false positive result generated with the affected lots. - A new VENTANA anti-CD10 (SP67) Rabbit Monoclonal Primary Antibody, Lot M10993 has been produced with additional testing measures and is now available, refer to the attachment for more information. | 2024-07-01 10:09
Ortho Clinical Diagnostics | 0.8% RESOLVE PANEL C SYSTEM | 0 | - | QuidelOrtho received complaints regarding weaker than expected reactivity from customers in the United Kingdom testing RESOLVE PANEL C Lot RC Cell against the NHSBT Weak Anti D control material on both automated and manual Ortho BioVue methods x D x D QuidelOrtho s investigation confirmed that the D antigen status of the donor used in the manufacture of Cell ID produces weaker than expected reactions however all other antigens in Cell are performing as expected Lot RC met and continues to meet all acceptance criteria the weaker reactivity noted for Cell is donor related due to a weaker expression of the D antigen compared to other D positive cells As a precaution this donor has been permanently deferred and will no longer be used in the manufacture of future lots of ORTHO RESOLVE PANEL C and all other red cell reagents x D x D QuidelOrtho plans to review the release process to look at potential improvements in the identification of donors that may not meet customer expectations with specific sample types x D x D All other antigens including the D antigen in Cells and are performing as expected Given the presence of D Antigen in the other panel members the appropriate x D identification of pre existing or unexpected D antibodies is still expected therefore customers may continue to use ORTHO RESOLVE PANEL C Lot RC x D x D During Quality Control QC testing a weaker than expected reaction when using a known Anti D positive control material may result in unnecessary troubleshooting and failed acceptance testing This situation may result in a delay in testing patient samples while the operator performs troubleshooting procedures however serious patient injury is not expected due to blood bank facilities emergency transfusion policies that ensure urgent transfusion is not on hold due to unavailable antibody screening identification results x D x D For patient Rhesus Rh antibody identification testing weaker than expected anti D reaction in Cell may cause confusion in result interpretation However pre existing or unexpected anti D antibodies would not be missed since the D antigen in Cells and are performing as expected x D | • Customers may continue to use ORTHO 0.8% RESOLVE PANEL C, Lot 8RC615, as all other antigens of Cell 2 are performing as expected and the D antigen of Cells 1, 3, 4 and 11 are performing as intended. • If your laboratory chooses not to continue to use ORTHO 0.8% RESOLVE PANEL C, Lot 8RC615, a replacement lot may be requested and will be fulfilled as inventory allows. • Complete and return the enclosed Confirmation of Receipt form | 2024-07-01 12:44
abbott | FILMARRAY TORCH System | 28 | - | When turning on BIOFIRE TORCH systems arcing inside of the power switch may result in x D carbon build up on the switch contacts The carbon build up may lead to excess heating x D inside of the power switch subsequently leading to deformation of the power switch case x D The deformation can create an opportunity for a power switch electrical short In addition to x D an electrical short deformation may result in an open circuit causing the power switch to fail | Required actions In this context, we request you to take the following actions. Please: • If a user determines that the BIOFIRE TORCH Base requires complete shutdown, each individual BIOFIRE TORCH Module must be powered down prior to turning on the BIOFIRE TORCH Base, • ‘Reset’ the BIOFIRE TORCH Base weekly using the ‘Reset’ button. • Move potentially flammable materials away from the TORCH system. • Post the Technical Note on or near the TORCH system in order to continuously be aware of the ‘Power On’ instructions and potential issues. • If your system has 6 or more modules installed and has been power-cycled frequently, or if the power entry switch has degraded to the point where the system should not be used, we recommend contacting the BIOFIRE support team (see contact information below) to send a field application specialist to install a replacement switch for you. • Please complete all fields of the accompanying Acknowledgement of Receipt Form and return it to bioMérieux so that bioMérieux may acknowledge your receipt of this notification. | 2024-07-02 13:52
Radiometer Medical ApS | ABL800 BASIC with AutoCheck | 787 | - | An internal technical study was carried out based on reports from internal and external users regarding pH sensitivity and pH drift errors for calibrations and measured QC pH values out of range x D x D The study concludes that there is a possibility of reporting biased out of specification pH results on blood samples x D x D This may occur if the calibration solution s pH value decreases during the in use period due to bacterial growth in the calibration solution bottles CAL and or CAL x D x D The described error may lead to a fetal patient with a scalp pH measurement result in the severe acidosis range being diagnosed as having moderate acidosis This may put the fetal patient at risk for not being delivered immediately by either instrumental vaginal delivery or urgent caesarean section and thus being at risk for experiencing permanent organ damage including brain damage due to hypoxia | With immediate effect, Radiometer requests that users implement either “Procedure 1” for customers who run daily quality controls or “Procedure 2” for customers who do not. The procedures are described within the Customer Letter. This must be carried out regardless of which patient groups are measured on the analyser. It will ensure that biased out-of-specification pH results that may lead to the risks above are not reported. | 2024-07-03 08:52
Boston Scientific | CapioTM SLIM Suture Capturing Device | 39 | - | Boston Scientific is initiating a removal of certain lots of Capio SLIM Suture Capturing Devices Capio SLIM due to an increase in reports of the device not catching the suture needle dart as expected during use Boston Scientific has determined through investigation that certain tooling used to manufacture Capio SLIM devices may have inadvertently affected component dimensions in specific lots impacting the functionality of the device Boston Scientific has addressed and corrected this finding in manufacturing | Specific batches removal | 2024-07-03 13:35
Medivance Inc. | Arctic Sun Temperature Management System | 124 | - | An internal investigation has determined that Alarm may not alert as expected when the device x D fails to reach the correct target water temperature while the device is operating in patient control mode x D x D Alarm is for Reduced Water Temperature Control The system has detected that the water temperature has not been controlled as accurately as expected in the last minutes This situation may be temporary due to sudden patient temperature changes interruption in water flow or blockage of air flow by an obstruction or dirty filter x D | Clinical User Actions The device can continue to be used in accordance with the following guidance: 1. The Arctic Sun™ Temperature Management System contains the following two fixed alarms when the device is operating in patient control mode: • Alarm 10 will alert if the patient’s temperature is below 31oC and the water temperature is below 31oC. • Alarm 08 will alert if the patient’s temperature is above 39.5oC and the water temperature is above 39.5oC. These alarms are already configured on your device, so you do not need to do anything to ensure that they are active. 2. BD also recommends that customers continue to use the customisable Patient Temperature Low (Alarm 11) and Patient Temperature High (Alarm 09) alarms on their devices. Both alarms are available when the device is operating in patient control mode. Customers should set these alarms to the lowest and highest temperatures they feel are acceptable for the specific patient being treated. These alarms come with default settings of 30oC and 44oC and alert independent of water temperature. 3. As instructed in the Operating Manual, always empty Arctic Gel Pads when therapy is ended, even if the stop in therapy is only temporary. Failure to do so may result in overfill of the water reservoir and this can affect temperature management. Actions taken by BD: As a result of our ongoing investigation, BD is developing a software update to correct this issue. Actions to be taken by BD: BD is developing a software solution to correct this issue and will provide a software upgrade to impacted devices, subject to the regulatory approvals, where applicable and BD will contact your facility in order to arrange upgrade. Customer Actions: • Review the information in Table 1 to determine if Arctic Sun™ 5000 Temperature Management System(s) in your possession are impacted. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 1st August 2024, clearly indicating the applicable contact person at your facility to support the software upgrade, when available. • Circulate this notice to all those who need to be aware within your organisation or to any organisation where the potentially affected product has been transferred. • If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: • Review the information in Table 1 and determine if the Arctic Sun™ 5000 Temperature Management System(s) in your possession are impacted. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 1st August 2024, clearly indicating the applicable contact person at their facility to support the software upgrade, when available. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. | 2024-07-03 16:41
Baxter Medical Systems GmbH + Co. KG | PST 500 | 31 | - | A sudden drop and or tilt of the surgical tabletop during a surgical procedure may cause unintended patient movement or fall This could lead to delay or interruption in therapy ineffective CPR and could cause critical patient harm such as major musculoskeletal or vital organ injuries Baxter has received six complaints related to this issue however no serious injury was reported | 1. Immediately cease all use of the impacted units until a correction has been completed. If there are alternative tables available, use them until the service kit is installed by the hospital technician, a Baxter technician, or an authorized representative. 2. Below are two options for correction of the impacted units: (to be adapted locally) a) Self-installation of the service kit (part number 2086594) executed by a hospital defined technician per the technical service bulletin included with the service kit. b) A Baxter representative will contact customers to schedule inspection of self-corrected units or correct the impacted units at your facility. Customers are instructed not to use the impacted units until the correction has been executed. 3. Customers are asked to acknowledge receipt of the Field Safety Notice 4. This information is to be provided to all users of PST 500 Surgical Table. If this product has been distributed to other facilities or departments, customers are asked to forward a copy of this communication to them in accordance with their customary procedures. | 2024-07-03 19:12
Welch Allyn, Inc. | BRAUN ThermoScan® PRO 6000 Ear Thermometer | 1772 | - | If cleaning instructions are not followed the device may be exposed to excess amounts of cleaning solution x D leading to fluid ingress If this occurs there is a risk of the probe tip overheating and potentially causing a burn to x D the user or the ear canal of the patient The population at greatest risk are patients who are unable to withdraw x D from the heat source or those who are unable to effectively communicate pain To date Baxter has not received x D any reports of serious injury related to this issue | Advice on actions to be taken by the distributor and the user 1. Clinicians may continue to use the PRO 6000 ear thermometers according to the current IFU (CD material number 421032). Please discard all versions of the outdated IFU (CD material number 419450). The material number is printed on the CD. 2. The Customer needs to complete the enclosed customer reply form and return it to Baxter by either faxing it or scanning and e-mailing it or sending it by post, even if you don’t have any inventory. Returning the customer reply form promptly will confirm your receipt of this notification and prevent you from receiving repeat notices. 3. If the customer distributes this product to other facilities or departments within their institution, they need to forward a copy of this communication to them. 4. If the customer purchased this product from a distributor, please note that the Baxter customer reply form is not applicable. If a reply form is provided by the distributor or wholesaler, they are asked to return it to the supplier according to their instructions. 5. If it is a dealer, wholesaler, distributor/reseller, or original equipment manufacturer (OEM) that distributed any affected product to other facilities, they need to notify their customers of this notification in accordance with their customary procedures | 2024-07-04 11:47
Biokit, S.A. | Alinity c Proteins Standard | 56 | - | It has identified that the use of Alinity c Proteins Standard affected lot results in Alinity c Proteins Control values out of acceptance range high when using any Alinity c A Antitrypsin reagent and controls lots x D | Please stop using the above affected lot and discard any remaining kits of it in your inventory. | 2024-07-04 13:28
SENTINELCH.S.p.A. | CALiaGold®pierceTube | 10 | - | a defect in the inner plastic crack that connects the two chambers of the collection tube | stop using the products. Dispose the affected lots in accordance with your national and local safety and environmental regulations. | 2024-07-04 15:23
Medtronic XOMED, Inc | NIM VITAL™ CONSOLE | 52 | - | There is a potential for a false negative response when using the NIM Vital Nerve Monitoring System x D Per IFU x D A false negative is the condition where the probe is on the nerve but you do not get an EMG tone There are potential failure modes that lead to false negatives x D False negative responses failure to identify nerve a condition where probe is on nerve but you do not get an EMG tone may result from x D Shorted EMG electrode or cabling conductive parts of applied needle electrodes or cables contacting each other x D Patient Interface fuse blown and not detected mA V Xomed Part No x D Patient Interface defective x D Inadequate stimulus current x D Inadequate current for stimulation of nerve through hardware such as stimulus dissection instruments may vary based on the physical size shape characteristics and design of the hardware and proximity to the nerve x D Simultaneous stimulation of the nerve and the surrounding tissue resulting in current shunting inadequate delivery of stimulus current to target nerve tissue x D Flatline on the EMG channel caused by shorted internal amplifier characterized by baseline activity of V peak to peak x D EMG electrodes not positioned properly in the target muscles x D | • Identify affected products within your inventory. The product is not required to be returned for this issue as Medtronic has deployed NIM Vital™ System software version 1.5.4. which is readily available to fix this issue. • Your Medtronic Representative will contact you to install the new software version 1.5.4 for correction of the impacted product in your possession. • Please complete and return the Customer Acknowledgement Form enclosed with this letter acknowledging receipt of this information even if you do not have the affected product. Note: Instructions on how to return that form to Medtronic can be found on the form itself. • Please share this communication within your organization, with other organizations where impacted devices have been transferred, and any other associated organizations that may be impacted by this action. Maintain a copy of this letter for your records. | 2024-07-04 15:41
OrthoPediatrics Canada ULC, dba Pega Medical. | The Fassier-Duval Telescopic IM System | 6 | - | An incorrect sub component was used during the assembly of the FRT cartridge Specifically the FRT HEX was used instead of the FRT HEX x D | - Identify in your inventory the rescue tray of the Fassier-Duval System - Identify in your inventory any additional spare parts of the FRT250 cartridge as shown in the attachment. - Contact OrthoPediatrics local representative to return parts. | 2024-07-07 13:13
Smiths Medical ASD, Inc. | BLUperc® Percutaneous Dilation Tracheostomy Kit | 1 | - | Smiths Medical has identified the potential for a disconnection of the pilot balloon from the tracheostomy inflation line within specific lots of the BLUSelect BLUgriggs and BLUperc products because of a manufacturing defect See the photo example of the issue below | Smiths Medical Actions: Smiths Medical is sending this notification to all BLUSelect®, BLUgriggs® and BLUperc® customers who received products from Smiths Medical listed in Attachment 1_Affected Product. Smiths Medical will provide replacement product(s) or credit to affected customers upon receipt of a completed response form to certify product destruction. Customer Required Actions: When using the device, all instructions, including warning and cautions contained in the Instructions for Use Documentation must be followed with heightened awareness. Please complete the following actions listed below 1. Check all inventory locations within your institution for the affected catalog numbers and lot numbers listed in the notification and discontinue use. Discard all affected products following your institution’s process for discarding. If discarding is not immediately possible at your facility, then the product should be quarantined until disposal. 2. Share this notification with all potential users of the device to ensure they are aware of this notification and proposed mitigations. If the devices are used at another location, please ensure this communication is delivered there. 3. Complete and return the attached Customer Response Form to EMEA-FSN@icumed.com within 10 days of receipt to acknowledge your understanding of this notification. | 2024-07-08 13:36
Philips | HeartStart Intrepid | 73 | - | When monitoring ECG using either a Lead or Lead ECG cable the HeartStart Intrepid Monitor Defibrillator may display intermittent ECG waveforms when the fourth limb lead is placed on the patient The ECG tracing is normal when only three limb leads are connected right arm left arm and left leg However when there are one or more poor ECG lead connections to the patient the ECG tracing either displays a dashed line or is intermittent between waveform and dashed line see Figures below This failure may occur any time the HeartStart Intrepid Monitor Defibrillator is being used to monitor Lead or Lead ECG | While a solution for this issue is in development, Philips is providing this Urgent Field Safety Notice to inform affected customers. Philips will notify you again to arrange a permanent resolution immediately upon release. Philips anticipates that the solution will be available in Q4 2024. | 2024-07-08 14:19
abbott | HeartMate 3 LVAS | 50 | - | the UI membrane is lifted | We have reinforced the Controller User Interface membrane inspection steps during the manufacturing process. However, as an additional measure, we are requesting the following from you on a temporary basis: • Carefully inspect the controller at the edge near the display button to confirm the UI membrane is not lifted (refer to provided visual examples). • Report Controllers with lifted UI membrane by contacting your Abbott Representative. • Reinforce requirement as listed on Page 2-11 “System Controller Warnings and Cautions” in the IFU to patients to protect controller from water. Advise them to follow instructions to avoid direct exposure to water; and communicate risk of continuous, steady exposure to water. • Reinforce requirement to patient to perform the daily self-test to verify the display works appropriately as listed on Page 2-29 “Performing a System Controller Self Test”. We will follow-up with a formal notification about the lifted UI membrane. In the meanwhile, we appreciate your partnership and sincerely apologize for any difficulties or inconvenience this may cause you and your patients | 2024-07-08 17:03
SAUDI MAIS CO FOR MEDICAL PRODUCTS | 3-Way Stopcock with/witout Extension Tube | 50636 | - | Leakage is due to some variation in product dimensions | All Products having this Lot. No: 230614 Reference No: 401-0010 shall be returned to the manufacturer. We will be replacing the remaining effected quantities with the good one | 2024-07-09 09:35
" Medtronic XOMED, Inc" | NIM EMG Endotracheal Tube | 903 | - | Medtronic has received complaint reports from customers experiencing issues with the NIM Standard Reinforced EMG Endotracheal Tube NIM CONTACT Reinforced EMG Endotracheal Tube during use It was reported that the cuff of the endotracheal EMG tube was over inflated causing it to herniate over the end of the tube leading to loss of ventilation due to airway blockage x D The EMG Endotracheal Tube are silicone tubes with the main shaft reinforced by a wire coil to prevent collapse while maintaining flexibility The cuffs are also manufactured with silicone Not following the Instructions for Use IFU and over inflating the cuff increases intra cuff pressure which can cause the silicone cuff to extend herniate or distort over the end of the tube and or the murphy eye potentially causing obstruction of the patient airway and loss of ventilation x D | • Immediately identify, segregate, and quarantine affected products within your inventory or control. Do not use the affected devices. Note: The list of impacted products is included in the product scope table in the Urgent Field Safety Notice. All lot numbers of NIM Standard and Contact EMG Endotracheal tubes are impacted. • Return affected products in your inventory to Medtronic. Please contact your Medtronic Representative if an alternative device is needed. Note: Instructions on how to return any impacted products to Medtronic can be found on the Customer Acknowledgement Form. Your local Medtronic Representative can assist you with the initiation of the return. • Complete and return the Customer Acknowledgement form included with this letter even if you do not have the affected product. Please Note: Instructions on how to return that form to Medtronic can be found on the form itself. Training on the NIM Standard and Contact EMG Endotracheal tube deployed with the February 2024 communication under Medtronic Academy is no longer required as part of this field action. | 2024-07-09 19:23
Ziehm Imaging GmbH | C-ARM , Ziehm Vision RFD | 7 | - | When carrying out the root cause analysis it was discovered that hand switches including a special type of microswitch are likely to fail when being exposed to mechanical stress e g dropped on the floor Even though Ziehm s risk managements procdures would currently not require risk mitigating activities based on the number of devices failed in the field Ziehm decided to proactively exchange potentially failing components in the field within the framework of the present Field Safety Corrective Action Exchanged parts hand switches will be recalled and discarded | replace all the affected hand switch | 2024-07-10 17:12
Medtronic, Inc. | DLP® Single Stage Venous Cannula | 170 | - | Medtronic received a complaint that prior to using the DLP Single Stage Venous Cannula with a right angle metal tip CFN it was reported that an incorrect device was put into the package cannula CFN Device was returned to Medtronic and the complaint was confirmed for the nonconformance During the complaint investigation the manufacturer identified that line clearance was not followed as the incorrect cannula found in the complaint was ran immediately prior to the complaint lot on the same equipment x D The incorporation of an incorrect component occurred again at the same manufacturer Therefore the issue regarding the DLP Single Stage Venous Cannulae CFN is documented under the same field action number the same Issue Impact Assessment IIA the same Corrective Action Preventive Action CAPA and the same Product Hold Order PHO as the issue regarding the DLP Vessel Cannulae x D | • Review your inventory for listed product. • Immediately identify and quarantine all unused, listed product in your inventory. • Return unused, listed product in your inventory to Medtronic. Your local Medtronic Representative can assist you with the initiation of the return. • Please complete and return the enclosed Customer Acknowledgment Form even if you do not have any affected product in your possession. • Please share this notification with others in your organization as appropriate. If product listed above has been forwarded to another facility, please notify the facility of this Urgent Field Safety Notice. • Please maintain a copy of this communication in your records. | 2024-07-10 18:41
Maquet Cardiopulmonary GmbH | QUADROX-i Adult | 5 | - | During internal investigation the manufacturer determined that the particle originates as result of abrasion of the gaskets used as part of a test device in production where each vacuum tight version of a Venous Hardshell Cardiotomy Reservoir is tested for over pressure vacuum and tightness It was found that the abrasion was caused by a too narrow diameter of the gasket x D As a corrective action the gasket has been replaced by a gasket with a wider inner diameter This change was implemented on June x D Consequently the scope of this field action was narrowed down to all products that contain a vacuum tight version of a Venous Hardshell Cardiotomy Reservoir manufactured before June | Product Removal | 2024-07-11 14:40
SCHILLER Medical SAS | Defigard touch 7 | 32 | - | Potential failures in the defibrillation shock delivery control circuit of the above affected device may lead to shocks not being delivered when the device is used | Please if a DEFIGARD Touch-7 were to cancel a shock for no apparent reason, follow the resuscitation protocol supported by the voice prompts and have the device inspected as soon as possible by your maintenance department. A new software version is being developed by Schiller Medical that will correct the fault. | 2024-07-11 16:11
Sociedad Española de Electromedicina y Calidad S.A. (SEDECAL) | Radiographic System PROTEUS XR/f | 11 | - | GE HealthCare has become aware that the XR system Proteus XR a does not have a de installation manual describing process steps for de installing the devices De installation instruction process steps must be followed when de installing these devices to avoid personnel injury during de install | You may continue to use the Proteus XR/a as intended for clinical use. This issue has no effect on clinical use of the equipment. If you intend to de-install your Proteus XR/a system or contract a 3rd party to de-install the system: 1. Ensure all instructions and guidance in the De-Installation Manual are followed. 2. Access the De-installation Manual from the product documentation portal: https://www.gehealthcare.com/support/manuals In the search bar, type either: “Proteus XR/a De-Installation Manual” or “5273026-8EN” 3. If the table dolly service tool (Part Number: 611-3700) that is used during the original system installation has been retained, ensure the bolts specified in Section 1.4.1 the De-installation Manual are used with the table dolly during de-installation and the instruction for use in Section 1.4.2 of the De-installation Manual is followed. 4. If you do not have this table dolly, follow Section 1.4.3 of the De-installation Manual for de-installation of the table without the dolly. Please ensure all potential service users in your facility or those contracted by your facility are made aware of this notification and the recommended actions. Please retain this document for your records. Please complete and return the attached acknowledgement form to SaudiArabiaServiceCenter@gehealthcare.com | 2024-07-14 13:19
Ortho Clinical Diagnostics Inc. | VITROS Chemistry Products ALB- TP Slides | 38 | - | QuidelOrtho investigated the issue and determined that during normal use of the product VITROS XT ALB TP Slides may shed dust and debris as they travel through the MicroSlide Subsystem This dust and debris may then settle onto VITROS Na Slides as they are processed through the MicroSlide Subsystem Results for both ALB and TP assays are unaffected by this issue however our investigation observed that dust and debris may cause positively or negatively biased Na results when VITROS Na Slides are processed on the same VITROS XT or XT System as the XT ALB TP Slides The direction and level of bias will depend on where the dust and debris settle on the Na Slide x D x D Until this issue is resolved QuidelOrtho will be pausing availability of the VITROS XT ALB TP Slides and we will shift all affected customers to the individual VITROS Chemistry Products ALB Slides and VITROS Chemistry Products TP Slides We have verified through internal testing that the individual ALB and TP slides do not show the same behavior as the XT ALB TP slides x D x D Upon receipt of this notification please perform the cleaning steps listed in Appendix VITROS XT System Cleaning Focus on your VITROS XT and or XT System s to remove any dust and debris which may have accumulated in the MicroSlide Subsystem x D x D x D Perform the cleaning steps listed in Appendix VITROS XT System Cleaning Focus again on your VITROS XT and or XT System s after your laboratory has discontinued using VITROS XT ALB TP Slides x D x D Customers may continue to use VITROS XT ALB TP Slides until the replacement VITROS ALB Slides and VITROS TP Slides are received After receipt of the replacement VITROS Slides and on completion of any necessary crossover studies validation please discontinue using render unusable and discard your remaining inventory of VITROS XT ALB TP Slides x D x D We are committed to ensuring that adequate supplies are available to meet all customer needs While we do not anticipate the need for allocation at this time we are closely monitoring demand and may implement allocation for VITROS ALB Slides and VITROS TP Slides if necessary in the future | • If your laboratory currently uses the VITROS XT ALB-TP Slides, perform a thorough cleaning of the MicroSlide Subsystem on your VITROS XT 3400 and/or VITROS XT 7600 System(s). Refer to Appendix 1: VITROS XT System Cleaning Focus for a list of specific areas which have been observed to accumulate dust and debris shed from VITROS XT ALB-TP Slides. • Prepare your laboratory’s transition from using VITROS XT ALB-TP Slides to the VITROS Chemistry Products ALB Slides and VITROS TP Slides products, including any crossover procedures which may be required. • Complete the enclosed Confirmation of Receipt form no later than Month DD, 2024. Upon receipt of your completed Confirmation of Receipt form, QuidelOrtho will replace your discarded inventory of VITROS XT ALB-TP Slides with the equivalent quantity of VITROS ALB Slides and VITROS TP Slides products. • Upon receipt of your replacement VITROS Chemistry Products ALB Slides and TP Slides products and completing the crossover procedure, discontinue using, render unusable, and discard your remaining inventory of VITROS XT ALB-TP Slides. • Save this notification with your User Documentation or post this notification by each VITROS XT 3400/XT 7600 System until the issue has been resolved. • Please forward this notification if the affected product was distributed outside of your facility. • If your laboratory has experienced the issue with this product and you have not already done so, please report the occurrence to your local Global Services Organization. | 2024-07-15 09:33
Philips Medical Systems Nederland B.V. | AZURION 3 M12 | 46 | - | Philips has identified two issues in software version R x of the Philips Azurion system that may affect x D x D a Management of the system Log Trace Files x D The mechanism that is present in the system to manage the number and size of the system Log Trace Files does not function properly as a result of which the Log Trace Files related to remote connection logging may occupy the full hard disk capacity of the Philips Azurion R x system When the full disk capacity is reached imaging X ray functionality will cease to be available without an advance warning to the user A restart of the system will not resolve the issue x D x D Note This issue only impacts the Philips Azurion R x systems that are currently connected to Philips Remote Service or that were connected in the past x D x D b Time to Perform a Cold Restart x D A cold restart of the Philips Azurion R x system may take up to minutes as stated in the Instructions for Use of the system Section This time includes the time required for the system to shut down and the system to start up However there are instances where the shutdown of the system may be prolonged with up to minutes resulting in the cold restart taking longer up to minutes in total | Philips is working on a software release (R2.2.10) that will correct these two issues (ref. FCOs: FCO72200548, FCO72200580, FCO72200582, FCO72200583, FCO72200584, FCO72200592). For those systems that have an Interventional Workspot (IW) and/or EchoNavigator, Philips will upgrade the software version of the IW and/or EchoNavigator to maintain compatibility with the updated Philips Azurion system software (R2.2.10). Your local Philips representative will contact you to schedule a visit to install the software update once available. Philips expects this software to be released by Q1 2025. | 2024-07-15 13:50
Philips Medical Systems Nederland B.V. | AZURION 3 M15 | 46 | - | Philips has identified two issues in software version R x of the Philips Azurion system that may affect x D x D a Management of the system Log Trace Files x D The mechanism that is present in the system to manage the number and size of the system Log Trace Files does not function properly as a result of which the Log Trace Files related to remote connection logging may occupy the full hard disk capacity of the Philips Azurion R x system When the full disk capacity is reached imaging X ray functionality will cease to be available without an advance warning to the user A restart of the system will not resolve the issue x D x D Note This issue only impacts the Philips Azurion R x systems that are currently connected to Philips Remote Service or that were connected in the past x D x D b Time to Perform a Cold Restart x D A cold restart of the Philips Azurion R x system may take up to minutes as stated in the Instructions for Use of the system Section This time includes the time required for the system to shut down and the system to start up However there are instances where the shutdown of the system may be prolonged with up to minutes resulting in the cold restart taking longer up to minutes in total | Philips is working on a software release (R2.2.10) that will correct these two issues (ref. FCOs: FCO72200548, FCO72200580, FCO72200582, FCO72200583, FCO72200584, FCO72200592). For those systems that have an Interventional Workspot (IW) and/or EchoNavigator, Philips will upgrade the software version of the IW and/or EchoNavigator to maintain compatibility with the updated Philips Azurion system software (R2.2.10). Your local Philips representative will contact you to schedule a visit to install the software update once available. Philips expects this software to be released by Q1 2025. | 2024-07-15 13:50
Philips Medical Systems Nederland B.V. | AZURION 7 B12 | 46 | - | Philips has identified two issues in software version R x of the Philips Azurion system that may affect x D x D a Management of the system Log Trace Files x D The mechanism that is present in the system to manage the number and size of the system Log Trace Files does not function properly as a result of which the Log Trace Files related to remote connection logging may occupy the full hard disk capacity of the Philips Azurion R x system When the full disk capacity is reached imaging X ray functionality will cease to be available without an advance warning to the user A restart of the system will not resolve the issue x D x D Note This issue only impacts the Philips Azurion R x systems that are currently connected to Philips Remote Service or that were connected in the past x D x D b Time to Perform a Cold Restart x D A cold restart of the Philips Azurion R x system may take up to minutes as stated in the Instructions for Use of the system Section This time includes the time required for the system to shut down and the system to start up However there are instances where the shutdown of the system may be prolonged with up to minutes resulting in the cold restart taking longer up to minutes in total | Philips is working on a software release (R2.2.10) that will correct these two issues (ref. FCOs: FCO72200548, FCO72200580, FCO72200582, FCO72200583, FCO72200584, FCO72200592). For those systems that have an Interventional Workspot (IW) and/or EchoNavigator, Philips will upgrade the software version of the IW and/or EchoNavigator to maintain compatibility with the updated Philips Azurion system software (R2.2.10). Your local Philips representative will contact you to schedule a visit to install the software update once available. Philips expects this software to be released by Q1 2025. | 2024-07-15 13:50
Philips Medical Systems Nederland B.V. | AZURION 7 B20 | 46 | - | Philips has identified two issues in software version R x of the Philips Azurion system that may affect x D x D a Management of the system Log Trace Files x D The mechanism that is present in the system to manage the number and size of the system Log Trace Files does not function properly as a result of which the Log Trace Files related to remote connection logging may occupy the full hard disk capacity of the Philips Azurion R x system When the full disk capacity is reached imaging X ray functionality will cease to be available without an advance warning to the user A restart of the system will not resolve the issue x D x D Note This issue only impacts the Philips Azurion R x systems that are currently connected to Philips Remote Service or that were connected in the past x D x D b Time to Perform a Cold Restart x D A cold restart of the Philips Azurion R x system may take up to minutes as stated in the Instructions for Use of the system Section This time includes the time required for the system to shut down and the system to start up However there are instances where the shutdown of the system may be prolonged with up to minutes resulting in the cold restart taking longer up to minutes in total | Philips is working on a software release (R2.2.10) that will correct these two issues (ref. FCOs: FCO72200548, FCO72200580, FCO72200582, FCO72200583, FCO72200584, FCO72200592). For those systems that have an Interventional Workspot (IW) and/or EchoNavigator, Philips will upgrade the software version of the IW and/or EchoNavigator to maintain compatibility with the updated Philips Azurion system software (R2.2.10). Your local Philips representative will contact you to schedule a visit to install the software update once available. Philips expects this software to be released by Q1 2025. | 2024-07-15 13:50
Philips Medical Systems Nederland B.V. | AZURION 7 M12 | 46 | - | Philips has identified two issues in software version R x of the Philips Azurion system that may affect x D x D a Management of the system Log Trace Files x D The mechanism that is present in the system to manage the number and size of the system Log Trace Files does not function properly as a result of which the Log Trace Files related to remote connection logging may occupy the full hard disk capacity of the Philips Azurion R x system When the full disk capacity is reached imaging X ray functionality will cease to be available without an advance warning to the user A restart of the system will not resolve the issue x D x D Note This issue only impacts the Philips Azurion R x systems that are currently connected to Philips Remote Service or that were connected in the past x D x D b Time to Perform a Cold Restart x D A cold restart of the Philips Azurion R x system may take up to minutes as stated in the Instructions for Use of the system Section This time includes the time required for the system to shut down and the system to start up However there are instances where the shutdown of the system may be prolonged with up to minutes resulting in the cold restart taking longer up to minutes in total | Philips is working on a software release (R2.2.10) that will correct these two issues (ref. FCOs: FCO72200548, FCO72200580, FCO72200582, FCO72200583, FCO72200584, FCO72200592). For those systems that have an Interventional Workspot (IW) and/or EchoNavigator, Philips will upgrade the software version of the IW and/or EchoNavigator to maintain compatibility with the updated Philips Azurion system software (R2.2.10). Your local Philips representative will contact you to schedule a visit to install the software update once available. Philips expects this software to be released by Q1 2025. | 2024-07-15 13:50
Philips Medical Systems Nederland B.V. | AZURION 7 M20 | 46 | - | Philips has identified two issues in software version R x of the Philips Azurion system that may affect x D x D a Management of the system Log Trace Files x D The mechanism that is present in the system to manage the number and size of the system Log Trace Files does not function properly as a result of which the Log Trace Files related to remote connection logging may occupy the full hard disk capacity of the Philips Azurion R x system When the full disk capacity is reached imaging X ray functionality will cease to be available without an advance warning to the user A restart of the system will not resolve the issue x D x D Note This issue only impacts the Philips Azurion R x systems that are currently connected to Philips Remote Service or that were connected in the past x D x D b Time to Perform a Cold Restart x D A cold restart of the Philips Azurion R x system may take up to minutes as stated in the Instructions for Use of the system Section This time includes the time required for the system to shut down and the system to start up However there are instances where the shutdown of the system may be prolonged with up to minutes resulting in the cold restart taking longer up to minutes in total | Philips is working on a software release (R2.2.10) that will correct these two issues (ref. FCOs: FCO72200548, FCO72200580, FCO72200582, FCO72200583, FCO72200584, FCO72200592). For those systems that have an Interventional Workspot (IW) and/or EchoNavigator, Philips will upgrade the software version of the IW and/or EchoNavigator to maintain compatibility with the updated Philips Azurion system software (R2.2.10). Your local Philips representative will contact you to schedule a visit to install the software update once available. Philips expects this software to be released by Q1 2025. | 2024-07-15 13:50
Medtronic, Inc. | Carelink Encore™ | 7239 | - | There is a rare potential for reduced or no energy output during high voltage HV therapy typically J in implantable cardioverter defibrillators ICDs and cardiac resynchronization therapy defibrillators CRT Ds manufactured with a specific glassed feedthrough including currently available ICDs and CRT Ds x D x D There is an increased potential for a reduced or no energy HV therapy in the AX B configuration when all the following conditions are met x D The device has a glassed feedthrough manufactured after July x D There is significant separation of the layers of insulation materials in the feedthrough components of the device header x D An unintended current pathway forms within the void created by the insulation separation capable of conducting high levels of current during HV therapy x D x D When an unintended current pathway is detected during HV therapy the Short Circuit Protection SCP feature may trigger This behavior can be intermittent both full energy and reduced energy HV therapies within the same episode have been observed SCP events may also be lead related for both lead related and device related unintended current pathways the defibrillation waveform is truncated early in the energy delivery sequence resulting in reduced or no energy being delivered J x D | Ensure that all Medtronic programmers are updated. Medtronic representatives can assist as necessary. | 2024-07-15 23:07
Medtronic, Inc. | Carelink® Programmer | 7239 | - | There is a rare potential for reduced or no energy output during high voltage HV therapy typically J in implantable cardioverter defibrillators ICDs and cardiac resynchronization therapy defibrillators CRT Ds manufactured with a specific glassed feedthrough including currently available ICDs and CRT Ds x D x D There is an increased potential for a reduced or no energy HV therapy in the AX B configuration when all the following conditions are met x D The device has a glassed feedthrough manufactured after July x D There is significant separation of the layers of insulation materials in the feedthrough components of the device header x D An unintended current pathway forms within the void created by the insulation separation capable of conducting high levels of current during HV therapy x D x D When an unintended current pathway is detected during HV therapy the Short Circuit Protection SCP feature may trigger This behavior can be intermittent both full energy and reduced energy HV therapies within the same episode have been observed SCP events may also be lead related for both lead related and device related unintended current pathways the defibrillation waveform is truncated early in the energy delivery sequence resulting in reduced or no energy being delivered J x D | Ensure that all Medtronic programmers are updated. Medtronic representatives can assist as necessary. | 2024-07-15 23:07
Medtronic, Inc. | Carelink® Programmer | 106 | - | The programmer was designed with the capability for cursor control using stylus or finger touch However cursor control in the original released design was limited to stylus only via software controls There is a limited availability of styli and as response to this Desktop software to enable finger touch interaction was developed During Desktop development issues with cursor control associated with enabling finger touch were noted Subsequent investigation into the behavior determined that grounding variability in the display assembly hardware could lead to autonomous cursor movement on some Desktop enabled programmers x D The cause of this behavior has been determined to be grounding variability in the display assembly hardware Because this capability had not been previously implemented in manufacturing nor had it ever been used high impedance grounding may have been present at the time of manufacturing or it may have increased over time through oxidation Based on the age of this population of programmers years old if a programmer is susceptible to grounding variability it would be present at the time of the update and mitigated through proper evaluation of the capability If this behavior is observed on a programmer the Medtronic Service Center performs a repair to improve the grounding variability in the display assembly resulting in a programmer with no observability of this behavior when evaluated before being returned to the field Through engineering evaluation programmers that have no observations of this behavior do not develop a significant change in performance over time x D | Software updates are necessary to maintain proper programmer function. Medtronic representatives will assist in performing the software update on all Medtronic CareLink™ 2090 programmers and assess proper function. Medtronic representatives will assist with returning programmers needing repair or replacement. • Following review of this letter, sign and return the enclosed Customer Acknowledgement Form. • Please forward this notice to others who need to be aware within your organization. | 2024-07-15 23:15
HOLOGIC N.V | QDR Series X-Ray Bone Densitometer | 6 | - | It has been identified a non conformance in Horizon DXA devices was identified The non conformance pertains to electromagnetic compatibility requirements according to the international technical standard IEC for the safety and essential performance of medical electrical equipment where the result from the Horizon DXA System exceeded the electromagnetic compatibility limit x D | Recommendations for customers: - Do NOT scan patients that have active implanted medical devices, including but not limited to neurostimulators, pacemakers, cardiac defibrillators, continuous glucose monitors, or other bio-wearable sensors. - Any operator who has an active implanted medical device should also refrain from operating the system at this time. - Do NOT scan patients that are currently being treated with an electronic medical device. Until the rectification is completed, the attachment and the specific warnings within it supersedes information provided on the Horizon DXA labelling and IFU pertaining to electromagnetic compatibility and electromagnetic interference, refer to the attachment for more information. | 2024-07-16 10:51
Biomerieux | BioFire® FilmArray® Gastrointestinal (GI) Panel | - | - | The cause for this risk is the presence of an increased level of non viable organism from Vibrio Vibrio Cholerae targets in Remel Cary Blair transport medium The presence of non viable organisms does not compromise the intended function of the transport medium However the BIOFIRE GI Panel detects nucleic acid from viable and non viable organisms alike | Required actions In this context, we request you take the following actions: • If the BIOFIRE GI Panel is used in conjunction with certain lots of Remel™ Cary-Blair (Lot # 743930, 769240, 712258, 782425, 732357, and 742049), positive results for Vibrio/Vibrio Cholerae should be confirmed by another method prior to reporting the test results. • Distribute this information to all appropriate personnel in your laboratory, retain a copy for your files, post this letter in or near the laboratory, and forward this information to all parties that may use this product, including others to whom you may have transferred our product. • Complete the Acknowledgement Form in Attachment A and return it to bioMérieux so that bioMérieux may acknowledge your receipt of this notification. It is important that you return the acknowledgment form to bioMérieux even if you determine this urgent field safety notice does not impact your facility. | 2024-07-17 18:00
stryker | ProCuity® beds missing electrical test data. | 11 | - | Stryker has identified that the manufacturing records for certain ProCuity beds x D are missing test values for electrical safety tests as required by IEC x D As a result of the missing test data the impacted units are not compliant to the x D above standard Therefore Stryker is initiating this voluntary recall to perform x D the necessary testing and documentation in the field | Actions needed 1. Locate the units listed on the attached business reply form and identify the address where they can be serviced. 2. Return the enclosed business reply form to confirm receipt of this notification by email miruna.proca@stryker.com 3. Upon receipt of the completed business reply form, Stryker will contact you to arrange a time to complete the required electrical tests: a. Dielectric voltage-withstand (HIPOT) test in accordance with IEC 60601- 1. b. Ground Continuity test as required annually per the PM checklist in the device Operations/Maintenance Manual. c. If you have questions about these tests, please contact your Stryker Representative below. 4. If you have loaned or sold any of the products listed in this letter, please forward a copy of this notice to the new users and advise us of their new location in the space provided on the business reply form. 5. If you have disposed of any of these units and they are no longer in use, please, advise us of their obsolescence by providing us with their serial number in the space provided on the business reply form. | 2024-07-17 18:07
Werfen | ACL TOP 750-CTS-LAS /350/550 CTS | 61 | - | The following software versions will contain the final mitigations to address a potential sample misidentification x D Software Version or earlier and Windows Operating System OS must be upgraded to Software x D Software Version or Parameters P and Windows OS must be upgraded to Software | Please take the following actions: • Share this information with your laboratory staff and update your internal procedures, as needed. • Forward this notification to all affected locations in your facility. • Post this notification on all affected systems in your facility | 2024-07-18 14:33
Becton Dickinson and Company | BACTEC™ MGIT™ 960 PZA Kit | 434 | - | Through internal BD complaint trending and subsequent internal raw material testing it was identified x D that the BD BACTEC MGIT PZA Kits listed in Appendix may intermittently produce falsely x D resistant results for pyrazinamide PZA during susceptibility testing of Mycobacterium tuberculosis x D isolates x D | Clinical User Actions • Customers should discard the affected product immediately. • There are no recommendations for retesting or reviewing previous patient test results. BD Actions: • BD will issue credit to the customers. • BD has identified the root cause and is taking action to prevent recurrence of this product issue. Customer Actions: • Cease use of any unused affected lots of BD BACTEC™ MGIT™ 960 PZA Kit. • Identify and quarantine all unused affected lots of BD BACTEC™ MGIT™ 960 PZA Kit. • Make a note of the lot numbers and destroy all unused affected units. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 30th August 2024. • Circulate this notice to all those who need to be aware within your organization or to any organization where the potentially affected products have been transferred. • If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: • Cease distribution. • Identify, quarantine, making a note of the lot numbers then destroy all unused affected lots of BD BACTEC™ MGIT™ 960 PZA Kit. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 30th August 2024. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. | 2024-07-18 16:19
Randox Laboratories Ltd | Immunoassay Speciality I – Level 3 | - | - | Randox can confirm a transcription error has occurred on the additional value sheet for x D Beckman Saudi for Immunoassay Speciality I Level lot EC The expiry has been x D incorrectly listed as the correct expiry is The additional value sheet x D has been updated and is available on www randox com please discard the incorrect version x D and download the corrected version | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2024-07-21 09:24
Randox Laboratories Ltd | RX Daytona Plus (with ISE) | - | - | Personnel process not followed x D Human error parameter sign off procedure not completely followed Parameter technical range checked x D against incorrect IFU | Download and update parameters, or manually enter updates to avoid losing any previously obtained data, for assays performed on your Rx Daytona Plus. Screenshots are available upon request for reference, please contact technical.services@randox.com for more information. Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2024-07-21 11:26
Randox Laboratories Ltd | RX Daytona Plus (without ISE) | - | - | Personnel process not followed x D Human error parameter sign off procedure not completely followed Parameter technical range checked x D against incorrect IFU | Download and update parameters, or manually enter updates to avoid losing any previously obtained data, for assays performed on your Rx Daytona Plus. Screenshots are available upon request for reference, please contact technical.services@randox.com for more information. Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2024-07-21 11:26
Beckman Coulter Inc. | Access HBc Ab | 167 | - | Beckman Coulter has determined that reagent packs from Access HBc Ab reagent kit lot number listed in the table above may contain an unusual amount of PMP clumping at the top of Well |  Before inversion, visually inspect all reagent packs from Access HBc Ab reagent kit lot number listed in the table above.  For any Access HBc Ab reagent pack where PMP clumping is observed at the top of Well 0: o Discontinue using the reagent pack and discard following your local laboratory protocols. o Contact your local Beckman Coulter representative to order replacement product for any affected reagent packs.  For any Access HBc Ab reagent pack where PMP clumping is not observed at the top of Well 0: o Continue using the reagent pack. o As per the Access HBc Ab reagent Instructions for Use (IFU), mix contents of new (unpunctured) reagent packs by gently inverting pack several times before loading on the instrument.  Ensure all Access HBc Ab reagent packs are stored in an upright position.  Beckman Coulter recommends sharing the content of this letter with your laboratory and/or Medical Director regarding the need to review previous patient test results. | 2024-07-21 14:56
Beckman Coulter Inc. | Access HBc Ab | 167 | - | Beckman Coulter has determined that reagent packs from Access HBc Ab reagent kit lot number listed in the table above may contain an unusual amount of PMP clumping at the top of Well |  Before inversion, visually inspect all reagent packs from Access HBc Ab reagent kit lot number listed in the table above.  For any Access HBc Ab reagent pack where PMP clumping is observed at the top of Well 0: o Discontinue using the reagent pack and discard following your local laboratory protocols. o Contact your local Beckman Coulter representative to order replacement product for any affected reagent packs.  For any Access HBc Ab reagent pack where PMP clumping is not observed at the top of Well 0: o Continue using the reagent pack. o As per the Access HBc Ab reagent Instructions for Use (IFU), mix contents of new (unpunctured) reagent packs by gently inverting pack several times before loading on the instrument.  Ensure all Access HBc Ab reagent packs are stored in an upright position.  Beckman Coulter recommends sharing the content of this letter with your laboratory and/or Medical Director regarding the need to review previous patient test results. | 2024-07-21 15:31
Biomerieux | VITEK® 2 Software 9.01-9.04 | 232 | - | Description of the issue x D bioM rieux has become aware that when using the Firefox web browser to access the VITEK x D Software versions and editing the Accession ID on an isolate or a batch of isolates an x D anomaly can occur The described combination above will lead to mismatched results if the arrow x D navigation keys in Isolate Detail View are used to move to the next isolate in the list after a change x D to the Accession ID field Currently in the Firefox web browser ONLY the use of the navigation arrow x D keys located immediately above the Accession ID field name pushes the x D edited field e g Accession ID to the next Isolate in the work list rather than the intended isolate | Required actions In this context, we request you to take the following actions. Please: • Distribute this information to all appropriate personnel in your laboratory, retain a copy in your files, and forward this information to all parties that may use this product, including others to whom you may have transferred our product. • Complete the Acknowledgement Form in Attachment A and return it to your local bioMérieux representative to confirm receipt of this notice. It is important that you return the acknowledgement form to bioMérieux even if you determine that this urgent product correction notice does not impact your facility. | 2024-07-21 16:48
Beckman Coulter Inc. | PE Dynamic Inlet-PLC | 32 | - | Customers have observed an issue where the metal part of the sample holder detaches from the plastic holder base when the robot arm grips a sample tube If this occurs at a centrifuge the metal part can inadvertently enter the centrifuge along with the sample tube The centrifuge is unable to detect the presence of the detached metal part and begins to rotate This can pose risks to both the equipment and the samples being processed | If this occurs, follow your laboratory procedure to handle biohazard materials and contact your Beckman Coulter representative for servicing your Power Express system. During the six months preventative maintenance all sample holders on a system are checked by Beckman Coulter for functionality and integrity. | 2024-07-21 23:10
Ortho Clinical Diagnostics | Folate Reagent Pack 1/2 | 480 | - | QuidelOrtho investigated customer complaints regarding generation of condition code TM x D MA or TM MB when testing the VITROS Folate on VITROS Immunodiagnostic x D Systems and VITROS XT Integrated Systems and confirmed that fibrinogen in x D patient plasma samples precipitates out of solution upon addition of folate stabilizer reagent x D as part of the pre treatment process which takes place in the MicroWell incubator middle x D ring Plasma from patients with elevated fibrinogen levels may be predisposed to causing x D TM MA or TM MB condition codes | If retesting the sample leads to subsequent “TM5-4MA” or “TM5-4MB” condition codes, it may not be possible to generate a result with this sample on the VITROS 3600 Immunodiagnostic Systems or VITROS 5600/XT 7600 Integrated Systems. • Use serum samples instead of plasma if repeat “TM5-4MA” or “TM5-4MB” condition codes are noted with a particular patient plasma sample. • Complete the enclosed Confirmation of Receipt form no later than July 23, 2024. • Please forward this notification if the affected product was distributed outside of your facility. | 2024-07-22 08:17
B. Braun Surgical, S.A. | NOVOSYN CHD | 216 | - | A manufacturing issue and some units of the mentioned references batches could have the package damaged consequently the product sterility could be compromised in addition to a lack of tightness of the package This lack of tightness of the package could accelerate the degradation of the suture thread not fulfilling the product specifications | Please identify and quarantine if you still have the listed product in your warehouse and send the product back to GMC immediately. | 2024-07-22 18:17
B. Braun Surgical, S.A. | NOVOSYN CHD | 216 | - | A manufacturing issue and some units of the mentioned references batches could have the package damaged consequently the product sterility could be compromised in addition to a lack of tightness of the package This lack of tightness of the package could accelerate the degradation of the suture thread not fulfilling the product specifications | Please identify and quarantine if you still have the listed product in your warehouse and send the product back to GMC immediately. | 2024-07-22 18:36
Randox Laboratories Ltd | All Randox RX Analysers | - | - | In line with IVDR Randox Laboratories have added the following statement to all instrument x D manuals advising all laboratories operating any Randox analysers of the procedure to follow x D in the event of an adverse incident x D | Action to be taken: Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Print a copy of the Technical Bulletin included with this notice and place in your analyser manuals. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. If you have any questions, please contact Randox Technical Services. | 2024-07-23 15:35
Randox Laboratories Ltd | All Randox Evidence Analysers | - | - | In line with IVDR Randox Laboratories have added the following statement to all instrument x D manuals advising all laboratories operating any Randox analysers of the procedure to follow x D in the event of an adverse incident x D | Action to be taken: Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Print a copy of the Technical Bulletin included with this notice and place in your analyser manuals. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. If you have any questions, please contact Randox Technical Services. | 2024-07-23 15:35
SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD. | V90 electronic vaporizer (for used with A9 Anesthesia system) | 40 | - | Mindray has identified a potential safety issue with batches of V electronic vaporizers The V electronic vaporizer is used to deliver inhalation anesthetic agents in the A anesthesia system During use the anesthetic injector of V electronic vaporizers may not be sealed tightly which may potentially result in the output of the anesthetic concentration is higher than the setting value This can be monitored by the anesthesia system in realtime and will give a relevant alarm such as Vaporizer Output Abnormal | Replace the V90 electronic Vaporizer | 2024-07-25 10:17
SCHILLER AG | CARDIOVIT AT-180 | - | - | Please find attached FSN letter | Please find attached FSN letter | 2024-07-25 11:54
Becton Dickinson and Company | ME0000002812SFDAA00205 | 125 | - | BD identified through post market surveillance an increase in the number of complaints for Escherichia coli E coli misidentification for clinical samples that were tested on BD Phoenix M instrument and BD Phoenix x D x D This issue impacts customers who are using the system for identification x D | Clinical User Actions 1. Continue use of the instrument. However, confirmatory testing should be performed on any suspected E. coli misidentifications until BD is able to perform the remediation to correct the issue. 2. It is not necessary to review previous test results and no additional clinical actions are recommended. BD Actions: 1. BD has identified the root cause and will implement appropriate corrective actions to prevent recurrence of this issue. 2. For BD Phoenix™ M50 users: BD will contact your facility to schedule a service visit to provide an update to the PUD (Phoenix Update Data) and software to correct this issue. 3. For BD Phoenix™ 100 users: A BD representative will contact your facility regarding next steps. Customer Actions: • Immediately inspect your facility to identify if you have the affected product. o Refer to Appendix 1 to confirm software version • Complete and return the Customer Response Form even if you no longer have any of the impacted product remaining in your facility by 30th August 2024. • Circulate this notice to all those who need to be aware within your organisation or to any organisation where the potentially affected product has been transferred. • If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: • Immediately inspect your facility to identify if you have the affected product in your inventory. o Refer to Appendix 1 to confirm software version • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 30th August 2024. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. | 2024-07-25 12:04
Philips Medical Systems Nederland B.V. | AZURION 3 M12 | 7 | - | A defective O Ring used in the level switch of the Certeray Generator can cause oil leakage from the tank inside the Cooling Unit leading to a low oil level The oil may accumulate in the drip tray located below the Cooling Unit in the technical room making it unlikely that the user will be aware of the leakage x D x D If this issue occurs during clinical use a low oil level will not affect the system functionality or patient treatment However during the next Azurion self test upon powering on of the system a low oil level will be detected triggering an emergency state where only low load fluoroscopy mode is possible In this mode the peak power for generation of fluoroscopy is lower than in normal operational mode resulting in a lower image quality x D | Philips will replace all Cooling Units in the systems affected by this issue (reference FCO72200579). Your local Philips representative will contact you to schedule a visit to replace the affected Cooling Unit. | 2024-07-25 13:35
Philips Medical Systems Nederland B.V. | AZURION 3 M15 | 7 | - | A defective O Ring used in the level switch of the Certeray Generator can cause oil leakage from the tank inside the Cooling Unit leading to a low oil level The oil may accumulate in the drip tray located below the Cooling Unit in the technical room making it unlikely that the user will be aware of the leakage x D x D If this issue occurs during clinical use a low oil level will not affect the system functionality or patient treatment However during the next Azurion self test upon powering on of the system a low oil level will be detected triggering an emergency state where only low load fluoroscopy mode is possible In this mode the peak power for generation of fluoroscopy is lower than in normal operational mode resulting in a lower image quality x D | Philips will replace all Cooling Units in the systems affected by this issue (reference FCO72200579). Your local Philips representative will contact you to schedule a visit to replace the affected Cooling Unit. | 2024-07-25 13:35
Philips Medical Systems Nederland B.V. | Azurion 5 M12 | 7 | - | A defective O Ring used in the level switch of the Certeray Generator can cause oil leakage from the tank inside the Cooling Unit leading to a low oil level The oil may accumulate in the drip tray located below the Cooling Unit in the technical room making it unlikely that the user will be aware of the leakage x D x D If this issue occurs during clinical use a low oil level will not affect the system functionality or patient treatment However during the next Azurion self test upon powering on of the system a low oil level will be detected triggering an emergency state where only low load fluoroscopy mode is possible In this mode the peak power for generation of fluoroscopy is lower than in normal operational mode resulting in a lower image quality x D | Philips will replace all Cooling Units in the systems affected by this issue (reference FCO72200579). Your local Philips representative will contact you to schedule a visit to replace the affected Cooling Unit. | 2024-07-25 13:35
Philips Medical Systems Nederland B.V. | Azurion 5 M20 | 7 | - | A defective O Ring used in the level switch of the Certeray Generator can cause oil leakage from the tank inside the Cooling Unit leading to a low oil level The oil may accumulate in the drip tray located below the Cooling Unit in the technical room making it unlikely that the user will be aware of the leakage x D x D If this issue occurs during clinical use a low oil level will not affect the system functionality or patient treatment However during the next Azurion self test upon powering on of the system a low oil level will be detected triggering an emergency state where only low load fluoroscopy mode is possible In this mode the peak power for generation of fluoroscopy is lower than in normal operational mode resulting in a lower image quality x D | Philips will replace all Cooling Units in the systems affected by this issue (reference FCO72200579). Your local Philips representative will contact you to schedule a visit to replace the affected Cooling Unit. | 2024-07-25 13:35
Philips Medical Systems Nederland B.V. | AZURION 7 B12 | 7 | - | A defective O Ring used in the level switch of the Certeray Generator can cause oil leakage from the tank inside the Cooling Unit leading to a low oil level The oil may accumulate in the drip tray located below the Cooling Unit in the technical room making it unlikely that the user will be aware of the leakage x D x D If this issue occurs during clinical use a low oil level will not affect the system functionality or patient treatment However during the next Azurion self test upon powering on of the system a low oil level will be detected triggering an emergency state where only low load fluoroscopy mode is possible In this mode the peak power for generation of fluoroscopy is lower than in normal operational mode resulting in a lower image quality x D | Philips will replace all Cooling Units in the systems affected by this issue (reference FCO72200579). Your local Philips representative will contact you to schedule a visit to replace the affected Cooling Unit. | 2024-07-25 13:35
Philips Medical Systems Nederland B.V. | AZURION 7 B20 | 7 | - | A defective O Ring used in the level switch of the Certeray Generator can cause oil leakage from the tank inside the Cooling Unit leading to a low oil level The oil may accumulate in the drip tray located below the Cooling Unit in the technical room making it unlikely that the user will be aware of the leakage x D x D If this issue occurs during clinical use a low oil level will not affect the system functionality or patient treatment However during the next Azurion self test upon powering on of the system a low oil level will be detected triggering an emergency state where only low load fluoroscopy mode is possible In this mode the peak power for generation of fluoroscopy is lower than in normal operational mode resulting in a lower image quality x D | Philips will replace all Cooling Units in the systems affected by this issue (reference FCO72200579). Your local Philips representative will contact you to schedule a visit to replace the affected Cooling Unit. | 2024-07-25 13:35
Philips Medical Systems Nederland B.V. | AZURION 7 M12 | 7 | - | A defective O Ring used in the level switch of the Certeray Generator can cause oil leakage from the tank inside the Cooling Unit leading to a low oil level The oil may accumulate in the drip tray located below the Cooling Unit in the technical room making it unlikely that the user will be aware of the leakage x D x D If this issue occurs during clinical use a low oil level will not affect the system functionality or patient treatment However during the next Azurion self test upon powering on of the system a low oil level will be detected triggering an emergency state where only low load fluoroscopy mode is possible In this mode the peak power for generation of fluoroscopy is lower than in normal operational mode resulting in a lower image quality x D | Philips will replace all Cooling Units in the systems affected by this issue (reference FCO72200579). Your local Philips representative will contact you to schedule a visit to replace the affected Cooling Unit. | 2024-07-25 13:35
Philips Medical Systems Nederland B.V. | AZURION 7 M20 | 7 | - | A defective O Ring used in the level switch of the Certeray Generator can cause oil leakage from the tank inside the Cooling Unit leading to a low oil level The oil may accumulate in the drip tray located below the Cooling Unit in the technical room making it unlikely that the user will be aware of the leakage x D x D If this issue occurs during clinical use a low oil level will not affect the system functionality or patient treatment However during the next Azurion self test upon powering on of the system a low oil level will be detected triggering an emergency state where only low load fluoroscopy mode is possible In this mode the peak power for generation of fluoroscopy is lower than in normal operational mode resulting in a lower image quality x D | Philips will replace all Cooling Units in the systems affected by this issue (reference FCO72200579). Your local Philips representative will contact you to schedule a visit to replace the affected Cooling Unit. | 2024-07-25 13:35
Zevex, Inc., 4314 Zevex Park Lane, Salt Lake City, UT, USA 84123 (doing business as Moog, Medical Devices Division) | FreeGo | 4522 | - | Moog as manufacturer of the FreeGo Enteral Feeding Pump is issuing a Field Safety Notice to provide information on a software issue which may cause a potential malfunction of the FreeGo Pump This Field Safety Notice applies to all FreeGo Enteral Feeding Pumps x D During internal testing it has been identified that the FreeGo Pump has the unlikely potential to malfunction due to a software issue This may lead to the pump not alarming for an occlusion or blockage even if an occlusion exists which could lead to no feed being delivered x D For patients at risk for hypoglycemia or hypovolemia please contact your healthcare provider for further advice as there could be a risk of harm to the patient s health due to the pump not delivering a feed as expected x D Pumps do not need to be returned at this time | Corrective Action Being Taken by the Manufacturer: Moog is developing a software update to resolve this issue for the FreeGo pumps. We anticipate this software update will be available in the first quarter of 2025 and an implementation of the update will follow. Users will be advised when the software update is ready for installation by an update to this Field Safety Notice. In the interim, an additional step should be taken by the user before the start of each feeding session as detailed below. The FreeGo pump operator manual will also be updated with more prescriptive instructions for use on priming technique. | 2024-07-26 15:50
Covidien llc | McGRATH MAC | 504 | - | Medtronic received customer reports related to the McGRATH V battery SKU installed in a video laryngoscope handle SKU First Generation reporting that the McGRATH V battery exploded Batteries that are depleted below the design threshold may increase the risk of the battery becoming unstable which may lead to a thermal event followed by a risk of explosion x D Medtronic s investigations to date have identified the root cause to these events as an internal short within the battery assembly that leads to a thermal run away which in turn leads to an explosive event There are three identified contributing factors that can work in combination to create an internal short within the battery x D A Mishandling Drop impact or excessive vibration of the battery x D B Battery Management Deep depletion of the battery results in increased movement of internal elements within the battery First generation McGRATH devices left on for minutes may increase the risk of internal short x D C Battery Manufacturing process variation Note No battery defects or non conformances have been identified as part of investigations or Medtronic s on site process controls review x D | For the 1st Generation Laryngoscopes model 300-000-000: • Immediately discontinue use of the devices being removed (300-000-000). • Uninstall and dispose of the battery assembly from the devices being removed. • Return the removed device (300-000-000) to Medtronic. • Pass this notice on to all who need to be aware within your organization or to any organization where the potentially affected product with the specified serial numbers have been transferred or distributed. • Please complete the enclosed Customer Acknowledgement Form even if you do not have affected inventory. For the 2nd Generation Laryngoscope model 301-000-000: • Continue using Next-Generation McGRATH™ Video Laryngoscopes Next-Generation (301-000-000). • Follow Important Battery Handling Information stated in the IFU and provided in this notice and maintain a copy of the addendum with the IFU for McGRATH™ MAC Video Laryngoscopes. • Pass this notice on to all who need to be aware within your organization or to any organization where the potentially affected product with the specified serial numbers have been transferred or distributed. • Please complete the enclosed Customer Acknowledgement Form even if you do not have affected inventory. | 2024-07-28 09:03
Randox Laboratories Ltd | Liquid Assayed Chemistry Premium Plus Level 1 | - | - | Randox Laboratories can confirm that there has been a transcription error on the Instructions For Use IFU x D for Liquid Assayed Chemistry Premium Plus Level lot UL The Target and range for Sodium ISE Indirect x D on the Abbott Alinity Architect c ci Systems have been incorrectly listed as Target Range and x D should be Target Range Updated IFUs are available on www randox com please discard the x D incorrect version and download the updated revision | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2024-07-28 09:55
Sysmex Corporation | Flow Cytometer XF-1600 | 0 | - | Due to an issue in the XF IVD software installer version and a folder which contains x D measurement settings for rotor position is missing The missing folder would cause the sample in x D rotor position to be measured with settings for position Measurement settings include x D Gain PMTv Threshold Channel Laser Flow rate Stop condition | - | 2024-07-28 11:04
Beckman Coulter Ireland Inc. | IRISpec CA/CB/CC | 161 | - | Beckman Coulter has become aware of an increase in customer complaints for intermittent IRISpec CB glucose control failure x D resulting in false positive results when used in conjunction with iChemVELOCITY Urine Chemistry Strips PNs and x D x D x D If a QC run fails the system will not allow the operator to run patient samples The QC run should be repeated and pass successfully before running patient samples x D x D This issue affects IRISpec CB control lot numbers and higher when analyzed with iChemVELOCITY Urine Chemistry Strips PN lot numbers M and higher and lot numbers M and higher | Beckman Coulter has updated its control ranges to move the negative Glucose control from IRISpec CB to IRISpec CC. Customers are to update the CB and CC control ranges for Glucose, instructions on how to complete this are provided in the customer letter. The updated Glucose ranges supersede the ranges listed in the IRISpec CB control Instructions for Use (IFU) PN C36362.AC, C36361.AC and the assay sheet C39722.AC and C39721.AC. Use only the specifications in the customer letter. No retrospective review of QC or patient sample results is necessary. A further investigation is taking place to prevent the re-occurrence of this problem. Customers will be provided more information when available. | 2024-07-28 22:46
Beckman Coulter Ireland Inc. | IRISpec CA/CB/CC | 227 | - | Beckman Coulter has become aware of an increase in complaints for intermittent IRISpec CB failures for Ketone on the x D iChemVELOCITY Urine Chemistry System Instructions For Use PN FH and DxU c Iris Urine Chemistry System x D Instructions For Use PN C AA x D x D If a QC run fails the system will not allow the operator to run patient samples The QC run should be repeated and pass successfully before running patient samples x D x D This issue affects IRISpec CB control lot numbers and higher when analyzed with iChemVELOCITY Urine Chemistry Strips PN lot numbers M and higher | • Beckman Coulter has updated its Ketone control range for IRISpec CB. • Update the CB control ranges for Ketone as indicated in Table 1 on the customer letter. For instructions, refer to Addendum A: “Updating the QC ranges for Ketone in the Analyzer”. • The above Ketone ranges supersede the ranges listed in the IRISpec CB control Instructions For Use (IFU) PN C36362.AC and the assay sheet C39722.AC. Use only the specifications in this letter. • No retrospective review of QC or patient sample results is necessary. • Beckman Coulter recommends posting this Recall Notification Letter on/or near the analyzer. | 2024-07-28 22:58
Abiomed Europe GmbH | Impella CP | 225 | - | Currently there is no Abiomed labeling IFU content related to this potential interaction so Abiomed will update the IFU to include information on this potential device interaction x D | "1-Notify Health Authorities within the stipulated timeline per local regulations. 2-Abiomed will only start communicating to customers when the FSN is ready. 3-the Field Safety Notice (FSN) is not available yet as Abiomed is still finalizing the language. 4-No product design or manufacturing issues are related to this potential interaction, and hemodynamic support will not be affected as a result of optical sensor damage. Instructions for Use (IFU) will be updated in order to reflect an additional warning and instructions to prevent this event from occurring." | 2024-07-29 19:53
Smith Medical | Nasopharyngeal Airway, Double Swivel Connector, 15mm Y Piece, Portex™ Orator Speaking Valve, Thermovent™ 1200 | 7620 | - | Smiths Medical is issuing this letter to notify you of a potential issue with the packaging sterile seals on specific Nasopharyngeal Airway Double Swivel Connector mm Y Piece Portex Orator Speaking Valve Thermovent products packaged between January and August This letter details the issue and the required steps for you to perform x D Issue Smiths Medical has identified a timeframe where there is uncertainty in the seal integrity of the sterile packaging of sterilized Nasopharyngeal Airway Double Swivel Connector mm Y Piece Portex Orator Speaking Valve and Thermovent products | Check all inventory locations within your institution for the affected catalog numbers and lot numbers listed in the notification and discontinue use. Destroy all affected products following your institution’s process for destruction. If destroying is not immediately possible at your facility, then the product should be quarantined until disposal. | 2024-08-01 18:45
Löwenstein Medical Technology GmbH + Co. KG | LUISA (LM150TD) | 16 | - | In exceptional rare cases an initialization problem with the ventilator s internal acoustic alarm system may result in only the visual alarm signal being emitted but not the acoustic alarm signal This only occurs at alarm volume x D When using alarm volume or higher a visual as well as an acoustic alarm is guaranteed The ventilator is delivered with an alarm volume of set as standard x D From a technical point of view this initialization problem can occur with firmware versions up to version x D There are known incidents from the market where this problem occurred at alarm volume This corresponds to a percentage probability of occurrence of with around million potential uses To date L wenstein Medical Technology has not received any reports of harm to patients in relation to this issue | Löwenstein Medical Technology will release a firmware update to rectify the cause of the error described above. The update will be made available to users free of charge. The firmware update is expected to be available in October 2024. | 2024-08-04 11:25
Beckman Coulter Inc. | IMMAGE Immunoglobulin M (IgM) | 77 | - | As a result of internal investigations Beckman Coulter has discovered that IMMAGE Immunochemistry Systems IGM Immunoglobulin M reagent lots listed above are not meeting the sensitivity statement in the IFU x D Sensitivity is defined as the lowest measurable concentration which can be distinguished from zero with confidence Sensitivity for IGM determination is mg dL |  Discontinue use of affected IMMAGE Immunochemistry Systems IGM Immunoglobulin M reagent lots and please follow product replacement instructions below.  Review the Safety Data Sheet (SDS) and discard according to your laboratory’s Standard Operating Procedures and/or your local regulations. The SDS is available on www.beckmancoulter.com.  Beckman Coulter recommends sharing the content of this letter with your laboratory and/or medical director regarding the need to perform a retrospective review of the patient results generated from the affected IMMAGE Immunochemistry Systems IGM Immunoglobulin M reagent lots to determine if retesting is necessary. | 2024-08-04 13:32
Randox Laboratories Ltd | Liquid Assayed Chemistry Premium Plus Level 2 | - | - | Randox Laboratories can confirm that the RX Series target and range for gamma GT in the x D Liquid Assayed Chemistry Premium Plus Level lot UN has been assigned incorrectly x D The Target and range for gamma GT is listed as Target Range and should be Target x D Range | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2024-08-04 14:16
Smiths Medical ASD, Inc. | Level 1 Fast Flow Fluid Warmer | 81 | - | ISSUE x D Smiths Medical is issuing this letter to notify you of two potential issues with Level Fast Fluid Flow Fluid Warmers and Level NORMOFLO Irrigation Fluid Warmers This letter details the issues the affected models and the required actions for users to perform x D x D Issues x D Smiths Medical has become aware of two issues related to the Level Fast Flow Fluid Warmers and Level NORMOFLO Irrigation Fluid Warmers x D The return tube assembly inside the warmers may exhibit stress cracking that can propagate over time and potentially result in leaking recirculating fluid Leaking recirculating fluid may potentially damage other components within the warmer such as the printed circuit board x D x D The cable inside the warmers that connects to the external on off switch may potentially exhibit delamination of insulation which may propagate over time and may lead to corrosion and a break in the electrical circuit | Required Actions for Users: To address the described risk, users should follow the instructions below: 1. Prior to use, confirm that the device powers on and passes all the alarm tests outlined in the Setup for Use section in the Operator’s Manual and described below in point A. If the device fails during these alarm tests or does not power on, remove the device from service, submit a complaint to Smiths Medical and await further instructions from Smiths Medical. 1. Alarm Tests outlined in Set Up for Use in Operators Manual: i. Press and hold the Alarm Test button on the Fluid Warmer’s Power and Alarm Test Panel. 1. All Fluid Warmer visual alarm LEDs illuminate and the audible alarm signals beep. ii. Release the Alarm Test button; the Over Temperature alarm continues. iii. Clear the Over Temperature alarm condition. 1. Turn the Fluid Warmer OFF and then back ON. 2. The Air Detector/Clamp runs a Power On Test. (Only on the H-1200 Fast Flow Fluid Warmer) 3. The Air Detector/Clamp goes into Automatic operation (Only on the H-1200 Fast Flow Fluid Warmer) Users should have back up devices such as manual pressure cuffs or additional Level 1 Fast Flow Fluid Warmer or Level 1 Normoflo Irrigation Fluid Warmer available for emergency situations. | 2024-08-04 16:46
Zoll medical Corporation | Automatic external defibrillator Powerheart AED G5 | 49 | - | Customers may not remove the protective film during the deployment of the product If the protective film is left adhered to the front bezel it may prevent the user from actuating the shock button | - Locate the affected devices. - Remove the protective film in accordance with Appendix A in the attachment. | 2024-08-05 13:04
Canon Medical Systems Corporation | INTERVENTIONAL ANGIOGRAPHY SYSTEM Alphenix | 12 | - | It has been discovered that there is a possibility that the fixing screws in C arm and table that are part of x D the X ray interventional systems may become loose x D As a result abnormal noises may be generated from the suspension devices x D Also sensor errors may be displayed x D The cause is the inadequacy of the work procedure in the factory x D However the fixing parts do not fall | Check for the tightening of the screws and properly fixing to correct this issue will be implemented on your system. When the preparation becomes available, your service representative will contact you for an appointment to schedule the installation. | 2024-08-05 18:13
Smith Medical | Portex™ Blue Line Siliconised PVC Tracheotomy Tube | 910 | - | This is an update to the previous communication dated April Smiths Medical identified additional lots that were potentially affected by the manufacturing defect identified in the original notice Therefore Smiths Medical is expanding the scope of this issue to include the additional lots in Table x D Content that was updated or differs from the previous communication date April is shown in red font x D Smiths Medical is issuing this letter to notify you of a potential issue with the Portex Blue Line Siliconised PVC Tracheotomy Tube The following information details the issue and the required steps for you to perform x D Issue x D Smiths Medical has identified an issue related to the neck plate flange of Portex Blue Line Siliconised PVC Tracheotomy Tube Specifically this failure mode can manifest itself during use as a complete or partial detachment of the neck plate from the tracheostomy tube on Portex Blue Line Classic Tracheotomy tubes | There are one hundred and thirty-seven (137) additional lots that were identified after 06 March 2024 which are potentially impacted by this manufacturing issue. The one hundred and thirty-seven additional lots identified were distributed between 19 FEB 2019 and 02 FEB 2024. Please refer to Table 1 below for a list of the additional lot numbers that were shipped to Saudi Arabia. Smiths Medical Actions: Smiths Medical is sending this notification to all customers who received product(s) from Smiths Medical listed above. Smiths Medical has initiated a global ship hold on impacted lots to ensure affected product is no longer distributed. Smiths Medical will provide credit to affected customers. Please contact your Smiths Medical customer service representative after the Customer Response Form has been provided to EMEA-FSN@icumed.com, to coordinate credit. Customer Required Actions: Please review the Portex™ Blue Line Siliconised PVC Tracheotomy Tube Instructions for Use with heightened awareness prior to use. 1) Check all inventory locations within your institution for the affected catalog numbers and lot numbers listed in the notification and discontinue use. Discard all affected products following your institution’s process for discarding. If discarding is not immediately possible at your facility, then the product should be quarantined until disposal. 2) Share this notification with all potential users of the device to ensure they are aware of this notification and proposed mitigations. If the devices are used at another location, please | 2024-08-07 16:45
Philips Medical Systems Nederland B.V. | AZURION 7 M20 | 4 | - | Philips has identified that the motorized longitudinal movement of the FlexArm stand may be inconsistent not smooth and eventually may become unavailable This issue is due to grease leakage from the bearings and or anti corrosion oil applied on the bearings because of which the lubrication on the ceiling rail may be excessive reducing the friction between the rails and the friction wheel see Figure x D When this issue occurs the system will display the messages shown below on a the Touchscreen Module TSM and b the FlexVision and or FlexSpot | Philips will replace the existing longitudinal bearings on the affected Azurion 7M20 systems with FlexArm stand with redesigned longitudinal bearings free of charge. Philips will contact all affected customers once the redesign longitudinal bearings are available (reference FCO72200585). As of the date of this Urgent Field Safety Notice, Philips expects the solution to be available by Q3 2025. | 2024-08-08 11:07
GE Healthcare | Centricity Universal Viewer Zero Footprint | 27 | - | GE HealthCare has become aware of an issue in Centricity Universal Viewer Zero Footprint Client ZFP versions v SP x and SP x where the latest addended report is not shown by default to the user The issue occurs when the addendum is created on the same day as the original report but at a different time and only when launched on the following browsers x D Internet Explorer IE x D Internet Explorer IE x D Firefox x D Users continue to have access to all reports via in the Series Selector However since the latest addendum may not be shown by default in the ZFP Viewer this could lead to potential misdiagnosis if a user relies only on the information in the original report and does not review the addended report x D | You can continue to use your device. Please ensure all potential users in your facility are made aware of this safety notification and the recommended actions below. Follow one of the below recommendations to ensure the latest addendum is being reviewed: 1. Use Microsoft Edge or Google Chrome browsers when launching ZFP Viewer. OR 2. If you cannot use the browsers listed in #1, to find the addendum / latest report, review all the reports available for the study in the Series Selector. Navigation to the series selection can be found in the User Manual under Section 2.3.3 of The Series Selector, with subheading as “Open a report/document using series selector”. | 2024-08-08 13:14
Datex-Ohmeda, Inc. | Aisys CS2 anesthesia devices equipped with the Auxiliary Common Gas Outlet (ACGO) | 1214 | - | GE HealthCare has become aware of a potential issue that may arise if the limb of a patient breathing circuit is mistakenly connected to the Auxiliary Common Gas Outlet ACGO port x D x D The ACGO when selected provides fresh gas flow to auxiliary manual breathing systems e g a Mapleson D circuit connected to the ACGO port x D x D If the expiratory limb of a patient breathing circuit is mistakenly connected to the ACGO port rather than to the Advanced Breathing System port it can result in excessive pressure within the breathing system The Anesthesia devices that are potentially impacted have several alarms that would alert the clinician of increased breathing system pressures However in the unlikely event that the misconnection to the ACGO port is not noticed the increased pressure can lead to patient injury x D | GE HealthCare will provide a customer-installable kit containing covers for the ACGO Port and ACGO Switch of impacted devices, at no cost to customers who indicate in the attached acknowledgement form that they do not have the covers. | 2024-08-11 15:21
Datex-Ohmeda, Inc. | Avance CS2 anesthesia devices equipped with the Auxiliary Common Gas Outlet (ACGO) | 1214 | - | GE HealthCare has become aware of a potential issue that may arise if the limb of a patient breathing circuit is mistakenly connected to the Auxiliary Common Gas Outlet ACGO port x D x D The ACGO when selected provides fresh gas flow to auxiliary manual breathing systems e g a Mapleson D circuit connected to the ACGO port x D x D If the expiratory limb of a patient breathing circuit is mistakenly connected to the ACGO port rather than to the Advanced Breathing System port it can result in excessive pressure within the breathing system The Anesthesia devices that are potentially impacted have several alarms that would alert the clinician of increased breathing system pressures However in the unlikely event that the misconnection to the ACGO port is not noticed the increased pressure can lead to patient injury x D | GE HealthCare will provide a customer-installable kit containing covers for the ACGO Port and ACGO Switch of impacted devices, at no cost to customers who indicate in the attached acknowledgement form that they do not have the covers. | 2024-08-11 15:21
Datex-Ohmeda, Inc. | Avance CS2 Pro, Aisys, Avance, Amingo, Aespire 100, Aespire View, Aespire 7900, and Aespire 7100 anesthesia devices equipped with the Auxiliary Common Gas Outlet (ACGO) | 1214 | - | GE HealthCare has become aware of a potential issue that may arise if the limb of a patient breathing circuit is mistakenly connected to the Auxiliary Common Gas Outlet ACGO port x D x D The ACGO when selected provides fresh gas flow to auxiliary manual breathing systems e g a Mapleson D circuit connected to the ACGO port x D x D If the expiratory limb of a patient breathing circuit is mistakenly connected to the ACGO port rather than to the Advanced Breathing System port it can result in excessive pressure within the breathing system The Anesthesia devices that are potentially impacted have several alarms that would alert the clinician of increased breathing system pressures However in the unlikely event that the misconnection to the ACGO port is not noticed the increased pressure can lead to patient injury x D | GE HealthCare will provide a customer-installable kit containing covers for the ACGO Port and ACGO Switch of impacted devices, at no cost to customers who indicate in the attached acknowledgement form that they do not have the covers. | 2024-08-11 15:21
GE Healthcare | Centricity PACS 4.0 | 51 | - | GE HealthCare has become aware of an issue where image storage is not possible when Centricity PACS reaches the maximum number of unique image identifiers Ckeys If this occurs it can result in acquisition of a partial study In an unlikely scenario that the partial study is unnoticed it can lead to misdiagnosis Centricity PACS is able to uniquely identify approximately billion images e g equivalent to million exams with images each However there is no notification to the user when the maximum number of unique image identifiers has been reached | You can continue to use your device if it has not reached the maximum number of Ckeys. In order to determine the remaining number of available Ckeys: • If GE HealthCare is remotely monitoring your Centricity PACS system as part of an existing service agreement, the number of available Ckeys is assessed routinely. If your device is close to maximum capacity, the GE HealthCare support team will contact you to discuss next steps. • If GE HealthCare is not remotely monitoring your system, please contact the GE HealthCare support team who will help you check your Ckey status GE HealthCare will provide an updated manual with CKey monitoring instructions at no cost to you. A GE HealthCare representative will contact you with further information | 2024-08-11 16:10
Fisher & Paykel Healthcare | Airvo 3 | 30 | - | This field safety corrective action relates to specific versions of software of the Airvo and how the x D device responds when flow alignment alarm occurs x D When this happens in Airvo devices set up with High Pressure Oxygen HPO and running software x D version or earlier the device will deliver room air only If this happens a patient may experience x D oxygen desaturation that could lead to hypoxia x D Software version updates the software algorithm and ensures the target therapy continues in the x D event that this alarm occurs | ACTIONS REQUIRED FOR AFFECTED PRODUCT IN YOUR INVENTORY Please immediately follow these steps to support this field safety corrective action. Step 1: Identify Affected Product a. Check the following on the product label underneath the base of the device or the label on the box, and confirm it is within the affected product range. i. Model Number reference (REF) ii. iii.Serial Number (SN) Manufacturing Date (YYYY-MM-DD) b. If the device is still within your inventory, please put the product in quarantine to prevent ongoing distribution until we have contacted you with next steps. Step 2: Complete and return the attached response form. a. Complete the First Consignee Field Safety Notice Response Form attached to this letter. b. Return the form to your F&P Representative. (Contact details are on the form). Step 3: Await instructions from your F&P Representative. a. After we receive your response form, an F&P Representative will contact you regarding the update of your software. ACTIONS FOR AFFECTED PRODUCT YOU MAY HAVE DISTRIBUTED Step 1: Identify Distributed Product a. Review your sales records and identify if any Airvo 3 devices manufactured prior to 14 March 2024 have been distributed to your customers. Step 2: Complete and return the attached response form. a. Complete the Second Consignee Field Safety Notice Response Form attached to this letter. b. Return the form to your F&P Representative. (Contact details are on the form). Step 3: Create list of affected customers. a. If you identify that Affected Product may have been distributed to your customers, create a list of affected customers for tracking purposes using the Second Consignee Spreadsheet provided in the email. Step 4: Notification Letter and Response Form. a. Create a Field Safety Notice Letter and Field Safety Notice Response Form for each customer. b. Use the templates listed below and edit only the text in red: i. Second_Consignee_ Field Safety Notice_Initial Letter ii. Second_Consignee_ Field Safety Notice_Response Form a. Send the Second Consignee Field Safety Notice Letter and Second Consignee Field Safety Notice Response Form to all affected customers within 10 business days of receiving this letter via email. Step 5: a. Update the relevant information on the Second Consignee Spreadsheet. NOTE: The Second Consignee Spreadsheet and all Response Forms must be kept and sent to your F&P Representative. Step 6: Follow-Up a. Where a customer fails to respond to the Field Safety Notice Letter within 10 business days of initial contact via email, please follow up minimum of three times, with a Reminder Letter once every subsequent 10 business days. b. Create the Reminder Letter using the Second Consignee_ Field Safety Notice _Reminder Letter template. Enter the type of reminder letter (First, Second or Final) and the date on which you will send the Reminder Letter. c. Enter the date and summary of letters sent and contact attempts in the Second Consignee Spreadsheet | 2024-08-12 10:03
HMC PREMEDICAL SPA | NASOGASTRIC FEEDING TUBE | 1 | - | Erasure of the graduation markings on tubes x D The problem was observed to be due to the use of incorrect cleaning and disinfection x D processes and excessive rubbing on the outer surface of the tubes | Product Defect Correction | 2024-08-12 12:55
Boston Scientific | Vercise Genus | 2 | - | Remind to follow the steps outlined in the labeling Instructions for Use IFU to implant Vercise Genus Deep Brain Stimulation DBS Implantable Pulse Generators IPGs within a subcutaneous pocket Device header feedthrough FT wire break s have occurred only in rechargeable Vercise Genus DBS IPGs that were implanted submuscular in the pectoral location Per device labeling IFU the Vercise Genus DBS IPG is intended to be implanted in a subcutaneous pocket | • When implanting Vercise Genus DBS IPGs: Follow existing device labeling/IFU regarding implant location within a subcutaneous pocket. • For any patient with a rechargeable Vercise Genus DBS IPG implanted submuscular in the pectoral region: Monitor per relevant IFU recommendations for any clinical observations of high monopolar impedances, undesired sensation, sudden loss of therapy, return of pre-implant symptoms and/or Bluetooth connectivity challenges, as these may be signals of potential FT wire break(s). | 2024-08-12 15:57
Alcon Laboratories, Inc. | DAILIES TOTAL1 | 392 | - | We identified an isolated quality issue with a material supplied by a third party vendor that was used in the production of these specific contact lens lots at a single manufacturing site As a result the identified contact lens lots may not meet Alcon quality and or performance standards for the entirety of their labeled shelf life | We ask that you stop any further distribution of contact lenses from the identified lots and request that you contact any eye care practitioners, patients, or others to whom you have distributed the identified lots. | 2024-08-12 17:09
Alcon Laboratories, Inc. | DAILIES TOTAL1 for Astigmatism | 392 | - | We identified an isolated quality issue with a material supplied by a third party vendor that was used in the production of these specific contact lens lots at a single manufacturing site As a result the identified contact lens lots may not meet Alcon quality and or performance standards for the entirety of their labeled shelf life | We ask that you stop any further distribution of contact lenses from the identified lots and request that you contact any eye care practitioners, patients, or others to whom you have distributed the identified lots. | 2024-08-12 17:09
werfen | HemosIL® AcuStar ADAMTS13 Activity | 3 | - | Issue Description and Impact x D Recently we have received complaints of serious injury in patients where HemosIL AcuStar ADAMTS Activity results were below the medical decision level activity and a comparator assay reported results above the medical decision level and in certain instances patient treatment was initiated based on the ACL AcuStar results The increased adoption of new Thrombotic Thrombocytopenic Purpura TTP therapies that potentially increase the risk of bleeding reinforces the need to take ADAMTS activity results in conjunction with other clinical and laboratory findings In addition some of these reports include use of HemosIL AcuStar ADAMTS Activity with pediatric patients despite the assay being labeled for use in the adult population x D Previous notifications emphasized the need to assess ADAMTS activity results with other clinical and laboratory findings with the following action x D A test result of will automatically be flagged by the instrument with the following notification Assay intended for diagnosis and monitoring of TTP Check sample integrity and verify results against other clinical and laboratory findings prior to reporting x D To further reinforce the appropriate usage of HemosIL AcuStar ADAMTS Activity we advise the following x D Results of this assay should always be interpreted in conjunction with other clinical and laboratory findings x D The use of the assay to guide patient therapy plan e g plasma exchange rituximab and caplacizumab has not been evaluated x D The assay has not been validated for use in the pediatric population x D The assay has not been validated for evaluating TTP relapse or recurrence | Based on the above, please take the following immediate actions: • Use HemosIL AcuStar ADAMTS13 Activity based on its labeled intended use. • Evaluate patient results based on the above limitations. • Assess the need to perform additional testing prior to taking clinical action for patient results that may be inconsistent with other clinical and laboratory findings. • Share this notification with your laboratory staff and follow your internal procedures. • Forward this notification to all affected locations within your facility. • Retain a copy of this notification for your records. | 2024-08-13 11:11
werfen | HemosIL® AcuStar ADAMTS13 Activity | 7 | - | Issue Description and Impact x D Recently we have received complaints of serious injury in patients where HemosIL AcuStar ADAMTS Activity results were below the medical decision level activity and a comparator assay reported results above the medical decision level and in certain instances patient treatment was initiated based on the ACL AcuStar results The increased adoption of new Thrombotic Thrombocytopenic Purpura TTP therapies that potentially increase the risk of bleeding reinforces the need to take ADAMTS activity results in conjunction with other clinical and laboratory findings In addition some of these reports include use of HemosIL AcuStar ADAMTS Activity with pediatric patients despite the assay being labeled for use in the adult population x D Previous notifications emphasized the need to assess ADAMTS activity results with other clinical and laboratory findings with the following action x D A test result of will automatically be flagged by the instrument with the following notification Assay intended for diagnosis and monitoring of TTP Check sample integrity and verify results against other clinical and laboratory findings prior to reporting x D To further reinforce the appropriate usage of HemosIL AcuStar ADAMTS Activity we advise the following x D Results of this assay should always be interpreted in conjunction with other clinical and laboratory findings x D The use of the assay to guide patient therapy plan e g plasma exchange rituximab and caplacizumab has not been evaluated x D The assay has not been validated for use in the pediatric population x D The assay has not been validated for evaluating TTP relapse or recurrence | Based on the above, please take the following immediate actions: • Use HemosIL AcuStar ADAMTS13 Activity based on its labeled intended use. • Evaluate patient results based on the above limitations. • Assess the need to perform additional testing prior to taking clinical action for patient results that may be inconsistent with other clinical and laboratory findings. • Share this notification with your laboratory staff and follow your internal procedures. • Forward this notification to all affected locations within your facility. • Retain a copy of this notification for your records. | 2024-08-13 11:11
werfen | HemosIL® AcuStar ADAMTS13 Activity | 7 | - | Issue Description and Impact x D Recently we have received complaints of serious injury in patients where HemosIL AcuStar ADAMTS Activity results were below the medical decision level activity and a comparator assay reported results above the medical decision level and in certain instances patient treatment was initiated based on the ACL AcuStar results The increased adoption of new Thrombotic Thrombocytopenic Purpura TTP therapies that potentially increase the risk of bleeding reinforces the need to take ADAMTS activity results in conjunction with other clinical and laboratory findings In addition some of these reports include use of HemosIL AcuStar ADAMTS Activity with pediatric patients despite the assay being labeled for use in the adult population x D Previous notifications emphasized the need to assess ADAMTS activity results with other clinical and laboratory findings with the following action x D A test result of will automatically be flagged by the instrument with the following notification Assay intended for diagnosis and monitoring of TTP Check sample integrity and verify results against other clinical and laboratory findings prior to reporting x D To further reinforce the appropriate usage of HemosIL AcuStar ADAMTS Activity we advise the following x D Results of this assay should always be interpreted in conjunction with other clinical and laboratory findings x D The use of the assay to guide patient therapy plan e g plasma exchange rituximab and caplacizumab has not been evaluated x D The assay has not been validated for use in the pediatric population x D The assay has not been validated for evaluating TTP relapse or recurrence | Page 2 of 2 • Mandatory Actions Based on the above, please take the following immediate actions: • Use HemosIL AcuStar ADAMTS13 Activity based on its labeled intended use. • Evaluate patient results based on the above limitations. • Assess the need to perform additional testing prior to taking clinical action for patient results that may be inconsistent with other clinical and laboratory findings. • Share this notification with your laboratory staff and follow your internal procedures. • Forward this notification to all affected locations within your facility. • Retain a copy of this notification for your records. | 2024-08-13 11:17
Becton, Dickinson and Company BD Biosciences | BD FACSDuet™ Sample Preparation System | 3 | - | Through a customer complaint investigation BD identified a specimen volume bias issue with the BD FACSDuet Sample Preparation System when using software version x D x D When the Multidispense feature is enabled in the BD FACSDuet Sample Preparation System software and the user is preparing a two tube assay the specimen volume dispensed into the first tube will be slightly higher than the specimen volume dispensed into the second tube For assays where absolute counts are calculated using BD Trucount Tubes an increase in the volume may result in higher than expected absolute counts in the first tube compared to the second tube x D x D This issue will be observed when preparing the following assays on the BD FACSDuet Sample Preparation System x D x D BD Multitest IMK Kit with BD Trucount Tubes x D x D Note If using the BD Multitest IMK Kit with BD Trucount Tubes a QC message will appear on your report if the bias is outside the acceptable range per assay requirements x D x D BD Multitest Color TBNK with BD Trucount Tubes only if consecutive replicates of two are run x D x D For all other quantitative assays utilizing BD Trucount Tubes in consecutive replicates of two please follow your institution s procedures to determine if a review of reported results is necessary x D x D This issue does not affect assays that are determining percent positive semi quantitative information x D | Clinical User Actions 1. Identify Software Version installed in BD FACSDuet™ Sample Preparation System at your site. Refer to Appendix 1 2. Cease use of the BD FACSDuet™ Sample Preparation System with software version 1.4.1. until the Multidispense feature has been disabled. If not already disabled, refer to Appendix 2 for instructions on how to disable the BD FACSDuet™ Sample Preparation System Multidispense feature. BD Actions: 1. On return of the completed response form, BD will schedule a mandatory upgrade of your BD FACSDuet™ Sample Preparation System software to version 1.4.2. 2. A corrective and preventive action was initiated to confirm root cause and identify actions to prevent reoccurrence. Customer Actions: • Complete and return the Customer Response Form even if you no longer have any of the impacted product remaining in your facility by 30th August 2024. • Circulate this notice to all those who need to be aware within your organisation or to any organisation where the potentially affected product has been transferred. • If you experience any issues, please report as a complaint as per your normal process. | 2024-08-13 13:56
Haemonetics Corporation | PLT & PLS Set w/ Filter | 6976 | - | Haemonetics has received reports from the field of instances in which there has been a significant leak of blood from the Latham bowl into the MCS equipment during the apheresis procedure The leak results when the bottom of the bowl becomes detached which then leads to a termination of the procedure and required cleaning and servicing of the equipment before the next use The objective of this FSN is to inform our customers of this possible occurrence and the options available x D Frequency of failures and complaints The estimated rate of occurrence is approximately per million x D Adverse events There have been no reports of adverse events associated with this leak | 1. Inform all your impacted customers of this FSN. For your convenience, we have attached a sample customer letter. If required, translate this letter into local language 2. If required by your local regulations, inform the appropriate authorities. Note: Haemonetics will inform the appropriate European authorities (including EEA and Turkey). 3. Complete the attached acknowledgement form in its entirety whether or not you have distributed affected product. Once complete, return the form to Haemonetics following the instructions on the form. Your response is vital to our monitoring of the effectiveness of this action. | 2024-08-15 09:47
Ortho Clinical Diagnostics | ORTHO Optix™ Reader | 0 | - | Starting with SW V customers may process multiple tests on a single Cassette followed by a single imaging event without having to image the Cassette for each order individually For the most efficient Cassette usage associate the first test to the columns and determine if additional tests need to be assigned to the remaining columns If so associate all and process all tests prior to the imaging event | • Until a Software Update is released to resolve the issue, a partial Cassette must not be re-used after the first imaging event on the same ORTHO Optix Reader. • Note: Starting with software version V2.0.0.2870 customers may process multiple tests on a single cassette followed by a single imaging event (without having to image the cassette for each order individually). • Complete and return the enclosed Confirmation of Receipt form no later than Month DD, 2024. | 2024-08-15 12:32
Siemens Medical Solution USA, Inc | ACUSON JUNIPER | 52 | - | The system incorrectly performs a unit conversion of liters to milliliters using a multiplier of one hundred instead of using the correct multiplier of one thousand in specific situations described below The result of this conversion error is the presentation of a x under estimated measured value This unit conversion issue displays an inaccurate value that is x smaller than the actual measurement ONLY after a system is changed from the factory default configuration units of measure to display the following x D Milliliters per second unit displayed as either ml sec or mL sec | Siemens Healthineers will correct this issue with a free-of-charge software update to your ACUSON Juniper | 2024-08-18 13:45
BIOMÉRIEUX | The VIDAS® KUBE™ system | 9 | - | Description of the issue x D In May bioM rieux received complaints from the field related to the inability to log into the VIDAS KUBE instrument x D VIDAS KUBE CLINICAL Ref VIDAS KUBE INDUSTRY Ref having installed VIDAS KUBE software x D version The customer could no longer use the VIDAS KUBE instrument until the database was restored with bioM rieux x D support x D | An internal investigation (INV-20646) was initiated to identify the root cause and resolve the issue. While the investigation is still ongoing, the following were identified: • The investigation confirmed that the issue only concerned VIDAS® KUBE™ CLINICAL (Ref. 423912) and VIDAS® KUBE™ INDUSTRY (Ref. 423905) instruments having installed VIDAS® KUBE™ V1.1 software. • This issue is due to corrupted files in the VAULT database, which are blocking the VAULT unsealing process. VAULT is a programming tool for creating, managing and storing secrets specific to a development project (login, encryption key, password, tokens, etc.). This makes it possible to verify and supervise access, and to reinforce authentication processes, including the generation of rotating passwords for database access. These corrupted files appear when VAULT writes to the database while VAULT software is stopped. • Due to this issue, the customer has no access to the VIDAS® KUBE™ V1.1 software and could no longer use the VIDAS® KUBE™ instrument. • A corrective patch "VIDAS KUBE - Patch to unlock VAULT.zip" has been developped and is already available on the GCS portal to correct and prevent the issue. It can be directly installed at customer site or via VILINK into the VIDAS® KUBE™ instrument or on the computer with bioMérieux support. • The risk associated to this issue are the followings: o a "loss of data" and "loss of traceability" between the last daily backup and when the issue occurred if a full system backup is restored with bioMérieux support. o an "inconvenience for the user" as the customer could no longer use the VIDAS® KUBE™ CLINICAL (Ref. 423912) and VIDAS® KUBE™ INDUSTRY (Ref. 423905) instruments. | 2024-08-19 11:59
CareFusion 303, Inc. | BD Alaris™ System | 2868 | - | BD has identified an increase in bent contacts on Female Left IUI connectors on the devices listed in Table BD was made aware of this issue through complaints identified during the device check in process and has not received any reports of this issue occurring during patient use A defective pin on certain female left IUI connectors that are not detected prior to use could result in an inoperable device or an operable device where adjoining device cannot be powered on at start up Examples of defective pins are provided in Appendix x D x D To date all instances of bent Female Left IUI connectors related to this issue have been identified prior to clinical use during the device check in process There are no reports of harm associated with this issue x D | Corrective actions by BD BD has implemented additional IUI inspection criteria during the manufacturing process for devices and spare parts. Actions to be taken by BD To reinforce existing instructions around IUI inspection, the following labeling is in the process of being updated: • BD Alaris™ System Maintenance User Manual v12.1.1 • BD Alaris™ System with Guardrails™ Suite MX For BD Alaris™ PCU Model 8015 and Alaris™ PCU Model 8015 Software Version 12.1.1 and Supported Modules Actions to be taken by Clinical Users 1. Per current labeling, customers should inspect all IUI connectors prior to use. a. Example images of defective pins are provided in Appendix 1. 2. In the event a bent Female/Left IUI connector is identified on affected serial numbers or parts, do not use the device and the send to Biomedical Engineers or Qualified Service Personnel for repair. Customer Actions: 1. Circulate this Field Safety Notice to all those within your organisation that may use the BD Alaris™ System devices listed in table 1. 2. Complete the Customer Response Form on page 3 and return it to abdulrahman.hanbazazah@bd.com or no later than the 30th September, 2024. 3. If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 30th September 2024. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. | 2024-08-19 13:34
STERIS Corporation | HOOKUP | 4 | - | Please find attched FSN letters | Please find attched FSN letters | 2024-08-19 14:23
.Siemens Healthcare Diagnostics Inc | Atellica® CH UCFP | 17780 | - | Issue Description x D Siemens Healthineers has confirmed the potential for biased quality control QC and patient results when using Atellica CH Urinary Cerebrospinal Fluid Protein UCFP lot on the Atellica CH and Atellica CI analyzers Siemens internal investigation confirmed when using lot QC may recover outside of the allowable control limits for urine chemistry and spinal fluid control levels x D All results generated using reagent Lot are considered impacted Siemens Healthineers is currently conducting a root cause investigation and has identified an issue with a specific lot of raw material used to manufacture UCFP reagent No other in date lots have demonstrated this issue however Siemens Healthineers has introduced additional quality testing of in date lots until the root cause investigation is complete x D Kindly review the attached UFSN | Customer Actions: • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Discontinue use of and discard the kit lot listed in the table above (Products Section). • Complete and return the Field Correction Effectiveness Check and indicate product replacement needs on the form attached to this letter within 30 days. • Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached UFSN | 2024-08-19 17:05
Siemens Medical Solution USA, Inc | ACUSON Redwood | 2 | - | The system incorrectly performs a unit conversion of liters to milliliters using a multiplier of one hundred instead of using the crrect multiplier of one thousand in specific situations described below The result of this conversion error is the presentation of a x under estimated measured value This unit conversion issue displays an inaccurate value that is x smaller than the actual measurement ONLY after a system is changed from the factory default configuration units of measure to display either of the following x D Milliliters per second unit displayed as either ml sec or mL sec x D Milliliters per minute unit displayed as either ml min or mL min | Software update | 2024-08-20 14:22
Medtronic Inc. | Vanta™ AdaptiveStim™ | 7 | - | Medtronic has received reports from Health Care Professionals HCP and patients indicating that the Model Vanta INS has depleted faster than HCPs or patients expected Review of complaint data and analysis of returned devices confirmed that the Vanta System is working as intended and there were no performance issues However Medtronic found that associated labeling may not be fully utilized to assess the impact of programmed settings on the battery life of the INS and is taking this field corrective action to re iterate warnings and instructions contained in labelling regarding the impact of programming on battery longevity and the use of battery longevity estimator tools to assess the impact of programming on battery life | • Ensure that the Device Eligibility screen in the Model A71300 Clinician Programmer application is used during the trialing evaluation to review compatible INSs that will allow the same parameters used during the trial evaluation. • Ensure that the battery longevity tools in the Model A71200 Clinician Programmer application are used during each patient visit, to keep track of the estimated longevity of the Vanta INS. • Ensure that the patient understands the impact of therapy changes on the battery life of the Vanta INS. • Please complete and return the Customer Acknowledgement Form enclosed in this letter acknowledging that you have received this information. | 2024-08-20 18:50
bioMérieux | FILMARRAY TORCH MODULE BOX | 7 | - | amplification leading to erroneous Tm detections Additionally in very rare cases the x D potential for False Negative Norovirus results on the GI Panel customer complaint out of x D million GI pouches sold and False Negative MRSA results on the BCID Panel x D customer complaint out of million BCID pouches sold due to Not Detected shifted x D Tm results Although this issue does not impact performance estimates due to the extremely x D low occurrence rate it has the potential to cause rare sporadic false results and therefore x D clinicians should always interpret test results in the context of other clinical findings and other x D laboratory test results BIOFIRE tests are intended to be an aid in diagnosis and results x D inconsistent with clinical evaluation should be further investigated in conjunction with other x D findings x D bioM rieux has identified serial numbers of BIOFIRE instruments that had temperature x D calibrations performed with the affected equipment between May and May These x D instruments will need to be returned to a BIOFIRE Service Center to perform an updated x D temperature calibration As the risk of incorrect results is extremely low you may still use x D your BIOFIRE instruments in accordance with product labeling while waiting for your x D scheduled temperature calibration For panel specific guidance please see the Required x D Actions section below x D bioM rieux has created a resource to help you identify impacted instruments by inputting the x D serial number of each of your instruments at the following link x D http biofiredx com fsca x D A comprehensive list of impacted serial numbers has also been provided in Attachment B In x D addition you will be contacted by bioM rieux soon with a list of the specific BIOFIRE x D instruments that are impacted by this correction at your account | Required actions: In this context, we request you to take the following actions: Please utilize Attachment B or http://biofiredx.com/fsca5788-1/ to determine which of your instruments are impacted. If your instrument is not on the impacted list, no action is required. You can continue to use your instrument in accordance with product labeling. If your instrument is on the impacted list, it needs to be returned to a BIOFIRE Service Center for recalibration; your local bioMérieux contact will reach out to you to coordinate the return of your impacted BIOFIRE instrument(s). In the meantime: • Continue to use all your BIOFIRE instruments in accordance with product labeling (instructions for use) while waiting for service. • For BIOFIRE GI Panel Users, you may continue using the product as directed in the package insert, but if a positive Salmonella or Giardia lamblia result is inconsistent with clinical presentation, consider confirming the positive using another method or rerun the sample on an unaffected instrument. In patients suspected of Norovirus infection, please 515 Colorow Drive, Salt Lake City, Utah 84108, USA +1-800-682-2666 local - www.biomerieux.com Field Safety Notice DX-REG-142810 consider confirming Norovirus results via another method prior to reporting those results or rerun the sample on an unaffected instrument. • For BCID2 users, you may continue using the product as directed in the package insert; however, in patients with a Staphylococcus aureus detected result on the BCID2 panel AND suspected of a MRSA infection, please confirm all negative MRSA results via another method prior to reporting or rerun the sample on an unaffected instrument. • For RP2.1, RP2.1plus, ME, PN, PNplus and JI users, you may continue using the product as directed in the package insert; No confirmation of results are required. • Complete all fields of the accompanying Acknowledgment of Receipt Form and return it to bioMérieux so that bioMérieux may acknowledge your receipt of this notification. • If you have further distributed the impacted BIOFIRE instrument(s), please identify any recipients and notify them at once by sending them a copy of this notice and acknowledgement form. | 2024-08-22 14:27
Philips | Allura Xper FD10 Allura Xper FD10/10 Allura Xper FD20/15 Allura Xper FD20 Allura Xper FD10 Allura Xper FD20 Allura Xper FD10/10 Allura Xper FD10/10 Allura Xper FD20 OR Table Allura Xper FD20 Biplane Allura Xper FD20 Allura Xper FD20 Biplane Allura Xper FD10 Allura Xper FD20/10 | 163 | - | Philips has identified that due to the forces applied when moving or turning the monitor of the system the bolts and plastic parts of the Cable Hose Carriers that hold the cable hose to the monitor ceiling suspension may become loose and or break see Figure If this issue occurs plastic parts and or metal bolts of the Cable Hose Carrier s may fall and or part of the cable hose may drop following the detachment of a Cable Hose Carrier | Philips is redesigning the Cable Hose Carriers. Once the redesigned cable hose carriers are available, Philips will contact you to schedule a visit to install them in the affected system(s) (reference FCO72200536, FCO72200576, and FCO72200577). As of the date of this urgent Field Safety Notice letter, Philips expects the solution to be available by Q3 2025. | 2024-08-23 14:25
Philips | Azurion 7 M20 Azurion 7 B20 Azurion 7 M20 Azurion 7 M12 Azurion 7 B20 Azurion 7 B12 Azurion 7 B12 Azurion 7 M12 Azurion 5 M20 | 163 | - | Philips has identified that due to the forces applied when moving or turning the monitor of the system the bolts and plastic parts of the Cable Hose Carriers that hold the cable hose to the monitor ceiling suspension may become loose and or break see Figure If this issue occurs plastic parts and or metal bolts of the Cable Hose Carrier s may fall and or part of the cable hose may drop following the detachment of a Cable Hose Carrier | Philips is redesigning the Cable Hose Carriers. Once the redesigned cable hose carriers are available, Philips will contact you to schedule a visit to install them in the affected system(s) (reference FCO72200536, FCO72200576, and FCO72200577). As of the date of this urgent Field Safety Notice letter, Philips expects the solution to be available by Q3 2025. | 2024-08-23 14:25
Philips | INTEGRIS Allura 9 (biplane) INTEGRIS Allura 15-12 (biplane) | 163 | - | Philips has identified that due to the forces applied when moving or turning the monitor of the system the bolts and plastic parts of the Cable Hose Carriers that hold the cable hose to the monitor ceiling suspension may become loose and or break see Figure If this issue occurs plastic parts and or metal bolts of the Cable Hose Carrier s may fall and or part of the cable hose may drop following the detachment of a Cable Hose Carrier | Philips is redesigning the Cable Hose Carriers. Once the redesigned cable hose carriers are available, Philips will contact you to schedule a visit to install them in the affected system(s) (reference FCO72200536, FCO72200576, and FCO72200577). As of the date of this urgent Field Safety Notice letter, Philips expects the solution to be available by Q3 2025. | 2024-08-23 14:25
Smith Medical | Infusion pump | 2220 | - | CADD Medication Cassette Reservoirs may exhibit a weakened weld joint between the medication bag and tubing due to a production equipment malfunction | product removal | 2024-08-26 13:58
Smith Medical | infusion pump | 384 | - | battery packs may have caused a short to a capacitor within the battery packs While the battery encasement is designed to be flame retardant a short to the capacitor may potentially lead to melting of the battery pack case | Customer should submit a complaint to Manufacture as per the FSN | 2024-08-26 16:19
CARL REINER GmbH | TwinStream | 1 | - | There is a possibility that the adapter does not sufficiently transmit the breathing gas to the ventilator for pressure measurement | The affected adapter will be replaced. | 2024-08-28 12:39
Dominion Biologicals Ltd. | NOVACLONE™ Anti-D IgM + IgG Monoclonal Blend | 70 | - | Some of IFU contain information for other product Misprinted IFU | No impact on the results in anticipated as the test methods descrition on the misprinted IFU are appropriate for NOVACLONE Anti-D IgM + IgG Monoclonal blend. Please discard the printed IFU that was packaged with the product. Please refer to “Customer Actions to be taken: in the attachment. | 2024-08-28 13:46
Allurion Technologies Inc. | Allurion Gastric Balloon | 2294 | - | IFU will be shared next week | IFU will be shared next week. | 2024-08-28 15:07
Abbott Ireland Diagnostics Division | Alinity i HBsAg Controls | 19 | - | Please find attached FSN | Please find attached FSN | 2024-08-28 15:42
Abbott Ireland Diagnostics Division | ARCHITECT HBsAg Controls | 19 | - | Please find attached FSN | Please find attached FSN | 2024-08-28 15:42
Welch Allyn, Inc. | Connex Spot Monitor | - | - | Baxter Healthcare Corporation is communicating important safety information to notify x D customers using the Welch Allyn Configuration Tool with Welch Allyn Connex products of x D the potential cybersecurity risks of using web browsers with password manager x D functionality The Configuration Tool is a web based portal that can be used to customize x D the configuration settings on your Welch Allyn Connex patient monitoring device x D Browsers with password manager functionality will auto populate saved login credentials x D in various fields within the configuration file that will be applied to the Welch Allyn Connex x D Spot Monitor Welch Allyn Connex Vital Signs Monitor and the Welch Allyn Connex x D Integrated Wall System which could allow non approved users access to the configuration x D file x D To address this vulnerability Baxter released an updated version of the Welch Allyn x D Configuration Tool x D Baxter is not aware of any exploitation of this vulnerability or of any personal or health x D information data being compromised as a result | Credentials that were input into Welch Allyn Configuration Tool files have the potential to be compromised and should be changed immediately. 2. Customers will have to reset their login credentials on the updated Configuration Tool. 3. Complete the enclosed Baxter Customer Reply Form and return it to Baxter by scanning and e-mailing it to Bandar_Alosaimi@Baxter.com Returning the customer reply form promptly will confirm your receipt of this notification and prevent you from receiving repeat notices. 4. If you distributed this product to other facilities or departments within your institution, please forward a copy of this communication to them. 5. If you purchased this product from a distributor, please note that the Baxter reply form is not applicable. If a response is requested by your distributor or wholesaler, please return it to your distributor/wholesaler according to their instructions. 6. If you are a dealer, wholesaler, distributor/reseller, or original equipment manufacturer (OEM) that distributed any affected product to other facilities, please notify your customers of this safety information in accordance with your customary procedures. | 2024-08-29 14:00
smith medical | infusion pump | 384 | - | The AC Adapter input plug can become damaged or broken If the input plug is damaged the metal contacts to the body of the AC Adapter may be exposed or one or more of the AC mains prongs may separate from the input plug | Locate all AC Adapters in your possession. Inspect the AC Adapter for damage to the input plug, missing AC prongs, or exposed metal contacts If damaged, remove the AC Adapter from service and submit a request for credit to Smiths Medical | 2024-08-29 15:02
Alcon Laboratories Inc. | TOTALPLUS | 673 | - | Alcon s investigation has determined that there was a process shift on one electropolish machine which resulted in the potential for a portion of knives trocar entry system within these lots to not meet sharpness expectations | • If you received sterile standalone knives or trocar entry system, we are asking that you dispose of the affected ophthalmic knife/trocar entry system and use a replacement. Alcon is sending sufficient replacement quantities to cover your estimated remaining impacted inventory. • If you received a Vitrectomy/Combined Procedure Pak ® with affected trocar entry system lot numbers, we are asking that, upon opening your Vitrectomy/Combined Procedure Pak for surgical use, you remove and dispose of the affected trocar entry system and use a replacement trocar entry system. Alcon is sending sufficient replacement quantities to cover your estimated remaining impacted inventory. | 2024-08-30 20:54
Medtronic Navigation, Inc. | Software App 9735762 Stealth S8 Application | 18 | - | Two software anomalies were identified while conducting internal testing Both anomalies were determined to only affect StealthStation S software application versions and x D Medtronic engineers discovered that no pop up message appears when patient exams have slice spacing greater than mm This issue only occurs during CT Fluoro Spine procedures x D Medtronic engineers discovered that when radiographic images are taken via the lateral orientation setting on the Ziehm FD system and transferred to the StealthStation S for navigation the navigation on the images is displayed incorrectly The Ziehm FD C arm is a third party imaging system used with the StealthStation S The StealthStation S software is responsible for converting the Ziehm FD C arm radiological scans into the navigation views that a surgeon utilizes on the StealthStation S during a procedure The navigation appears with a degree rotation around the superior inferior S I axis with the left right L R axis and anterior posterior A P axis appearing as mirror images The exam image orientation and related indication labels on the exams remain correct This issue only occurs with use of the Ziehm FD C arm system s The technical cause of this anomaly was determined to be due to the method used to calculate the calibration transform x D | • Please review this notification, including the Medtronic Recommendations detailed in the Issue Description section above with all physician users. If you have any questions related to this issue, please contact your Medtronic Representative • Please complete and return the Customer Acknowledgement Form enclosed with this letter acknowledging receipt of this information. o Medtronic will work expeditiously with your facility to coordinate a software update to version 2.1 at no charge to resolve these anomalies. • This notice needs to be passed on to those who need to be aware within your organization or to any organization where the potentially affected devices have been transferred. Please maintain a copy of this notice in your records. | 2024-08-31 18:22
Philips Medical Systems Nederland B.V. | Ingenia 1.5T | 4 | - | Philips has identified an issue in the MR System Patient Support where the Integrated Radio Frequency IRF Carrier assembly screws may come loose and protrude beyond the surface Figure This could cause the screws to interfere with other moving parts potentially blocking the horizontal tabletop movement x D The IRF Carrier Figure Item and a component of the Patient Support within the MR system supports the tabletop where the patient lies As the patient moves into the MRI tunnel the tabletop supported by the IRF Carrier forms a container for positioning the posterior coil x D There have been no reports of adverse events reported to Philips regarding the issue as of March | Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to inspect the IRF carrier assembly and provide a correction if necessary (reference FCO78100588). Philips plans to start implementing corrections in Q4 2024. | 2024-09-02 08:59
Philips Medical Systems Nederland B.V. | Ingenia 3.0T | 4 | - | Philips has identified an issue in the MR System Patient Support where the Integrated Radio Frequency IRF Carrier assembly screws may come loose and protrude beyond the surface Figure This could cause the screws to interfere with other moving parts potentially blocking the horizontal tabletop movement x D The IRF Carrier Figure Item and a component of the Patient Support within the MR system supports the tabletop where the patient lies As the patient moves into the MRI tunnel the tabletop supported by the IRF Carrier forms a container for positioning the posterior coil x D There have been no reports of adverse events reported to Philips regarding the issue as of March | Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to inspect the IRF carrier assembly and provide a correction if necessary (reference FCO78100588). Philips plans to start implementing corrections in Q4 2024. | 2024-09-02 08:59
Philips Medical Systems Nederland B.V. | Ingenia 3.0T CX | 4 | - | Philips has identified an issue in the MR System Patient Support where the Integrated Radio Frequency IRF Carrier assembly screws may come loose and protrude beyond the surface Figure This could cause the screws to interfere with other moving parts potentially blocking the horizontal tabletop movement x D The IRF Carrier Figure Item and a component of the Patient Support within the MR system supports the tabletop where the patient lies As the patient moves into the MRI tunnel the tabletop supported by the IRF Carrier forms a container for positioning the posterior coil x D There have been no reports of adverse events reported to Philips regarding the issue as of March | Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to inspect the IRF carrier assembly and provide a correction if necessary (reference FCO78100588). Philips plans to start implementing corrections in Q4 2024. | 2024-09-02 08:59
Philips Medical Systems Nederland B.V. | INGENIA AMBITION S | 4 | - | Philips has identified an issue in the MR System Patient Support where the Integrated Radio Frequency IRF Carrier assembly screws may come loose and protrude beyond the surface Figure This could cause the screws to interfere with other moving parts potentially blocking the horizontal tabletop movement x D The IRF Carrier Figure Item and a component of the Patient Support within the MR system supports the tabletop where the patient lies As the patient moves into the MRI tunnel the tabletop supported by the IRF Carrier forms a container for positioning the posterior coil x D There have been no reports of adverse events reported to Philips regarding the issue as of March | Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to inspect the IRF carrier assembly and provide a correction if necessary (reference FCO78100588). Philips plans to start implementing corrections in Q4 2024. | 2024-09-02 08:59
Philips Medical Systems Nederland B.V. | INGENIA AMBITION X | 4 | - | Philips has identified an issue in the MR System Patient Support where the Integrated Radio Frequency IRF Carrier assembly screws may come loose and protrude beyond the surface Figure This could cause the screws to interfere with other moving parts potentially blocking the horizontal tabletop movement x D The IRF Carrier Figure Item and a component of the Patient Support within the MR system supports the tabletop where the patient lies As the patient moves into the MRI tunnel the tabletop supported by the IRF Carrier forms a container for positioning the posterior coil x D There have been no reports of adverse events reported to Philips regarding the issue as of March | Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to inspect the IRF carrier assembly and provide a correction if necessary (reference FCO78100588). Philips plans to start implementing corrections in Q4 2024. | 2024-09-02 08:59
Philips Medical Systems Nederland B.V. | INGENIA ELITION X | 4 | - | Philips has identified an issue in the MR System Patient Support where the Integrated Radio Frequency IRF Carrier assembly screws may come loose and protrude beyond the surface Figure This could cause the screws to interfere with other moving parts potentially blocking the horizontal tabletop movement x D The IRF Carrier Figure Item and a component of the Patient Support within the MR system supports the tabletop where the patient lies As the patient moves into the MRI tunnel the tabletop supported by the IRF Carrier forms a container for positioning the posterior coil x D There have been no reports of adverse events reported to Philips regarding the issue as of March | Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to inspect the IRF carrier assembly and provide a correction if necessary (reference FCO78100588). Philips plans to start implementing corrections in Q4 2024. | 2024-09-02 08:59
Philips Medical Systems Nederland B.V. | INGENIA ELITION S | 4 | - | Philips has identified an issue in the MR System Patient Support where the Integrated Radio Frequency IRF Carrier assembly screws may come loose and protrude beyond the surface Figure This could cause the screws to interfere with other moving parts potentially blocking the horizontal tabletop movement x D The IRF Carrier Figure Item and a component of the Patient Support within the MR system supports the tabletop where the patient lies As the patient moves into the MRI tunnel the tabletop supported by the IRF Carrier forms a container for positioning the posterior coil x D There have been no reports of adverse events reported to Philips regarding the issue as of March | Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to inspect the IRF carrier assembly and provide a correction if necessary (reference FCO78100588). Philips plans to start implementing corrections in Q4 2024. | 2024-09-02 08:59
Philips Medical Systems Nederland B.V. | MR 5300 | 4 | - | Philips has identified an issue in the MR System Patient Support where the Integrated Radio Frequency IRF Carrier assembly screws may come loose and protrude beyond the surface Figure This could cause the screws to interfere with other moving parts potentially blocking the horizontal tabletop movement x D The IRF Carrier Figure Item and a component of the Patient Support within the MR system supports the tabletop where the patient lies As the patient moves into the MRI tunnel the tabletop supported by the IRF Carrier forms a container for positioning the posterior coil x D There have been no reports of adverse events reported to Philips regarding the issue as of March | Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to inspect the IRF carrier assembly and provide a correction if necessary (reference FCO78100588). Philips plans to start implementing corrections in Q4 2024. | 2024-09-02 08:59
Philips Medical Systems Nederland B.V. | MR 7700 | 4 | - | Philips has identified an issue in the MR System Patient Support where the Integrated Radio Frequency IRF Carrier assembly screws may come loose and protrude beyond the surface Figure This could cause the screws to interfere with other moving parts potentially blocking the horizontal tabletop movement x D The IRF Carrier Figure Item and a component of the Patient Support within the MR system supports the tabletop where the patient lies As the patient moves into the MRI tunnel the tabletop supported by the IRF Carrier forms a container for positioning the posterior coil x D There have been no reports of adverse events reported to Philips regarding the issue as of March | Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to inspect the IRF carrier assembly and provide a correction if necessary (reference FCO78100588). Philips plans to start implementing corrections in Q4 2024. | 2024-09-02 08:59
Philips | Evolution Upgrade 3.0T,SmartPath to dStream for 1.5T,SmartPath to dStream for XR and 3.0T,SmartPath to Ingenia Elition X | 4 | - | Philips has identified an issue in the MR System Patient Support where the Integrated Radio Frequency IRF Carrier assembly screws may come loose and protrude beyond the surface Figure This could cause the screws to interfere with other moving parts potentially blocking the horizontal tabletop movement x D The IRF Carrier Figure Item and a component of the Patient Support within the MR system supports the tabletop where the patient lies As the patient moves into the MRI tunnel the tabletop supported by the IRF Carrier forms a container for positioning the posterior coil x D There have been no reports of adverse events reported to Philips regarding the issue as of March | Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to inspect the IRF carrier assembly and provide a correction if necessary (reference FCO78100588). Philips plans to start implementing corrections in Q4 2024. | 2024-09-02 08:59
Philips | IntelliVue Patient Monitor MX400 IntelliVue Patient Monitor MX430 IntelliVue Patient Monitor MX450 IntelliVue Patient Monitor MX500 IntelliVue Patient Monitor MX550 | 6 | - | Electrical grounding is crucial to ensure that hospital patient monitoring equipment is safe reliable and in compliance with industry standards Grounding provides a path for electrical current to return to the ground during faults thereby minimizing risk of electric shock and protecting equipment from damage It also controls leakage currents and dissipates static electricity that could pass through a patient s body or impair device function Additionally grounding helps to reduce electromagnetic interference EMI ensuring accurate readings of patient vitals and reduction of interference with other electrical equipment in the healthcare environment During a production process Philips became aware of one IntelliVue power supply with a broken ground bolt upon disconnection of the ground cable from the equipotential ground connector Image depicted below | A Philips representative will contact you to schedule a visit from a Philips Field Service Engineer who will replace the power supply. | 2024-09-02 09:27
GS Elektromedizinische Geräte G. Stemple GmbH | Arm Corpuls CPR | 39 | - | Sporadically the specified life span of the battery corpuls cpr LiPo REF is not x D reached x D Cause for this is the cell voltage which falls below a certain voltage threshold due to deep x D discharging or longer storage periods if the battery is not charged regularly as prescribed in x D the user manual This promotes the progressive degradation of the electrode materials This x D results in the battery going into error statuses and the life span is shortened | A new firmware 1.4-006 software will be installed & updated on battery corpuls CPR | 2024-09-02 10:22
Baxter Healthcare Corporation | EXACTAMIX Inlet | 80 | - | Not identified yet | An expalantory letter directed to the end users is attached to guide End users on how to contain and detect this issue as a Correction/ Preventive action, no further actions are planned at this stage. | 2024-09-03 22:51
Siemens Healthcare Diagnostics Inc. | Atellica IM CA 19-9 | 7600 | - | Issue Description x D Siemens Healthineers has confirmed through the investigation of customer complaints a positive bias for some samples around a concentration of U mL with the lots listed in the table below as compared to the previous lot The sample bias observed was greatest in the Asia Pacific sample population The observed bias is not proportional across the assay measuring interval and is less pronounced for sample results above a concentration of U mL x D This issue is isolated to the product listed in the table below x D See Appendix for data in Figures x D Kindly review the attached UFSN | Customer Actions • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Perform the instructions provided below: o Discontinue use of and discard the kit lots listed in the table above (Products Section). o If you experience this issue, you may request no-charge replacement product from your local Siemens Healthineers or distributor office. Please review your inventory of these products to determine your laboratory’s replacement needs and to provide information to Siemens Healthineers for reporting to the authorities. o Complete and return the Field Correction Effectiveness Check and indicate product replacement needs on the form attached to this letter within 30 days. o Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached UFSN | 2024-09-09 07:18
.Siemens Healthcare Diagnostics Inc | Atellica IM CA 19-9 | 7600 | - | Issue Description x D Siemens Healthineers has confirmed through the investigation of customer complaints a positive bias for some samples around a concentration of U mL with the lots listed in the table below as compared to the previous lot The sample bias observed was greatest in the Asia Pacific sample population The observed bias is not proportional across the assay measuring interval and is less pronounced for sample results above a concentration of U mL x D This issue is isolated to the product listed in the table below x D See Appendix for data in Figures x D Kindly review the attached UFSN | Customer Actions • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Perform the instructions provided below: o Discontinue use of and discard the kit lots listed in the table above (Products Section). o If you experience this issue, you may request no-charge replacement product from your local Siemens Healthineers or distributor office. Please review your inventory of these products to determine your laboratory’s replacement needs and to provide information to Siemens Healthineers for reporting to the authorities. o Complete and return the Field Correction Effectiveness Check and indicate product replacement needs on the form attached to this letter within 30 days. o Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached UFSN | 2024-09-09 07:18
Siemens Healthcare Diagnostics Inc. | CA 19-9 | 7600 | - | Issue Description x D Siemens Healthineers has confirmed through the investigation of customer complaints a positive bias for some samples around a concentration of U mL with the lots listed in the table below as compared to the previous lot The sample bias observed was greatest in the Asia Pacific sample population The observed bias is not proportional across the assay measuring interval and is less pronounced for sample results above a concentration of U mL x D This issue is isolated to the product listed in the table below x D See Appendix for data in Figures x D Kindly review the attached UFSN | Customer Actions • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Perform the instructions provided below: o Discontinue use of and discard the kit lots listed in the table above (Products Section). o If you experience this issue, you may request no-charge replacement product from your local Siemens Healthineers or distributor office. Please review your inventory of these products to determine your laboratory’s replacement needs and to provide information to Siemens Healthineers for reporting to the authorities. o Complete and return the Field Correction Effectiveness Check and indicate product replacement needs on the form attached to this letter within 30 days. o Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached UFSN | 2024-09-09 07:18
Siemens Healthcare Diagnostics Inc. | Atellica CH μALB_2 | 312480 | - | Issue Description x D Siemens Healthineers has confirmed through an investigation the Atellica CH Microalbumin ALB lots listed in the table below are not meeting the High Dose Hook Effect claim as stated in the Instructions for Use IFU on the Atellica CH and Atellica CI Analyzers x D The Atellica CH ALB Measuring Interval is mg dL mg L The IFU states that High microalbumin levels can cause a paradoxical decrease in signal as a result of the high dose hook effect In the Atellica CH ALB assay microalbumin levels as high as mg dL mg L will read mg dL mg L x D For the lots in the table below Products Section the high dose hook effect claim begins to fail at concentrations greater than mg dL mg L x D Kindly review the attached UFSN | Customer Actions • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Customers can continue to use the impacted μALB_2 lots in the table above (Products Section), with the understanding that patient samples with values above 9,500 mg/dL (95,000 mg/L) can result in falsely depressed results. • Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. • Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached UFSN | 2024-09-09 07:36
Siemens Healthcare Diagnostics Inc. | Atellica IM aTPO | 7570 | - | Issue Description x D Siemens Healthineers has confirmed customer complaints of a positive bias in patient results x D at and below the cut off of U mL with the lots listed in the table below when compared to x D alternate lots on the Atellica IM analyzer Atellica CI does not exhibit the same positive bias as x D observed with the Atellica IM analyzer x D The observed bias is not proportional across the assay measuring interval and becomes less x D significant as sample concentration increases The Expected Values cut off of U mL has x D been verified with euthyroid patient samples on the Atellica CI and Atellica IM analyzers x D Please note this issue is limited to only the kit lots ending in as listed in the table below x D Kindly review the attached UFSN | Customer Actions • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Perform the actions provided below: o Discontinue use of and discard the kit lots listed in the table above (Products Section). o You may request no-charge replacement product from your local Siemens Healthineers or distributor office. Please review your inventory of these products to determine your laboratory’s replacement needs and to provide information to Siemens Healthineers for reporting to the authorities. o Complete and return the Field Correction Effectiveness Check and indicate product replacement needs on the form attached to this letter within 30 days. o Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached UFSN | 2024-09-09 09:41
Siemens Healthcare Diagnostics Inc. | Atellica IM aTPO +O | 7570 | - | Issue Description x D Siemens Healthineers has confirmed customer complaints of a positive bias in patient results x D at and below the cut off of U mL with the lots listed in the table below when compared to x D alternate lots on the Atellica IM analyzer Atellica CI does not exhibit the same positive bias as x D observed with the Atellica IM analyzer x D The observed bias is not proportional across the assay measuring interval and becomes less x D significant as sample concentration increases The Expected Values cut off of U mL has x D been verified with euthyroid patient samples on the Atellica CI and Atellica IM analyzers x D Please note this issue is limited to only the kit lots ending in as listed in the table below x D Kindly review the attached UFSN | Customer Actions • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Perform the actions provided below: o Discontinue use of and discard the kit lots listed in the table above (Products Section). o You may request no-charge replacement product from your local Siemens Healthineers or distributor office. Please review your inventory of these products to determine your laboratory’s replacement needs and to provide information to Siemens Healthineers for reporting to the authorities. o Complete and return the Field Correction Effectiveness Check and indicate product replacement needs on the form attached to this letter within 30 days. o Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached UFSN | 2024-09-09 09:41
SLE, LTD. | SLE6000 Infant Ventilator | 7 | - | After investigation it was identified that the root cause of the reported failures was an EEPROM component on the Monitor Printed Circuit Board PCB of the SLE This specific component generated Sub ambient pressure and Out of calibration alarms at start up not related to the status of the system | If customers identify affected devices in the Customer Response Form, Device Technologies will contact customers to arrange the replacement of affected Monitor PCBs. NB: Customers may continue to use the impacted devices until replacement PCB have been received and installed. | 2024-09-09 17:45
AGFA HealthCare | Enterprise Imaging XERO Viewer | 24 | - | Incorrect handling of non square Image pixel spacing in Enterprise Imaging XERO Viewer at x D initial load leads to incorrect placement of reference lines | AGFA HealthCare will provide the software corrections for the affected versions. A patch is available for each of the following XERO Viewer version streams: | 2024-09-13 15:52
B. Braun Melsungen AG | Spaceplus Perfusor® | 3 | - | Please find attached FSN letter | Please find attached FSN letter | 2024-09-14 16:01
Cepheid | Xpert Xpress CoV-2 plus | 100 | - | Cepheid has received reports that some of our customers have experienced a higher number of x D E Probe Check Too Low errors than would typically be expected when using these tests and x D their intended use specimen types E yields a non determinate result Error | If you experience an E5007 error when using a cartridge, repeat testing on another cartridge. If another E5007 error occurs, or if you have been experiencing intermittent E5007 errors on tests from these lots, Cepheid will replace any remaining tests that you currently have in inventory from lots. If you are experiencing issues and you would like to request replacement product, you should document (see Response Form), stop using, and then dispose of any remaining tests from the applicable lots that are in your inventory. | 2024-09-14 19:31
GE Healthcare | Centricity Universal Viewer Zero Footprint | 55 | - | GE HealthCare has become aware of an issue in Centricity Universal Viewer Zero Footprint Client ZFP version v SP x where the latest report is not shown by default to the user The issue may occur when a study has multiple DICOM structured reports i e preliminary final and or addendum stored in different series x D Users continue to have access to all reports in the Series Selector However since the latest report may not be shown by default in the ZFP Viewer this could lead to potential misdiagnosis if a user relies only on the information in the original report and does not review the latest report x D | You can continue to use your device by following the recommendation below. To find the latest report, review all the reports available for the study in the Series Selector. Navigation to the series selection can be found in the User Manual under Section 2.4.3 of The Series Selector, with the subheading of “Open a report/document.” Please ensure all potential users in your facility are made aware of this safety notification and the recommended action. Please retain this document for your records. GE HealthCare will correct all affected products at no cost to you. A GE HealthCare representative will contact you to arrange for the correction. After the software has been corrected, be sure to destroy the installation media for affected software at your site. | 2024-09-15 11:44
Maquet GmbH | MEERA CL | 8 | - | it is possible that the error code occurs on the IR Hand Control when a MEERA table is controlled and the table stops the movement This error may occur sporadically when the MEERA table is controlled via an IR Hand Control With a cable hand Control the error does not occur If the error occurs the movement is not permanently blocked The MEERA table turns off after the key is released and restarts when the key is pressed again the table restarts The MEERA table works normally again until the error occurs again x D We are not aware of instances where a patient was injured due to one of the described issues x D | A solution that will correct this issue has been developed. Getinge will initiate an immediate field action of all affected device units. You will be contacted by your Getinge sales or service representative to plan for the update of your device The device can be used in accordance with the instructions for use. | 2024-09-15 12:12
Ohmeda Medical, a Division of Datex-Ohmeda, Inc. | Giraffe OmniBed Carestation CS1 | 1161 | - | GE HealthCare has become aware that for certain Giraffe OmniBed and Giraffe OmniBed Carestation devices the screw that secures the doors that cover the warmer heaters may not have been torqued to specification This may lead to the doors becoming loose If the doors become loose the system activates a high priority alarm and the canopy movement will stop If this situation is encountered the user manual instructs that the canopy should not be moved and the system should not be put into clinical use until service is performed The product user manual requires pre use checkout steps to be followed that includes checking for proper functioning of the heater doors If these instructions are followed no patient injury would be expected However if these instructions are not followed continued attempts to force the canopy to move may damage the canopy and in rare cases it may result in the heater door to fall potentially resulting in patient injury x D x D No injuries have been reported as a result of this potential issue x D | You may continue to use your devices. Please always pay attention to product alarms and follow the instructions in the product user manual. Download the Service Manual Addendum (5971733) from the GE HealthCare Customer Documentation Portal (see impacted device list below). https://www.gehealthcare.com/support/manuals Inspect the heater door mounting screw on your device per Figure 1 in the attahced letter. If the screw is not appropriately tightened (as shown in Figure 1 as “incorrect assembly”), complete steps 9 and 10 of Section 1.1 of the Service Manual Addendum (5971733). If you have any questions or need assistance in following the instructions, please contact your GE HealthCare Service Representative. Please ensure all potential users in your facility are made aware of this safety notification and the recommended actions. Retain this document for your records. | 2024-09-15 14:39
Olympus | SOLTIVE™ Premium SuperPulsed Laser System | 9 | - | Olympus received a complaint regarding the Soltive Laser System preset laser settings The term Bladder Stone was incorrectly translated in both Spanish and Portuguese to Kidney Stone C lculo renal on the systems Graphical User Interface GUI As a result of the incorrect Spanish and Portuguese translation from Bladder Stone to Kidney Stone on the GUI there is a potential to deliver an incorrect or unintended amount of energy to the patient s anatomy This issue is specific to the preset laser settings and does not affect users who are not utilizing the preset option Olympus has not received any additional complaints related to this issue | Acknowledge of the FSN | 2024-09-15 16:12
Randox Laboratories | mALB | 4 | - | Randox Laboratories can confirm there has been a transcription error in the Instructions for Use IFU for the Microalbumin calibrator series MA MA lots MA MA packed into batches The mg dl values are incorrect by a factor of for the two lots all other values are correct | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2024-09-16 09:04
Randox Laboratories | mALB CAL | 4 | - | Randox Laboratories can confirm there has been a transcription error in the Instructions for Use IFU for the Microalbumin calibrator series MA MA lots MA MA packed into batches The mg dl values are incorrect by a factor of for the two lots all other values are correct | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2024-09-16 09:04
Shimadzu Corporation | X RAY R/F SYSTEM FLUOROspeed | 2 | - | PLEASE FIND ATTACHED FSN LETTER | PLEASE FIND ATTACHED FSN LETTER | 2024-09-16 11:17
Abbott Molecular Division Inc. | Alinity m SYSTEM | 1 | - | Please find attached FSN letter | Please find attached FSN letter | 2024-09-16 11:57
SAHHA FACTORY COMPANY FOR SOLUTION MEDICAL PREPARATIONS | SACO- POVIDONE IODINE | 75000 | - | leakage from the packaging due to a defect in the packaging Material | Recall | 2024-09-16 13:59
Cepheid | Xpert® Xpress CoV-2 plus | 100 | - | As with all Xpert tests no patient results are reported when the test yields a non determinate x D result Error and the user is instructed to retest on a new cartridge Customers who conduct a x D retest will experience a delay in final test results | If you experience an E5007 error when using a cartridge, repeat testing on another cartridge. If another E5007 error occurs, or if you have been experiencing intermittent E5007 errors on tests from these lots, Cepheid will replace any remaining tests that you currently have in inventory from lots. If you are experiencing issues and you would like to request replacement product, you should document (see Response Form), stop using, and then dispose of any remaining tests from the applicable lots that are in your inventory. | 2024-09-16 15:12
Olympus | Gastrointestinal Videoscope Olympus EVIS X1™ ,Gastrointestinal Videoscope Olympus EVIS X1, Bronchovideoscopes | 11 | - | Olympus is writing to inform you of a Medical Device Correction pertaining to the abovelist of scopes and serial numbers of the Olympus EVIS EXERA III GIF H N Gastrointestinal Videoscope Olympus EVIS X GIF Gastrointestinal Videoscope Olympus EVIS X BF H and BF TH Bronchovideoscopes and Olympus EVIS EXERA III SIF H Small Intestinal Videoscope These videoscopes work in conjunction with the Olympus video system light sources monitors etc for visualization during endoscopy and endoscopic surgery | 1. Carefully read the content of this notification. 2. Examine your inventory and identify the above-listed device with the affected serial number. 3. Olympus representative will reach out to you to arrange a mutually convenient time for the return of your device to an Olympus Repair Center to receive a color adjustment. 4. If you have further distributed this product, identify your customers, and forward them this notification. | 2024-09-18 17:29
X-Nav Technologies, LLC | X-Guide Handpiece Adaptor Sleeve | - | - | The Handpiece Adaptor may have a manufacturing defect causing incorrect geometry The incorrect geometry in turn causes the Handpiece Adaptor to not fit onto the doctor s dental handpiece | Identify if you possess any of the affected devices. The lot numbers of the affected devices are the following: 2308010012 2306020012 2310010012 If the lot number on your device is not one of the above, then that Adaptor Sleeve is not part of this Recall. If you have a defective device, quarantine the device so that it is prevented from being used. Finally, destroy the device or return to X-Nav Technologies, LLC. Return the response form. | 2024-09-19 14:29
Saudi Mais Medical Products | I.V CANNULA | 279076 | - | Manufacturing defects lead to particles in I V CANNULA | The product I.V CANNULA with reference 072-2419 having this Lot 240124 delivered shall be returned to the manufacturer | 2024-09-23 16:22
Respironics, Inc. | Trilogy Evo International | 297 | - | Nebulized aerosols that accumulate over time have the potential to permanently impact the internal flow sensor Any Trilogy Evo Trilogy Evo O and Trilogy EV devices that have historically been used with an in line nebulizer in certain configurations may be impacted If your device has never been used with an in line nebulizer it is not affected by this issue and can continue to be used If using an in line nebulizer continue to use in accordance with the guidance in this notice x D x D Circumstances that may result in aerosol deposition x D When the in line nebulizer is used with passive circuits for tidal volumes greater than or equal to mL or x D When the in line nebulizer is placed at the dry side of the heated humidifier or x D When the in line nebulizer is placed at the inspiratory port to patient device outlet or x D When the in line nebulizer is placed in any location other than those identified in the images in Section x D Those most vulnerable to this issue include ventilator dependent patients infants and pediatric patients who are being ventilated in a volume control mode | At this time, this communication is intended to provide awareness and understanding of the issue and immediate actions to be taken by the customer when using an in-line nebulizer. Philips Respironics understands that use of in-line nebulizers is common amongst ventilator patients and is working diligently to further understand the interaction between in-line nebulizers and the flow sensor within the Trilogy Evo devices. Philips Respironics is continuing to investigate this issue and will follow- up with customers to provide additional guidance and solutions as it becomes available during the next few months. | 2024-09-23 16:39
Respironics, Inc. | Trilogy Evo O2, International | 297 | - | Nebulized aerosols that accumulate over time have the potential to permanently impact the internal flow sensor Any Trilogy Evo Trilogy Evo O and Trilogy EV devices that have historically been used with an in line nebulizer in certain configurations may be impacted If your device has never been used with an in line nebulizer it is not affected by this issue and can continue to be used If using an in line nebulizer continue to use in accordance with the guidance in this notice x D x D Circumstances that may result in aerosol deposition x D When the in line nebulizer is used with passive circuits for tidal volumes greater than or equal to mL or x D When the in line nebulizer is placed at the dry side of the heated humidifier or x D When the in line nebulizer is placed at the inspiratory port to patient device outlet or x D When the in line nebulizer is placed in any location other than those identified in the images in Section x D Those most vulnerable to this issue include ventilator dependent patients infants and pediatric patients who are being ventilated in a volume control mode | At this time, this communication is intended to provide awareness and understanding of the issue and immediate actions to be taken by the customer when using an in-line nebulizer. Philips Respironics understands that use of in-line nebulizers is common amongst ventilator patients and is working diligently to further understand the interaction between in-line nebulizers and the flow sensor within the Trilogy Evo devices. Philips Respironics is continuing to investigate this issue and will follow- up with customers to provide additional guidance and solutions as it becomes available during the next few months. | 2024-09-23 16:39
Respironics, Inc. | Trilogy EV300 INTL | 297 | - | Nebulized aerosols that accumulate over time have the potential to permanently impact the internal flow sensor Any Trilogy Evo Trilogy Evo O and Trilogy EV devices that have historically been used with an in line nebulizer in certain configurations may be impacted If your device has never been used with an in line nebulizer it is not affected by this issue and can continue to be used If using an in line nebulizer continue to use in accordance with the guidance in this notice x D x D Circumstances that may result in aerosol deposition x D When the in line nebulizer is used with passive circuits for tidal volumes greater than or equal to mL or x D When the in line nebulizer is placed at the dry side of the heated humidifier or x D When the in line nebulizer is placed at the inspiratory port to patient device outlet or x D When the in line nebulizer is placed in any location other than those identified in the images in Section x D Those most vulnerable to this issue include ventilator dependent patients infants and pediatric patients who are being ventilated in a volume control mode | At this time, this communication is intended to provide awareness and understanding of the issue and immediate actions to be taken by the customer when using an in-line nebulizer. Philips Respironics understands that use of in-line nebulizers is common amongst ventilator patients and is working diligently to further understand the interaction between in-line nebulizers and the flow sensor within the Trilogy Evo devices. Philips Respironics is continuing to investigate this issue and will follow- up with customers to provide additional guidance and solutions as it becomes available during the next few months. | 2024-09-23 16:39
Covidien LLC | McGrath MAC Video Laryngoscope | 126 | - | Medtronic received customer reports related to the McGRATH V battery SKU installed in a video laryngoscope handle SKU First Generation reporting that the McGRATH V battery exploded Batteries that are depleted below the design threshold may increase the risk of the battery becoming unstable which may lead to a thermal event followed by a risk of explosion x D Medtronic s investigations to date have identified the root cause to these events as an internal short within the battery assembly that leads to a thermal run away which in turn leads to an explosive event There are three identified contributing factors that can work in combination to create an internal short within the battery x D A Mishandling Drop impact or excessive vibration of the battery x D B Battery Management Deep depletion of the battery results in increased movement of internal elements within the battery First generation McGRATH devices left on for minutes may increase the risk of internal short x D C Battery Manufacturing process variation Note No battery defects or non conformances have been identified as part of investigations or Medtronic s on site process controls review x D | • Immediately discontinue use of all 300-000-000 devices. All 300-000-000 devices, irrespective of age, present the same risk to health outlined above. • Uninstall the battery assembly from the devices and dispose of both the device and battery per local procedures. • Pass this notice on to all who need to be aware within your organization or to any organization where the potentially affected devices have been transferred or distributed. • Please complete the Customer Acknowledgement Form acknowledging receipt of this notification | 2024-09-24 08:25
OLYMPUS | FLUSHING PUMP OFP-2 | 16 | - | Olympus identified an intermittent loss of function of the OFP Flushing Pump caused by an internal component connection failure In the event that the connection failure occurs the OFP device will return to or remain in the Off condition and the operator will not be able to use the flushing function Through its investigation Olympus determined that this issue resulted from an alternate tool being used during the manufacturing of one lot of the subject component Olympus has received three complaints that are potentially applicable to this issue No adverse events have been reported x D To address this issue an Olympus representative will perform an on site inspection of your OFP device s If the device fails the inspection it will be repaired by a qualified Field Service Technician on site You may continue to use your OFP pump until an Olympus representative performs the inspection x D | if identified during preparation of the device this can lead to a delay in initiating the procedure, or if noticed during use it can prolong the procedure to either replace the device and/or the user can elect to complete the procedure with alternative methods of endoscopic irrigation available in the procedural room setting. | 2024-09-24 13:54
Randox Laboratories | URNAL CONTROL 2 | 12 | - | Randox Laboratories can confirm that the urinalysis control Level lot UC may show negative results for Leukocytes within the claimed storage and stability period for some Methods If you are experiencing this issue please contact technical services randox com with data to demonstrate the issue and replacements will be arranged No action is required if you are reporting positive results for Leucocytes | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2024-09-25 14:31
Spacelabs Healthcare | Ultraview DM3 Monitor and Spacelabs DM4 | 125 | - | The lithium ion battery used in the DM and DM is susceptible to failure and overheating when used beyond its recommended lifetime or otherwise damaged | The Operations Manual and Service Manual for the DM3 and DM4 have been updated to convey the risks associated with continuing to use the lithium batteries past the specified use life and to clarify that batteries older than two years, damaged, or that are no longer holding a charge must be replaced. | 2024-09-25 16:36
Philips Medical Systems Nederland B.V. | Zenition 50 | 28 | - | Philips has identified an issue with the Zenition and Zenition systems where the Ampere Mains Control Unit MCU board fuse may blow out during the system start up or during an ongoing procedure This issue may occur due to fluctuations in the power supply outside of the specifications mentioned in the Instructions for Use Section Power Supply leading to a fuse blowout and to subsequent loss of power and eventually to the system becoming unavailable | Philips will replace the 5-Ampere MCU board fuse with a 10-Ampere MCU board fuse in all affected systems. Philips will contact customers to schedule a visit to perform this action free of charge (reference FCO71800109). | 2024-09-30 15:04
Philips Medical Systems Nederland B.V. | Zenition 70 | 28 | - | Philips has identified an issue with the Zenition and Zenition systems where the Ampere Mains Control Unit MCU board fuse may blow out during the system start up or during an ongoing procedure This issue may occur due to fluctuations in the power supply outside of the specifications mentioned in the Instructions for Use Section Power Supply leading to a fuse blowout and to subsequent loss of power and eventually to the system becoming unavailable | Philips will replace the 5-Ampere MCU board fuse with a 10-Ampere MCU board fuse in all affected systems. Philips will contact customers to schedule a visit to perform this action free of charge (reference FCO71800109). | 2024-09-30 15:04
Randox Laboratories | HUM ASY CONTROL 2 | 420 | - | Randox Laboratories can confirm there has been a transcription error in the Instructions for Use IFU for the Human Assayed Control Level HN and HS lots UN UN UN and The analyte NEFA in the method section is listed as Immunoturbidimetric and should be Colorimetric The IFUs have been updated and are available on www randox com Please discard the incorrect IFUs and download the updated IFUs Please note the targets and ranges have not been affected | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2024-10-01 10:39
William Cook Europe APS | Cook Celect® Platinum Vena Cava Filter | 1 | - | During deployment of COOK Celect Platinum Vena Cava Filter jugular approach filter twisted | Initial Action: Isolation of the affected products (Lot No. E4307793) Field Safety Corrective Action: Due to be communicated - after investigation | 2024-10-01 16:22
Baxter Medical Systems GmbH + Co. KG | TruSystem 7000 | - | - | the power supply cable which runs along the battery was incorrectly positioned under the battery after replacement | Baxter will implement a design improvement to decrease the likelihood of incorrectly positioning battery and power supply cable during replacement. Baxter will contact customers who recently had their batteries replaced to inspect and confirm accurate placement of the battery and power supply cable. | 2024-10-02 11:34
Waldemar Link GmbH & Co. KG | Universal Handle, with quick coupling, Stainless Steel, straight General Hip Instruments | 1 | - | The purpose of the instrument set is to enable the user to use the associated implant system within the scope of the procedures described in the associated surgical technique Any other use of the instruments is not permitted The instrument set consists of defined combinable instruments All instruments in the instrument set are intended for temporary use This instrument is used in the context of femur preparation to impact rasp stems and bone compressors into the bone and extract them | Identify Device Quarantine Device Return Device Destroy Device On-site device inspection Follow patient management recommendations Take note of amendment / reinforcement of Instructions For Use (IFU) Other - see instructions below in red None If you have any of the affected product in your inventory, please follow instructions: Identify the potentially affected products. Rasp handles that were used intraoperatively are not affected from the recall. Identification is possible based on visible marks on the strike plate. Return the reply form to us and confirm for each product the respective status. Used intraoperatively: You can keep the instrument. As good as new: Return the product to Waldemar Link GmbH & Co. KG. Replacement of affected products will not incur any costs to you. Should you have any questions on acquiring replacements for forthcoming surgeries, please contact your local sales representative or customer service for Link products. Please return the reply form to us in any event until the 16.09.2024 as documentation of the recall. This applies even if you have none of the listed products in stock or if these products do not exhibit the defect in question. | 2024-10-03 09:50
Ortho Clinical Diagnostics | VITROS Immunodiagnostic Products Anti‑HTLV I/II Controls | 18 | - | QuidelOrtho received complaints related to quality control QC failure with certain lots listed above of VITROS Anti HTLV I II Control generating lower than expected results below the lower limit of the established range of means Our investigation determined that the root cause of this issue is the instability of Control due to a specific lot of raw material used in the manufacture of the affected lots of VITROS Anti HTLV I II Control x D x D Customers are advised to discontinue using render unusable and discard your inventory of the affected lots of VITROS Anti HTLV I II Controls and QuidelOrtho will credit or replace your discarded inventory Please note that VITROS Anti HTLV I II Control is performing as expected however for credit replacement purposes both Control and Control must be discarded | • Discontinue using, render unusable, and discard your inventory of the affected lots of VITROS Immunodiagnostic Products Anti-HTLV I/II Controls 1 and 2. • Complete and return the enclosed Confirmation of Receipt form no later than Month DD, 2024. Upon receipt of your completed Confirmation of Receipt form, QuidelOrtho will credit or replace your discarded inventory. • Save this notification with your User Documentation or post this notification in your laboratory until the issue has been resolved by your laboratory’s receipt of credit or a replacement lot of VITROS Immunodiagnostic Products Anti-HTLV I/II Controls. • Please forward this notification if the affected product was distributed outside of your facility. • If your laboratory has experienced the issue and you have not already done so, please report the occurrence to your local Global Services Organization. | 2024-10-06 08:51
Smith Medical | Ventilator | 26 | - | Smiths Medical became aware of an issue involving a paraPAC Plus P and P ventilator where there is potential for inadvertent tidal volume knob movement from the original setting when set at high mL and low mL settings | check attached FSN | 2024-10-07 14:11
Olympus | UroPass | - | - | Olympus conducted an investigation after receiving complaints reporting broken dilator tips in the package and in patients during surgical procedures The investigation determined that exposing the UroPass product to Ultraviolet UV Radiation can cause brittleness of the device dilator tip which may lead to breakage Since April Olympus has received adverse event complaints reporting broken UroPass dilator tips for devices still within their shelf life x D To reduce the risk of UV exposure to your device s Olympus instructs users to implement the following actions x D Store individual UroPass Ureteral Access Sheath pouches away from ultraviolet UV light sources including sunlight and artificial light Exposure to UV light during storage periods may cause embrittlement of this device increasing potential for breakage and patient injury x D Risk to Health x D Exposure to Ultraviolet light may cause embrittlement of this device potentially leading to breakage of the UroPass tip Tip breakage may lead to a delay in initiating a procedure if the tip is broken in the package or discovered during use or may result in a foreign body remaining in the patient resulting in potentially prolonged operative time or an additional procedure to locate and remove the broken piece Additionally tissue damage or perforation of the ureter could occur due to exposed sharp edges x D | "1. Carefully read the content of this notification. 2. Inspect your inventory and determine if any of these devices with the model name specified above remain in inventory. Please check all areas of the facility/hospital. Add a copy of this notification with your remaining inventory. You may continue to use the products in accordance with the instructions regarding UV exposure: Store individual UroPass Ureteral Access Sheath pouches away from ultraviolet (UV) light sources (including sunlight and artificial light). Exposure to UV light during storage periods may cause embrittlement of this device, increasing potential for breakage and patient injury. 3. Olympus is not requiring the return of your UroPass device(s) as a result of this action. However, if you want to return the UroPass device(s) in your inventory, please contact Olympus representative with regard to return and reimbursement procedure. Olympus will issue a credit to your facility upon return of your affected product. 4. Ensure all personnel are completely knowledgeable and thoroughly trained on the content of this notification, including the new information regarding UV exposure. Olympus is in the process of updating the Instructions for Use with this information. " | 2024-10-07 18:33
Smith Medical | Ventilator | 26 | - | Smiths Medical became aware of an issue related to a potential for the patient outlet connector to loosen detach from the paraPAC Plus P and P ventilators impacting the active ventilation function | see attached FSN | 2024-10-08 14:43
Philips | Intera 1.5T Omni/Stellar | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Philips | SmartPath to dSream for 1.5T | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Philips Medical Systems Nederland B.V. | Ingenia 1.5T CX | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Philips | SmartPath to dStream for XR and 3.0T | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Philips Medical Systems Nederland B.V. | Ingenia 3.0T CX | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Philips Medical Systems Nederland B.V. | Achieva 3.0T | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Philips | Intera 1.5T | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Philips Medical Systems Nederland B.V. | Ingenia 1.5T | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Philips Medical Systems Nederland B.V. | Ingenia 3.0T | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Philips Medical Systems Nederland B.V. | Ingenia 1.5T S | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Philips Medical Systems Nederland B.V. | INGENIA ELITION X | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Philips Medical Systems Nederland B.V. | INGENIA AMBITION X | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Philips Medical Systems Nederland B.V. | INGENIA AMBITION S | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Philips | Evolution upgrade 1.5T | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Philips | Evolution upgrade 3.0T | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Philips | SmartPath to Ingenia Elition X | 19 | - | The tabletops of the Intera Achieva SmartPath to dStream Ingenia and Evolution upgrade MR systems have a docking interface which is located underneath the tabletop and attached with screws To establish a safe patient transportation the Tabletop is latched to the Trolley via a docking interface The screws fixating the tabletop docking interface may come loose potentially leading to disengagement or some instability of the tabletop when clamped on the FlexTrak Trolley | Philips Service will contact customers to arrange a visit by a Field Service Engineer to check their MR system and correct it if required. This will be planned via Field Change Order FCO78100553. Philips will begin contacting customers in October 2024. | 2024-10-08 14:55
Drägerwerk AG & Co. KGaA | Atlan | 60 | - | Failure of the ventilator motor | Phase 1) Send the safety notice to all customers as per the location list by 14.10.2024 Phase 2) Replacement of the ventilator motor by 2025-09-30 | 2024-10-08 16:00
Randox Laboratories Ltd | RX Imola | 6 | - | The RX Imola assay parameters for Urea Shot ApoA ApoB have been further updated in Parameter version v following on from previous customer notification REC The rerun settings sample rerun high prozone and rerun low and dilutions for auto rerun have been disabled ApoA ApoB diluent saline has been enabled to allow serial dilution during calibration | Contact Randox Service department service@randox.com or your local approved distributor engineer to arrange update to User Interface software version (129). Download and update parameters, or manually enter updates to avoid losing any previously obtained data, for assays performed on your Rx Imola. Screenshots are available upon request for reference, please contact technical.services@randox.com for more information. Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2024-10-09 08:39
Philips Image Guided Therapy Devices | OmniWire Pressure Guide Wire | 195 | - | An internal wire component could become damaged due to the wire potentially not meeting an internal specification | Philips is notifying all impacted customers of this issue. Philips is working with its component supplier to resolve this issue as soon as possible. All impacted customers will be notified by Philips sales organization once the issue is resolved. | 2024-10-09 10:24
Saudi Mais Co, for Medical Products | 3 way stopcock | 102580 | - | leakage issue from the connecter | We request you to please check your inventory for the product within the scope of this field action. Stop the distribution of stock of this Lot and Quarantine immediately. We will be replacing the remaining effected quantities with the good one. | 2024-10-09 14:26
GE Healthcare | Centricity Universal Viewer Zero Footprint | 2 | - | GE HealthCare has become aware of an issue in Centricity Universal Viewer Zero Footprint Client ZFP where there is a potential security vulnerability which could allow a malicious actor to access the system and potentially manipulate patient data x D x D There have been no injuries or unauthorized access to patient data reported as a result of this issue x D | You can continue to use your device. Please ensure all potential users in your facility are made aware of this safety notification. Please implement one of the actions below until your system has been corrected. Option 1: Restrict access to these systems from the internet. or Option 2: Implement additional safeguards such as a VPN. Please retain this document for your records. GE HealthCare will correct all affected products at no cost to you. A GE HealthCare representative will contact you to arrange for the correction. | 2024-10-10 13:44
Roche Diagnostics GmbH | HCYS, 200 tests | -1 | - | The root cause is still under investigation by Roche and after its completion appropriate corrective preventive x D actions will be defined and implemented | No customer-specific action is required because the issue can be limited to lot 747029, which expired 30-Sept- 2024. So far, no adverse medical events have been alleged by customers. In this case, no general recommendations with respect to the review of previous results can be given, taking into account that the issue may have occurred only under certain, rare circumstances, using one of the affected cassettes. Customers should follow their standard laboratory operating procedures. Any specific questions raised by the users should be addressed individually, considering all relevant clinical information. | 2024-10-10 16:43
Smith Medical | Portex Venn Reusable Tracheal Tube Introducer | 2477 | - | Smiths Medical has identified a potential for ingress of fluid into the device during reprocessing The ingress takes place at the rear of the device between the end cap and the tube This could lead to staining of the device or allowing the fluid to remain in the device In addition the Hypochlorite Solution ppm and the Acetic Acid disinfectants recommended in the IFU may be inadequate according to the disinfection standards for this type of device | Smiths Medical Actions: Smiths Medical is sending this notification to all customers who received affected product as listed above. Smiths Medical no longer distributes any of the affected products. Smiths Medical will provide full credit to affected customers. Customer Required Actions: 1) Check all inventory locations within your institution for the affected product numbers listed in the notification and discontinue use. Destroy all affected products following your institution’s process for destruction. If destroying is not immediately possible at your facility, then the product should be quarantined until disposal. 2) Share this notification with all potential users of the device, to ensure they are aware of this notification. If the devices are used at another location, please ensure this communication is delivered there. 3) Complete and return the attached Customer Response Form to EMEA-FSN@icumed.com within 10 days of receipt to acknowledge your understanding of this notification, even if you do not have the affected product. | 2024-10-11 16:02
Smith Medical | Portex Tracheal Tube Exchange Guide | 2477 | - | Smiths Medical has identified a potential for ingress of fluid into the device during reprocessing The ingress takes place at the rear of the device between the end cap and the tube This could lead to staining of the device or allowing the fluid to remain in the device In addition the Hypochlorite Solution ppm and the Acetic Acid disinfectants recommended in the IFU may be inadequate according to the disinfection standards for this type of device | Smiths Medical Actions: Smiths Medical is sending this notification to all customers who received affected product as listed above. Smiths Medical no longer distributes any of the affected products. Smiths Medical will provide full credit to affected customers. Customer Required Actions: 1) Check all inventory locations within your institution for the affected product numbers listed in the notification and discontinue use. Destroy all affected products following your institution’s process for destruction. If destroying is not immediately possible at your facility, then the product should be quarantined until disposal. 2) Share this notification with all potential users of the device, to ensure they are aware of this notification. If the devices are used at another location, please ensure this communication is delivered there. 3) Complete and return the attached Customer Response Form to EMEA-FSN@icumed.com within 10 days of receipt to acknowledge your understanding of this notification, even if you do not have the affected product. | 2024-10-11 16:02
Olympus | Single Use Biopsy Valve | 4 | - | Olympus has become aware that the lots identified above may be missing product identification and an expiration date on the sterile packaging The information that may be missing are the sterile and manufacturing lot numbers and expiration dates This information is clearly printed and can be found on the outer box or zipper bag label The affected lots were distributed from March to May Olympus has not received any complaints associated with this issue x D Olympus recommends that you do not use any of these products if you are unable to determine the expiration date x D | 1.Carefully read the content of this notification. 2.Inspect your inventory and identify any MAJ-1555 devices with the lot number(s) specified above. Please check all areas of the hospital to determine if any of these devices remain in inventory. 3.If you have the affected lots in stock, inspect the sterile packages to confirm the sterile lot, manufacturing lot and expiration date are clearly printed as illustrated in FSN. 4.The information missing from the sterile packages can be found on the outer box or zip bag as illustrated below: Carton Box Zipper Bag 5.In the event the sterile package is missing the sterile lot, manufacturing lot and/or expiration date, maintain the outer box or zip bag for your future reference. Ensure all personnel are completely knowledgeable and thoroughly trained on the content of this notification. 6.Olympus further recommends that you do not use any of these products if you are unable to determine the expiration date, please contact Olympus representative with regard to return and reimbursement procedure. Olympus will issue a credit to your facility upon return of your affected product. 7.If you have further distributed this product, identify your customers, and forward them this notification. | 2024-10-13 10:42
B. Braun Melsungen AG | Infusomat Space Line | 6 | - | Please find attached FSN letter | Please find attached FSN letter | 2024-10-13 12:26
Hyaltech Ltd | ophthalmic viscosurgical device | 2230 | - | Hyaltech has identified a hole defect within a single blister OVD product during routine shelf life confirmatory testing post product release x D The sterility of the delivery system within the blister may be compromised if the hole defect is present However the likelihood of receiving a product with a hole is very low It should be noted that the sterility of the OVD within the syringe is not compromised In addition there have been no reports of this hole defect from the field x D | • The Customer Field Safety Notification instructs customers to visually inspect product before use. • This is in line with the instruction in the IFU “ not to be used if sterile packaging has been damaged” • In the event they identify a hole defect in the blister they should not use the product • Customers should return to GMC for disposal | 2024-10-13 17:45
United Medical Industries Co. Ltd | UNIPRO | 1 | - | The user King Fahad Medical City Riyadh has logged a report to SFDA on ade sfda gov sa An e mail has been received from the Assigned NCMDR Officer as following x D x D From Muwayed S Al Khnaini MSKhnaini sfda gov sa x D Sent Tuesday September AM x D To mjallad info unimed com sa qa unimed com sa x D Cc AE MD ae md sfda gov sa ae sfda md gmail com x D Subject Incident UNIPRO SD D NCMDR SFDA x D x D Dear United Medical Industries Co Ltd x D x D As a part of mutual collaboration on keeping and maintaining safety performance and quality of medical devices placed in services within Saudi Market The National Center for Medical Devices Reporting NCMDR kindly request to review and investigate the below issue x D x D Description It was reported that x D x D The medical supply is not sticking to the skin tendency that it will loosen and the attached tube will not be secured causing disruption of procedure x D x D Device Medical Product Manufacturer Model Serial Lot No Location x D UNIPRO United Medical Industries Co Ltd SD D King Fahad Medical City | 1) Visited the user, King Fahad Medical City, met with the user to understand their concern. 2) Sent an Acknowledgement Letter on the quantities supplied and quantities available with the user. 3) Sent a letter to other customers who are using same product with different batch to gather feedback. 4) Initiated the investigation on the reported Batch | 2024-10-15 15:20
MERIVAARA | Q-FLOW SURGICAL LIGHTS | 7 | - | Due to an assembly error the device s support x D yokes can become loose from their fixings In the worst case if these fixing screws become completely loose the luminaire x D may fall off its fixture | Attached | 2024-10-16 15:16
Boston Scientific | POLARX BALLOON CATHETER | 6 | - | AE fistula is a known and inherent risk for patients undergoing catheter ablation for atrial fibrillation Although uncommon esophageal injury is a potentially life threatening complication due to proximity of the esophagus to the posterior left atrium Since commercial introduction of the POLARx Cryoablation System in Boston Scientific has received seven reports worldwide of AE fistula occurring following atrial fibrillation ablations four of these reports were associated with a patient death x D Detailed investigation of the available data associated with these AE fistula events did not identify product performance related issues with any component of the cryoablation system however frequency and intensity of cryoablation applications were observed as possible contributing factors Therefore Boston Scientific is updating the POLARx and POLARx FIT cryoablation balloon catheter IFUs to emphasize the risk of AE fistula as well as practices that may reduce this risk These IFU updates align with the FROzEN AF clinical trial Boston Scientific is communicating these IFU updates to all global customers and affected worldwide regulatory authorities to minimize the risk of AE fistula associated with use of the POLARx Cryoablation System Following applicable regulatory approval updated IFUs will be packaged and shipped with corresponding POLARx Cryoablation System devices | - Review the IFU Updates related to AE fistula and cryoablation application parameters from the FROzEN AF clinical trial, which demonstrated safety and effectiveness of the POLARx Cryoablation System. - Provide awareness of this information, share this communication with clinicians that use the Boston Scientific POLARx Cryoablation System, including the POLARx Catheter, the POLARx FIT Catheter and the SMARTFREEZE Console. | 2024-10-16 16:12
Philips Medical Systems Nederland B.V. | Vue PACS | 2 | - | A software issue affecting Philips Vue PACS may cause incorrect ischemic map and table value calculations when using the export function in the Perfusion Application for a Siemens CT perfusion study x D When a Siemens CT imaging study is initially loaded into the perfusion application and perfusion analysis is performed the ischemic map calculation is correct However after an export action is performed and the study is reloaded in the Perfusion Application the software issue may occur and the following warning message is generated | Actions that should be taken by the customer / user in order to prevent risks for patients or users • To prevent the issue from occurring, Philips recommends that users not use the perfusion application Export function (as shown in Figure 4). When used without export functionality, the perfusion application performs as expected. • You may continue to use your system(s) in accordance with the intended use and by following the recommendation above. • If a study is reloaded in the Perfusion Application after an export action is performed, and the Perfusion/Diffusion Warning message is observed (shown in Figure 1), the ischemic map and table value calculations should not be used for diagnostic purposes. • Circulate this notice to all users of this device so that they are aware of the potential issue. Please retain this letter with your system(s) until a solution is installed on your system; ensure the letter is in a place likely to be seen/viewed. • Please complete and return the attached response form to Philips promptly and no later than 30 days from receipt of this letter. Actions planned by Philips to correct the problem A Philips representative will contact you to schedule a time to install a software solution on your system(s) to resolve the issue (reference FCO78000003). Philips plans to begin installing the solution on affected systems in October 2024. Please be assured that maintaining the highest level of safety and quality is our greatest priority. If you need any further information or support concerning this issue, please contact your local Philips representative: met.quality@philips.com | 2024-10-17 10:48
Smiths Medical ASD, Inc. | Level 1® Normothermic I.V. Fluid Administration Set | 20 | - | Smiths Medical has become aware of users not being able to fit some Level DI Normothermic I V Fluid Administration Sets properly into the interlock block of the Level Fast Flow Fluid Warmer | Required Actions for Users: Please complete the following actions listed below: 1. Check all inventory locations within your institution for the affected product code listed in the notification and discontinue use. Destroy all DI-50 administration sets following your institution’s process for discarding. If destruction is not immediately possible at your facility, then the product should be quarantined until disposal. 2. Share this notification with all potential users of the devices to ensure they are aware of this notification and proposed mitigations. If the devices are used at another location, please ensure this communication is delivered there. 3. Complete and return the attached Customer Response Form to EMEA-FSN@icumed.com within 10 days of receipt to acknowledge your understanding of this notification, even if you do not have the affected product. 4. DISTRIBUTORS: If you have distributed potentially affected products to your customers, please immediately forward this notice to them and request that they complete the response form and return it to YOU. Then the DISTRIBUTOR must complete a SINGLE form with the required details and return to EMEA-FSN@icumed.com Follow up Actions by Smiths Medical: Smiths Medical can offer alternate products in place of the DI-50. Kindly contact customer service to learn about the alternate products and how to order them. Smiths Medical will provide full credit to affected customers. | 2024-10-17 15:59
Baxter Healthcare SA | Minicap Extended Life PD Transfer Set with Twist Clamp | 144 | - | The source of the NDL PCBAs and or NDL PCBs in those recalls was due to the manufacturing process of the silicone tubing | Baxter is in the process of evaluating whether these same risks are present with its PD transfer sets. At this time Baxter does not have data to definitively conclude whether there is a safety risk. Therefore, Baxter is informing you of the potential patient safety risk while our evaluation is in process. | 2024-10-20 11:36
Biomerieux | VIDAS® Anti HBs TOTAL II | 2 | - | bioM rieux received complaints from the field for negative control C issue out of range high with VIDAS Anti HBs TOTAL II T ref Lots x D The only risk associated to the referenced issue is a potential delay in reporting result there is no risk of false results x D The investigation confirmed that only VIDAS Anti HBs TOTAL II T ref Lot is impacted by the issue x D A continuous monitoring of all other lots available on the market is implemented and shows that all other lots of VIDAS Anti HBs TOTAL II T ref available on the market are conform x D The exact root cause is still under investigation x D | Upon receipt of this letter, if possible please order a new kit of VIDAS® Anti-HBs Total II (AHBS) Reference 30318 or replace kit of lot 1010308790 by another one in your inventory. There is no risk of false results, the kit of lot 1010308790 may continue to be used with a valid calibration. Discard the kits of lot 1010308790 remaining in your inventory before the VIDAS® instrument requires a new calibration | 2024-10-20 11:51
Olympus | Gastrointestinal Videoscope | 9 | - | Reason for Action x D Olympus previously issued revised cleaning instructions for the GIF TH in October to issue new updated channel drying steps Olympus conducted user evaluation through human factors validation and determined further clarification is required to improve the instructions for easier understanding of each drying step As a result Olympus is issuing a revised IFU addendum that informs customers that the channel drying time has been extended by minutes to purge water from the channels This is accomplished by drying each channel for minutes in the sequence described in the addendum Ensure all reprocessing personnel are completely knowledgeable and thoroughly trained on the new water purge requirements Olympus has received reportable complaints related to this issue | Actions Required: Our records indicate that your facility has purchased one or more of the affected products. The range of affected serial numbers is included at the top of this letter. Olympus requires you to take the following actions: 1. Carefully read the content of this notification. 2. Implement use of the enclosed Addendum stating that the channel drying time has been extended by 12 minutes to purge water from the channels. This is accomplished by drying each channel for 3 minutes, in the sequence described in the addendum. Ensure all reprocessing personnel are completely knowledgeable and thoroughly trained on the new water purge requirements. 3. Olympus requests that you acknowledge receipt of this letter. Indicate on the Reply Form that you have received and understood this notification by filling out and returning the completed enclosed Reply Form back to your local Olympus representative ra@olympus-mea.com latest by 15.11.2024. 4. If you have further distributed this product, identify your customers, and forward them this notification. | 2024-10-22 09:01
3M | 3M™ V.A.C.® Ulta Therapy Unit | 406 | - | Solventum formerly M Health Care values the care of our patients and the quality of our products As a result Solventum is initiating a voluntary field corrective action on the occlusion detection sensor located on all models and serial numbers of the M V A C Ulta Therapy Unit This voluntary field corrective action will include specific steps to be performed to test and if necessary replace the occlusion detection sensor on the therapy unit | 1. Locate all serial numbers and Therapy Unit Software version of V.A.C.® Ulta Therapy Units. 2. Perform the steps outlined in the attached test procedure for the occlusion detection sensor 3. Record the test results on the provided Field Notice Receipt and Customer Response Form. Complete all remaining fields, sign and return this form via email to kci3mfieldactionresponse@solventum.com. Additional copies of the response form may be required based on the number of units in your facility. 4. For any therapy units with a “FAIL” test result, contact your Solventum representative to coordinate the repair/replacement of the occlusion detection sensor with your facility. If any additional issues are identified during the repair, you will be notified, and it will be the responsibility of the unit owner to determine whether to proceed with the additional services and associated costs. | 2024-10-22 13:53
Baxter Healthcare SA | Minicap Extended Life PD Transfer Set with Twist Clamp | 178 | - | Separation of the female connector from the main body of the transfer set These separations may occur while connecting or disconnecting from PD therapy products | Baxter Healthcare Corporation is issuing a Safety Alert for the MiniCap Extended Life PD Transfer Sets Baxter has received reports of patients subsequently cutting the patient line and/or not following aseptic technique when trying to address the separation | 2024-10-22 15:42
Medtronic MiniMed | MiniMed® 640G | 5606 | - | Medtronic MiniMed received a complaint where a user reported an issue with the MiniMed G Pump in which the pump was depleting AA batteries every days Per the Product Requirements the pump battery life shall exceed days when using a new AA battery when tested per Nominal Use x D CAPA investigation determined that an increase in the electronic sleep current is the primary root cause of early battery depletion There were damaged capacitors identified on both printed circuit board assembly PCBA and PCBA particularly those located at position C on PCBA which were leaking current and the leakage current caused the increased sleep current The damage observed on the capacitors was determined to be primarily caused by electronic damage sustained through mechanical impact Returned pumps with confirmed moisture damage were excluded from the CAPA investigation scope as moisture damage is a known cause of early battery depletion x D | • Ensure you carry an extra set of new AA lithium or alkaline batteries or fully charged NiMH batteries should you need to replace your pump battery sooner than expected. • Always pay attention to alerts and alarms displayed on your pump and be prepared to replace the battery as soon as you receive the “Low Battery Pump” alert. Do not wait for the later battery alerts as these may occur when less battery life remains than the User Guide states. Refer to your pump’s User Guide for instructions on how to replace the battery. • If you notice any significant decrease in battery life, regardless of whether your pump has been dropped, bumped, or has experienced physical impact, contact our Helpline your Medtronic contact for troubleshooting support to determine if a replacement pump is needed. • Ensure you have back-up insulin therapy available at all times in your emergency kit in case it's needed. | 2024-10-22 17:20
Medtronic MiniMed | MiniMed 670G | 5606 | - | Medtronic MiniMed received a complaint where a user reported an issue with the MiniMed G Pump in which the pump was depleting AA batteries every days Per the Product Requirements the pump battery life shall exceed days when using a new AA battery when tested per Nominal Use x D CAPA investigation determined that an increase in the electronic sleep current is the primary root cause of early battery depletion There were damaged capacitors identified on both printed circuit board assembly PCBA and PCBA particularly those located at position C on PCBA which were leaking current and the leakage current caused the increased sleep current The damage observed on the capacitors was determined to be primarily caused by electronic damage sustained through mechanical impact Returned pumps with confirmed moisture damage were excluded from the CAPA investigation scope as moisture damage is a known cause of early battery depletion x D | • Ensure you carry an extra set of new AA lithium or alkaline batteries or fully charged NiMH batteries should you need to replace your pump battery sooner than expected. • Always pay attention to alerts and alarms displayed on your pump and be prepared to replace the battery as soon as you receive the “Low Battery Pump” alert. Do not wait for the later battery alerts as these may occur when less battery life remains than the User Guide states. Refer to your pump’s User Guide for instructions on how to replace the battery. • If you notice any significant decrease in battery life, regardless of whether your pump has been dropped, bumped, or has experienced physical impact, contact our Helpline your Medtronic contact for troubleshooting support to determine if a replacement pump is needed. • Ensure you have back-up insulin therapy available at all times in your emergency kit in case it's needed. | 2024-10-22 17:20
Medtronic MiniMed | MiniMed 740G | 5606 | - | Medtronic MiniMed received a complaint where a user reported an issue with the MiniMed G Pump in which the pump was depleting AA batteries every days Per the Product Requirements the pump battery life shall exceed days when using a new AA battery when tested per Nominal Use x D CAPA investigation determined that an increase in the electronic sleep current is the primary root cause of early battery depletion There were damaged capacitors identified on both printed circuit board assembly PCBA and PCBA particularly those located at position C on PCBA which were leaking current and the leakage current caused the increased sleep current The damage observed on the capacitors was determined to be primarily caused by electronic damage sustained through mechanical impact Returned pumps with confirmed moisture damage were excluded from the CAPA investigation scope as moisture damage is a known cause of early battery depletion x D | • Ensure you carry an extra set of new AA lithium or alkaline batteries or fully charged NiMH batteries should you need to replace your pump battery sooner than expected. • Always pay attention to alerts and alarms displayed on your pump and be prepared to replace the battery as soon as you receive the “Low Battery Pump” alert. Do not wait for the later battery alerts as these may occur when less battery life remains than the User Guide states. Refer to your pump’s User Guide for instructions on how to replace the battery. • If you notice any significant decrease in battery life, regardless of whether your pump has been dropped, bumped, or has experienced physical impact, contact our Helpline your Medtronic contact for troubleshooting support to determine if a replacement pump is needed. • Ensure you have back-up insulin therapy available at all times in your emergency kit in case it's needed. | 2024-10-22 17:20
Medtronic MiniMed | MiniMed 780G | 5606 | - | Medtronic MiniMed received a complaint where a user reported an issue with the MiniMed G Pump in which the pump was depleting AA batteries every days Per the Product Requirements the pump battery life shall exceed days when using a new AA battery when tested per Nominal Use x D CAPA investigation determined that an increase in the electronic sleep current is the primary root cause of early battery depletion There were damaged capacitors identified on both printed circuit board assembly PCBA and PCBA particularly those located at position C on PCBA which were leaking current and the leakage current caused the increased sleep current The damage observed on the capacitors was determined to be primarily caused by electronic damage sustained through mechanical impact Returned pumps with confirmed moisture damage were excluded from the CAPA investigation scope as moisture damage is a known cause of early battery depletion x D | • Ensure you carry an extra set of new AA lithium or alkaline batteries or fully charged NiMH batteries should you need to replace your pump battery sooner than expected. • Always pay attention to alerts and alarms displayed on your pump and be prepared to replace the battery as soon as you receive the “Low Battery Pump” alert. Do not wait for the later battery alerts as these may occur when less battery life remains than the User Guide states. Refer to your pump’s User Guide for instructions on how to replace the battery. • If you notice any significant decrease in battery life, regardless of whether your pump has been dropped, bumped, or has experienced physical impact, contact our Helpline your Medtronic contact for troubleshooting support to determine if a replacement pump is needed. • Ensure you have back-up insulin therapy available at all times in your emergency kit in case it's needed. | 2024-10-22 17:20
Integra LifeSciences Production Corporation | CODMAN® Surgical Patties | 1629 | - | During an internal investigation Integra LifeSciences identified higher than expected levels of endotoxin within the raw material used to produce Codman Surgical Patties and Strips that may have resulted in out of specification levels of endotoxin in those finished goods Consequently while the endotoxin levels identified were higher than expected the possibility of adverse health consequences actually occurring remains remote | 1. Please review and understand the information provided in this letter. 2. If you do have affected products in your warehouse: a. Quarantine them immediately. b. Check the box “I do have affected unit(s)” in the enclosed reply form. c. Record on the table 2. at the bottom of the reply form the total quantity of affected product(s) and lot number(s) that you have. 3. If you do not have affected product(s) in your warehouse, check the box, “I do not have affected unit(s)”. 4. Please check your customer traceability records for shipments of affected products. 5. Forward a copy of the enclosed Field Safety Notice to any of your customers that have purchased the affected products. 6. Please return the completed Reply form by email to emea-fsca-neuro@integralife.com, or Fax to +33 (0)4.37.47. 59.30. By filling in this form, you confirm that you have received this Safety Notice and you intend to fully comply with this notification. We expect a response within 3 weeks. You also confirm that this notification has been forwarded to every concerned person in your organization. 7. At receipt of the reply form, and if it is noted that you or your customers have affected product(s), Customer Service will contact you and provide an RMA number and directions to return the product(s). A credit will be processed upon receipt of returned goods. 8. We recommend that you retain a copy of the form for your records. | 2024-10-22 20:47
Welch Allyn, Inc. | Welch Allyn® HScribe™ | 25 | - | Baxter Healthcare Corporation is issuing an Urgent Field Safety Notice for the HScribe Q Stress RScribe Vision Express Welch Allyn Diagnostic Cardiology Suite Connex Cardio and XScribe products listed below due to the potential for exam files being assigned duplicate Unique Identifiers UIDs The UID as defined by the Digital Imaging and Communications in Medicine DICOM standard is created based on the date and time of acquisition Separate files created at the same time down to the second and exported to the same DICOM storage system are assigned the same UID If the system receiving the DICOM file e g Picture Archiving and Communication System PACS relies solely on the UID to accept exams this could lead to a mismatch of the patient s identification with their physiological data x D x D Baxter has identified that the following events must occur for the potential mismatch of a patient s identification with their physiological data x D x D Two or more exams are acquired with the affected Baxter devices at the same second and exported to the same DICOM storage system e g PACS which results in the same UID x D The receiving system that communicates with the affected Baxter devices via DICOM relies solely on the UID to accept the exam transmission x D | 1. Contact the system IT administrator to confirm that your system is impacted by this issue. This action is only necessary if the interfacing system (e.g., PACS) is configured to receive DICOM results and relies solely on the UID for results - patient/order matching. 2. For the products listed below, a correction can be implemented by the system IT administrator to address this issue and is already available in your current version of software. The New Series Instance UID option will need to be activated. Baxter will work with you to confirm, if applicable, that the correction has been implemented. NOTE: Baxter recommends that before proceeding with the corrections noted below, any changes to your interface configurations should be evaluated in a test environment before deploying in a clinical environment. 3. For the products listed below, please contact Baxter Technical Service to request and schedule the correction. Please note you will be receiving this correction from Baxter at no charge. 4. Complete the enclosed customer reply form and return it to Baxter by scanning and e-mailing it to Bandar_Alosaimi@Baxter.com, even if you don’t have any inventory. Returning the customer reply form promptly will confirm your receipt of this notification and prevent you from receiving repeat notices. 5. If you distribute this product to other facilities or departments within your institution, please forward a copy of this communication to them. 6. If you purchased this product from a distributor, please note that the Baxter customer reply form is not applicable. If a reply form is provided by your distributor or wholesaler, please return it to the supplier according to their instructions. 7. If you are a dealer, wholesaler, distributor/reseller, or original equipment manufacturer (OEM) that distributed any affected product to other facilities, please notify your customers of this Correction in accordance with your customary procedures. | 2024-10-22 21:10
Welch Allyn, Inc. | Welch Allyn®RScribe™ | 25 | - | Baxter Healthcare Corporation is issuing an Urgent Field Safety Notice for the HScribe Q Stress RScribe Vision Express Welch Allyn Diagnostic Cardiology Suite Connex Cardio and XScribe products listed below due to the potential for exam files being assigned duplicate Unique Identifiers UIDs The UID as defined by the Digital Imaging and Communications in Medicine DICOM standard is created based on the date and time of acquisition Separate files created at the same time down to the second and exported to the same DICOM storage system are assigned the same UID If the system receiving the DICOM file e g Picture Archiving and Communication System PACS relies solely on the UID to accept exams this could lead to a mismatch of the patient s identification with their physiological data x D x D Baxter has identified that the following events must occur for the potential mismatch of a patient s identification with their physiological data x D x D Two or more exams are acquired with the affected Baxter devices at the same second and exported to the same DICOM storage system e g PACS which results in the same UID x D The receiving system that communicates with the affected Baxter devices via DICOM relies solely on the UID to accept the exam transmission x D | 1. Contact the system IT administrator to confirm that your system is impacted by this issue. This action is only necessary if the interfacing system (e.g., PACS) is configured to receive DICOM results and relies solely on the UID for results - patient/order matching. 2. For the products listed below, a correction can be implemented by the system IT administrator to address this issue and is already available in your current version of software. The New Series Instance UID option will need to be activated. Baxter will work with you to confirm, if applicable, that the correction has been implemented. NOTE: Baxter recommends that before proceeding with the corrections noted below, any changes to your interface configurations should be evaluated in a test environment before deploying in a clinical environment. 3. For the products listed below, please contact Baxter Technical Service to request and schedule the correction. Please note you will be receiving this correction from Baxter at no charge. 4. Complete the enclosed customer reply form and return it to Baxter by scanning and e-mailing it to Bandar_Alosaimi@Baxter.com, even if you don’t have any inventory. Returning the customer reply form promptly will confirm your receipt of this notification and prevent you from receiving repeat notices. 5. If you distribute this product to other facilities or departments within your institution, please forward a copy of this communication to them. 6. If you purchased this product from a distributor, please note that the Baxter customer reply form is not applicable. If a reply form is provided by your distributor or wholesaler, please return it to the supplier according to their instructions. 7. If you are a dealer, wholesaler, distributor/reseller, or original equipment manufacturer (OEM) that distributed any affected product to other facilities, please notify your customers of this Correction in accordance with your customary procedures. | 2024-10-22 21:10
Welch Allyn, Inc. | Welch Allyn®XScribe™ | 25 | - | Baxter Healthcare Corporation is issuing an Urgent Field Safety Notice for the HScribe Q Stress RScribe Vision Express Welch Allyn Diagnostic Cardiology Suite Connex Cardio and XScribe products listed below due to the potential for exam files being assigned duplicate Unique Identifiers UIDs The UID as defined by the Digital Imaging and Communications in Medicine DICOM standard is created based on the date and time of acquisition Separate files created at the same time down to the second and exported to the same DICOM storage system are assigned the same UID If the system receiving the DICOM file e g Picture Archiving and Communication System PACS relies solely on the UID to accept exams this could lead to a mismatch of the patient s identification with their physiological data x D x D Baxter has identified that the following events must occur for the potential mismatch of a patient s identification with their physiological data x D x D Two or more exams are acquired with the affected Baxter devices at the same second and exported to the same DICOM storage system e g PACS which results in the same UID x D The receiving system that communicates with the affected Baxter devices via DICOM relies solely on the UID to accept the exam transmission x D | 1. Contact the system IT administrator to confirm that your system is impacted by this issue. This action is only necessary if the interfacing system (e.g., PACS) is configured to receive DICOM results and relies solely on the UID for results - patient/order matching. 2. For the products listed below, a correction can be implemented by the system IT administrator to address this issue and is already available in your current version of software. The New Series Instance UID option will need to be activated. Baxter will work with you to confirm, if applicable, that the correction has been implemented. NOTE: Baxter recommends that before proceeding with the corrections noted below, any changes to your interface configurations should be evaluated in a test environment before deploying in a clinical environment. 3. For the products listed below, please contact Baxter Technical Service to request and schedule the correction. Please note you will be receiving this correction from Baxter at no charge. 4. Complete the enclosed customer reply form and return it to Baxter by scanning and e-mailing it to Bandar_Alosaimi@Baxter.com, even if you don’t have any inventory. Returning the customer reply form promptly will confirm your receipt of this notification and prevent you from receiving repeat notices. 5. If you distribute this product to other facilities or departments within your institution, please forward a copy of this communication to them. 6. If you purchased this product from a distributor, please note that the Baxter customer reply form is not applicable. If a reply form is provided by your distributor or wholesaler, please return it to the supplier according to their instructions. 7. If you are a dealer, wholesaler, distributor/reseller, or original equipment manufacturer (OEM) that distributed any affected product to other facilities, please notify your customers of this Correction in accordance with your customary procedures. | 2024-10-22 21:10
Randox Laboratories | HUM ASY CONTROL 3 | 1264 | - | Randox Laboratories can confirm there has been a transcription error in the RX Series section of the Instructions for Use IFU for Human Assayed Multi Sera Level HE Lots UE UE and UE | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. | 2024-10-23 11:18
Abbott GmbH. & Co. KG | Alinity ci-series System Control Module | 302 | - | Please find attached FSN letter | Please find attached FSN letter | 2024-10-23 11:38
Gambro Lundia AB | AK 98 | 179 | - | The source of the NDL PCBAs and or NDL PCBs in those recalls was due to the manufacturing process of the silicone tubing which used a chlorinated peroxide initiator | Baxter is in the process of evaluating whether these risks are present with the AK 98 machines. . Available information indicates that NDL PCBA and NDL PCB levels will decrease over time with machine use. | 2024-10-23 14:29
Maquet Cardiopulmonary GmbH | HLS Set Advanced | 1 | - | negative pressure sensor at initiation of ECMO was reading then the reading decreased until it reached Zero even after that it started to go up until it reach At the same time all the other pressures Pre membrane and post membrane was normal x D the patient is okay and still alive under support of ECMO x D the team in hospital tried to calibrate the Venous pressure sensor but still reading is positive then they checked the pressure cable and it was intact and in place x D the team decided to change the machine and after that the problem still existing the pre pump reading in positive x D and the team is going now to change the membrane | the team decided to change the machine and after that the problem still existing .( the pre pump reading in positive ) and the team is going now to change the membrane . | 2024-10-24 18:15
Radiometer Medical ApS | TCM5 | 0 | - | Radiometer has become aware of an issue with the TCM monitor with software versions or for application freezing | Per the customer advisory letter, users are requested to: • Immediately stop using the affected product, as detailed above, until the TCM5 software has been downgraded. • Complete the Recall Response Form (on the last page of this letter) and return it to your Radiometer representative within one week of receiving it. • For TCM5 monitors used in sleep studies only: o If the application freezes, switch the monitor off and then on again to restart the application. | 2024-10-25 15:27
Beckman Coulter Inc. | Access Substrate | 10035 | - | In February Beckman Coulter issued an Urgent Medical Device Recall letter regarding Access Substrate bottles that may have been received without labels This letter serves as an update to include additional affected lots beyond those initially listed | Verify upon receipt that all Access Substrate bottles have a label.  Discard all Access Substrate bottles without labels, according to your local regulations.  Contact your Beckman Coulter representative and request a replacement order if your laboratory receives Access Substrate bottles that are missing a label.  For customers in United States and Canada: Complete the attached Replacement Order Form. | 2024-10-27 22:24
Getinge | 1009.25A0 | 7 | - | Getinge has been determined that with Getinge IR remote control A in some cases no function occurs and therefore no movement of the operating table is possible Please refer to the attached customer letter for details | Software update for all affected Universal Remote Controls 100925A0 by a Getinge service technician according to Service Manual 100925A0, chapter 6.3. | 2024-10-28 12:14
ADVANCED MEDICAL SOLUTIONS LTD. | ACTIVHEAL Non-Adhesive PHMB Foam | 1880 | - | defect on primary packaging pouches in which minor missing patches of Polyethylene have x D been detected The defect on these pouches could compromise the device s ability to maintain a sterile barrier | Immediately inspect your internal inventory for the aforementioned packaging defect and quarantine all affected Product pending safe destruction. | 2024-10-28 20:16
NOxBOX Ltd | NOxBOXi® | 32 | - | Malfunctions related to manifold check valves for NOxBOXi devices x D | 1. Check your device inventory and any devices presently being used on patients to determine if you have an affected NOxBOXi device, using the list provided in Attachment A. 2. Please complete and return the Customer Response Form provided in Attachmen | 2024-10-29 11:00
BK Medical Aps | bk medical | 17 | - | BK Medical a GE HealthCare company has become aware of an issue where the internal power connections at the power inlet of the device can come loose for bk bk bk bkActiv ultrasound devices with the battery option In the unlikely event that this goes unnoticed there is the potential for excessive heat which could lead to a thermal event or the device becoming inoperable x D There have been no injuries reported as a result of this issue x D | You can continue to use your device by following the instructions below: Until GE HealthCare implements a permanent correction on your device for this issue, please inspect the power cord connection at the power inlet to the ultrasound device (See Figure 1). - If a cord retaining bracket is not present, ensure the power cord is fully pushed into the power inlet before each use and while the ultrasound device is charging. - If a cord retaining bracket is present, no further action is required. | 2024-10-29 13:12
KARL STORZ SE & Co. KG | Cutter, sterile | 3 | - | During an internal validation it was determined that there was holes in the sterile barrier x D system and due to the compromised sterile packaging the affected products are therefore x D being recalled | The user shall immediately discontinue the use of the affected products and shall send back the affected Products. The action needs to be acknowledged by returning the Acknowledgement Form. | 2024-10-29 14:56
Xiantao Tongda Nonwoven Products Co Ltd | BRYCEMED | - | عدم حصول المستلزم على إذن تسويق | - | Recall of all medical devices affected by the safety alert. | 2024-10-30 12:06
TG MEDICAL SDN. BHD. | Safari Discovery Series Latex medical examination gloves | - | - | Not conforming to the technical specification ISO Annex A Water tightness testing for Medical | • Stop using all affected products. • Return all affected products. | 2024-10-30 12:42
GE Medical Systems, LLC | Signa Architect | 1 | - | GE HealthCare has become aware that gradient coils for certain MR systems see affected product list in this letter under specific conditions can produce elevated acoustic noise during scanning As a result acoustic levels may exceed dB the limit established by International Electrotechnical Commission IEC when using hearing protection with a Noise Reduction Rating NRR of dB as currently required in the Operator Manual Prolonged periods of exposure to elevated acoustic noise could potentially lead to hearing impairment | You can continue to use your MR system by following the instructions below: Ensure hearing protection with a Noise Reduction Rating (NRR) of at least 33dB is used during exams on these systems. 200 pairs are provided with this letter. Please post this letter in your facility on or near the MR operator console. Please ensure all potential users in your facility are made aware of this safety notification and the recommended actions. Please retain this document for your records | 2024-11-03 13:24
GE Hualun Medical Systems Co. Ltd | Proteus XR/a | 5 | - | GE HealthCare has become aware that steel cables in the Wall Stand for certain Proteus XR a systems serviced by GE HealthCare have not been replaced at the intervals detailed in the Preventative Maintenance Action List GEHC Service Manual If the cables are not replaced as required in the GEHC Service Manual a fatigue related failure could result causing the bucky to drop downward on the track toward the floor This could cause injury during some types of exams No injuries have been reported related to this issue | You can continue to use the system for table exams without limitations. • You can continue to use the wall stand for exams where the patient does not have any part of their body underneath the bucky. • Do not perform exams where the patient has any body part underneath the bucky until the Wall Stand steel cable has been inspected and if necessary, has been replaced. Please ensure all potential users in your facility are made aware of this safety notification and the recommended actions. Please retain this document for your records. | 2024-11-03 14:20
Becton Dickinson and Company | Product Cybersecurity | 620 | - | Through our cybersecurity monitoring tools and processes BD recently identified unauthorised access to a limited portion of its information technology IT environment BD terminated the unauthorised access and applied additional security measures After a thorough investigation BD confirmed that product service credentials intended for use by BD technical support teams for certain BD products were accessed by an unauthorised actor Until these product service credentials are updated there is a risk of unauthorised access that may impact the confidentiality integrity and or availability of the relevant products and associated system or data x D x D To date BD has not been made aware of any unauthorised use of these product service credentials and has received no reports of these credentials being used for unauthorised access to any BD device x D | Through our cybersecurity monitoring tools and processes, BD recently identified unauthorised access to a limited portion of its information technology (IT) environment. BD terminated the unauthorised access and applied additional security measures. After a thorough investigation, BD confirmed that product service credentials intended for use by BD technical support teams for certain BD products were accessed by an unauthorised actor. Until these product service credentials are updated, there is a risk of unauthorised access that may impact the confidentiality, integrity and/or availability of the relevant products and associated system or data. To date, BD has not been made aware of any unauthorised use of these product service credentials and has received no reports of these credentials being used for unauthorised access to any BD device. | 2024-11-04 09:17
Canon | Interventional Angiography System | 1 | - | When SPOT Fluoro is first turned on the default ROI is selected and the collimator blades move to the boundary of the ROI and the static image is overlayed around the ROI covering the collimated area If customer changes the ROI the collimator blades and static images are adjusted accordingly However when the acquisition program is changed the collimator returns to the default ROI position which may be larger than the specified one leaving the X ray exposed area partially covered by the static image This is a software problem with the SPOT Fluoro program and will not occur if the function is not used | A software update is being developed to correct the problem. Customers will be contacted when it is available. Until the software update is available customers are advised to not use the Spot Fluoro function. This is a software problem with the SPOT Fluoro program and will not occur if the function is not used. | 2024-11-04 14:58
Philips Medical Systems Nederland B.V. | Achieva 1.5T | 61 | - | Philips has identified an issue with patient set up while using the MR system breast coil where if the cross section of the prone patient breast coil and patient table exceeds the internal diameter of the magnet bore the patient may be compressed between the breast coil and top of the magnet bore potentially resulting in harm to the patient see Image Additionally in cases of compromised bone strength the pressure points created between the prone patient and the breast coil could be exaggerated due to exam duration the machines vibrations and or patient positioning that induces pain resulting in harm to the patient x D Philips has received reports globally of patient harm associated with this issue as of September | Philips is providing this Field Safety Notice (FSN) Letter which contains recommendations for continued use of the systems referenced in Section 4(Actions that should be taken by the customer / user in order to prevent risks for patients). | 2024-11-04 15:43
Philips Medical Systems Nederland B.V. | Ingenia 1.5T | 61 | - | Philips has identified an issue with patient set up while using the MR system breast coil where if the cross section of the prone patient breast coil and patient table exceeds the internal diameter of the magnet bore the patient may be compressed between the breast coil and top of the magnet bore potentially resulting in harm to the patient see Image Additionally in cases of compromised bone strength the pressure points created between the prone patient and the breast coil could be exaggerated due to exam duration the machines vibrations and or patient positioning that induces pain resulting in harm to the patient x D Philips has received reports globally of patient harm associated with this issue as of September | Philips is providing this Field Safety Notice (FSN) Letter which contains recommendations for continued use of the systems referenced in Section 4(Actions that should be taken by the customer / user in order to prevent risks for patients). | 2024-11-04 15:43
Philips Medical Systems Nederland B.V. | Ingenia 1.5T S | 61 | - | Philips has identified an issue with patient set up while using the MR system breast coil where if the cross section of the prone patient breast coil and patient table exceeds the internal diameter of the magnet bore the patient may be compressed between the breast coil and top of the magnet bore potentially resulting in harm to the patient see Image Additionally in cases of compromised bone strength the pressure points created between the prone patient and the breast coil could be exaggerated due to exam duration the machines vibrations and or patient positioning that induces pain resulting in harm to the patient x D Philips has received reports globally of patient harm associated with this issue as of September | Philips is providing this Field Safety Notice (FSN) Letter which contains recommendations for continued use of the systems referenced in Section 4(Actions that should be taken by the customer / user in order to prevent risks for patients). | 2024-11-04 15:43
Philips Medical Systems Nederland B.V. | Ingenia 3.0T | 61 | - | Philips has identified an issue with patient set up while using the MR system breast coil where if the cross section of the prone patient breast coil and patient table exceeds the internal diameter of the magnet bore the patient may be compressed between the breast coil and top of the magnet bore potentially resulting in harm to the patient see Image Additionally in cases of compromised bone strength the pressure points created between the prone patient and the breast coil could be exaggerated due to exam duration the machines vibrations and or patient positioning that induces pain resulting in harm to the patient x D Philips has received reports globally of patient harm associated with this issue as of September | Philips is providing this Field Safety Notice (FSN) Letter which contains recommendations for continued use of the systems referenced in Section 4(Actions that should be taken by the customer / user in order to prevent risks for patients). | 2024-11-04 15:43
Philips Medical Systems Nederland B.V. | INGENIA AMBITION S | 61 | - | Philips has identified an issue with patient set up while using the MR system breast coil where if the cross section of the prone patient breast coil and patient table exceeds the internal diameter of the magnet bore the patient may be compressed between the breast coil and top of the magnet bore potentially resulting in harm to the patient see Image Additionally in cases of compromised bone strength the pressure points created between the prone patient and the breast coil could be exaggerated due to exam duration the machines vibrations and or patient positioning that induces pain resulting in harm to the patient x D Philips has received reports globally of patient harm associated with this issue as of September | Philips is providing this Field Safety Notice (FSN) Letter which contains recommendations for continued use of the systems referenced in Section 4(Actions that should be taken by the customer / user in order to prevent risks for patients). | 2024-11-04 15:43
Philips Medical Systems Nederland B.V. | INGENIA AMBITION X | 61 | - | Philips has identified an issue with patient set up while using the MR system breast coil where if the cross section of the prone patient breast coil and patient table exceeds the internal diameter of the magnet bore the patient may be compressed between the breast coil and top of the magnet bore potentially resulting in harm to the patient see Image Additionally in cases of compromised bone strength the pressure points created between the prone patient and the breast coil could be exaggerated due to exam duration the machines vibrations and or patient positioning that induces pain resulting in harm to the patient x D Philips has received reports globally of patient harm associated with this issue as of September | Philips is providing this Field Safety Notice (FSN) Letter which contains recommendations for continued use of the systems referenced in Section 4(Actions that should be taken by the customer / user in order to prevent risks for patients). | 2024-11-04 15:43
Philips Medical Systems Nederland B.V. | INGENIA ELITION X | 61 | - | Philips has identified an issue with patient set up while using the MR system breast coil where if the cross section of the prone patient breast coil and patient table exceeds the internal diameter of the magnet bore the patient may be compressed between the breast coil and top of the magnet bore potentially resulting in harm to the patient see Image Additionally in cases of compromised bone strength the pressure points created between the prone patient and the breast coil could be exaggerated due to exam duration the machines vibrations and or patient positioning that induces pain resulting in harm to the patient x D Philips has received reports globally of patient harm associated with this issue as of September | Philips is providing this Field Safety Notice (FSN) Letter which contains recommendations for continued use of the systems referenced in Section 4(Actions that should be taken by the customer / user in order to prevent risks for patients). | 2024-11-04 15:43
Philips | Options for R5 dStream | 61 | - | Philips has identified an issue with patient set up while using the MR system breast coil where if the cross section of the prone patient breast coil and patient table exceeds the internal diameter of the magnet bore the patient may be compressed between the breast coil and top of the magnet bore potentially resulting in harm to the patient see Image Additionally in cases of compromised bone strength the pressure points created between the prone patient and the breast coil could be exaggerated due to exam duration the machines vibrations and or patient positioning that induces pain resulting in harm to the patient x D Philips has received reports globally of patient harm associated with this issue as of September | Philips is providing this Field Safety Notice (FSN) Letter which contains recommendations for continued use of the systems referenced in Section 4(Actions that should be taken by the customer / user in order to prevent risks for patients). | 2024-11-04 15:43
Philips Medical Systems Nederland B.V. | Panorama HFO | 61 | - | Philips has identified an issue with patient set up while using the MR system breast coil where if the cross section of the prone patient breast coil and patient table exceeds the internal diameter of the magnet bore the patient may be compressed between the breast coil and top of the magnet bore potentially resulting in harm to the patient see Image Additionally in cases of compromised bone strength the pressure points created between the prone patient and the breast coil could be exaggerated due to exam duration the machines vibrations and or patient positioning that induces pain resulting in harm to the patient x D Philips has received reports globally of patient harm associated with this issue as of September | Philips is providing this Field Safety Notice (FSN) Letter which contains recommendations for continued use of the systems referenced in Section 4(Actions that should be taken by the customer / user in order to prevent risks for patients). | 2024-11-04 15:43
Philips | SmartPath to dStream for 1.5T | 61 | - | Philips has identified an issue with patient set up while using the MR system breast coil where if the cross section of the prone patient breast coil and patient table exceeds the internal diameter of the magnet bore the patient may be compressed between the breast coil and top of the magnet bore potentially resulting in harm to the patient see Image Additionally in cases of compromised bone strength the pressure points created between the prone patient and the breast coil could be exaggerated due to exam duration the machines vibrations and or patient positioning that induces pain resulting in harm to the patient x D Philips has received reports globally of patient harm associated with this issue as of September | Philips is providing this Field Safety Notice (FSN) Letter which contains recommendations for continued use of the systems referenced in Section 4(Actions that should be taken by the customer / user in order to prevent risks for patients). | 2024-11-04 15:43
ARGUS 717V and ARGUS 718V | ARGUS 717V and ARGUS 718V | 70 | - | This problem involves the power board mainboard female on board Micro Match connector the flat x D ribbon cable and the Micro Match male connector If any of these components are defective the x D connection between the mainboard and the power board can be compromised x D The flat ribbon cable and connectors are particularly vulnerable to damage if the cable is pulled out by x D tugging on the wires during servicing x D Additionally during a Standard Safety Check SSC or repairs if the back and front parts of the A XV x D are opened too far without disconnecting the cable it can create tension on the cable and connectors x D This excessive mechanical stress can damage the crimping between the male connector and the cable x D leading to a compromised connection between the mainboard and power board x D In case this problem happens the device could in rare cases unexpectedly shutdown and generate a x D technical error TE and or TE | the flat ribbon cable (REF: 601342 - flat ribbon cable 24-pole) shall always be removed from the female connector by holding and pulling the male connector, not by tugging on the wires. • if the cable of the identified device has been subjected to high mechanical stress at any point during the pump's lifetime, either through stretching during repairs or improper disconnection, the cable should be replaced with a new one (REF: 601342 - flat ribbon cable 24-pole). • in the case of any future service, it is recommended to replace the flat ribbon cable with a new one (REF: 601342 - flat ribbon cable 24-pole), once it has been disconnected from the power board/mainboard | 2024-11-04 16:07
TG MEDICAL SDN. BHD. | Safari Discovery Series Latex medical examination gloves | 2996 | - | Non conformance to ISO Annex A Water tightness testing for Medical | 1- company already implement the requested ISO from 2023. 2-NO available stock in the customers or market as it is non moving item | 2024-11-05 09:33
Siemens Healthcare Diagnostics Inc. | Atellica CH Reaction Cuvette Segments | 1034 | - | Issue Description x D Siemens Healthineers has confirmed through the internal investigation of customer complaints x D that there is a potential for imprecision with seven assays identified to have optical sensitivity x D when using a subset of lots of Atellica CH Reaction Ring Cuvette Segments on some Atellica CH x D analyzers See Appendix A Table The lot numbers of potentially impacted Atellica CH Reaction x D Ring Cuvette Segments begin with N on the packaging or on the underside of the cuvette x D segment Not all Atellica CH analyzers are impacted by this issue x D Quality Control QC failures may alert customers to the issue before patient samples are x D processed See Table in the Appendix A for impacted assay bias data x D This issue does not impact Atellica CH Reaction Ring Cuvette Segments which have lot numbers x D that begin with x D Kindly review the attached Urgent Field Safety Notice | Customer Actions  Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable.  Review your current Atellica CH Reaction Ring Cuvette Segment stock. All Atellica CH Reaction Ring Cuvette Segments beginning with “11” are not part of this communication and can be used without restriction. Any other lot in your stock must be immediately discarded.  Locate your analyzer’s serial number within Appendix B Serial Number List. If your analyzer’s serial number is in the provided list perform the instructions to assess, inspect and/or replace Atellica CH Reaction Ring Cuvettes Segments provided in the Appendix section.  Determine your laboratory’s replacement needs and provide detailed information to Siemens Healthineers for reporting to authorities. Complete and return the Field Correction Effectiveness Check Form attached to this letter within thirty (30) days.  Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached Urgent Field Safety Notice | 2024-11-05 17:11
M.A.Med Alliance SA | SELUTION SLR™ 018 PTA BALLOON CATHETER | 5 | - | Cordis has identified that there is potential for deflation difficulty or slow deflation of lots of SELUTION SLRTM PTA mm X mm balloon catheters | Read this Urgent Field Safety Notice (Removal) letter. Immediately check your inventory to confirm that you do not have any units from the affected lots in your possession. Identify and set aside any units from the identified lots in a manner that ensures the affected product will not be used. Check all storage and usage locations. Review, complete, sign and return the enclosed Acknowledgement Form directly to Cordis at the fax number on the form or email to: GMB-CordisFieldAction@cordis.com Return any affected product to the address listed on the form, with reference to your Customer Number which is listed on the form. Share this letter with others in your facility who need to be made aware of this removal and with any other facility that may have been sent the affected units of product from your facility. If any units of the affected lots are found to be at the other facility, please arrange the return of the units. Maintain awareness of this notice until all affected product has been returned to Cordis and keep a copy of this notice with the affected product. | 2024-11-06 10:18
Medtronic Inc. | External Drainage and Monitoring Kit | 44 | - | Medtronic has received customer complaints reporting cracks and leaks on the stopcock of Becker and Exacta EDMS devices Complaint analysis shows a higher complaint rate associated with the p n Four Way Stopcock used in Becker and Exacta EDMS devices | Advice on actions to be taken by the user: Recommended Mitigations*: • Prior to use, inspect all stopcocks and connections to ensure that connections are secure and that there are no visible cracks in the stopcocks. If cracks or leaks are identified, do not use the device and return it to Medtronic. • The system must be pre-filled with sterile isotonic saline solution prior to connecting to the patient. • Check all connections to ensure that fittings (connections) are tight and leak-free. • All connections should be finger tightened. Over tightening can cause cracks and leaks to occur. * The mitigations described are consistent with the instructions for use (IFU) for Becker and Exacta. Please refer to the IFU for full details. Patient Management Recommendations: • If the system develops cracks or leaks during use, replace it using sterile technique and return the damaged system to Medtronic. If the leak occurs with the Patient Line Stopcock (see Figure 1), consider application of a hemostat or other clamp to the proximal patient line while coordinating a replacement if occlusion of the patient line does not create a risk to the patient. • As stated in the product labeling, all patients with EDMS should be monitored for evidence of infection. If a system is found to have cracks or leaks, the system should be replaced, and the patient should continue to be monitored for evidence of infection. Required Actions: Please complete and return the enclosed Customer Acknowledgment Form to acknowledge that you have read and understand this letter. Please post a copy of this notification near impacted product as a reminder of the issue and recommended mitigating actions. Retain a copy of this letter for your records. | 2024-11-06 15:55
Siemens Healthcare Diagnostics Inc. | Atellica CH B2M | 5600 | - | Issue Description x D Siemens Healthineers has confirmed the potential for falsely depressed auto diluted results for samples above the measuring interval of mg L up to mg L when using the Atellica CH Microglobulin B M reagent Results within the measuring interval mg L are not impacted x D This issue applies to all in date and future Atellica CH B M reagent lots for use on Atellica CH and Atellica CI Analyzers The Instructions for Use IFU will be updated to support manual dilution only with saline for samples mg L x D Kindly review the attached Urgent Field Safety Notice | Customer Actions • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Manually dilute patient samples that are > 18.00 mg/L: 1. Navigate to the CH Test Definition screen. 2. Under Measuring Interval, uncheck “Repeat when Outside Measuring Interval.” 3. For samples that generate test results > 18.00 mg/L, prepare a 1:2 manual dilution with saline solution (0.9% unbuffered Isotonic saline). 4. From the Create Patient Order screen, order the sample and enter “2” in the manual dilution field on the left of the screen. Note: The system will apply the 1:2 dilution factor and calculate the final result. 5. Process the manually diluted sample. 6. Report results greater than 36.00 mg/L as “>36.00 mg/L.” • Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. • Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached Urgent Field Safety Notice | 2024-11-06 16:59
GE Healthcare Japan Corporation | Revolution Ascend | 1 | - | GE HealthCare has become aware of a potential issue on certain Discovery Optima Revolution series CT systems and certain operator console upgrades that can result in rotated CT images Particularly with helical cine and cardiac if applicable scan modes after the first image of an exam each subsequent image can be rotated by an increasing amount The amount of progressive image rotation is dependent on the gantry speed and the duration of the scan The final image in the exam can be rotated up to a maximum of degrees The entire anatomy rotates as a whole by equal degrees without distortion and the anatomical positioning and z axis of the images remains correct Small degrees of image rotation might not be noticed If image rotation is noticed and necessitates a rescan of the patient the same image rotation will reoccur x D x D NOTE This issue does not impact scout or axial scan modes x D x D There have been no injuries reported as a result of this issue x D | You can continue to use your system. When utilizing helical, cine, and cardiac scan modes, review the image data set for any potential rotation. When using these modes during interventional procedures or radiation therapy planning, adhere to standard clinical quality checks and workflows. If image rotation is detected, consult with the physician to determine whether a rescan is necessary or if the images need to be corrected during post-processing to ensure accurate interpretation and planning. Please ensure all potential users in your facility are made aware of this safety notification and the recommended actions. GE HealthCare will correct all affected products at no cost to you. A GE HealthCare representative will contact you to arrange for the correction. | 2024-11-07 12:26
D.O.R.C Dutch Ophthalmic Research Center (International) B.V | Silicone Oil | - | - | It has been concluded that elevated infusion pressure may be one of the potential causes | we direct that the maximum silicone oil infusion pressure be lowered to 4 bar until further notice. Review of the investigation has determined that reducing the pressure used during injection will result in reducing the risk of syringe breakage. | 2024-11-07 14:20
Siemens Healthcare Diagnostics Inc. | Atellica IM EPO | 600 | - | Issue Description x D Siemens Healthineers is communicating the resolution to Urgent Field Safety Notice UFSN AIMC A OUS Atellica IM and ADVIA Centaur Erythropoietin EPO Negative Bias versus WHO Standardization released on Jan regarding the negative bias as compared to the rd World Health Organization WHO International Standard NIBSC code x D Siemens Healthineers has restored alignment to the WHO Standard and verified the reference interval as claimed in the Instructions for Use IFU starting with the products listed in the table below x D Customers using Atellica IM and ADVIA Centaur EPO WHO realigned assay must ensure use of WHO realigned Quality Control and Master Curve Material products as specified in the Appendix Table x D See Appendix for Additional Data in Tables x D x D Kindly review the attached Follow up Urgent Field Safety Notice | Customer Actions • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Perform the instructions provided below: o Users may evaluate and implement the restored Atellica IM and ADVIA Centaur EPO Assay kit lots. o Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. o Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached Follow-up Urgent Field Safety Notice | 2024-11-07 17:25
DiaSorin Inc. | LIAISON Q.S.E.T. Device Plus | 40 | - | Internal investigation has determined that of the tubes in the lots identified may have loose clear caps The loose clear caps may allow leakage of the buffer from the tube | Devices should be inspected for loose clear caps prior to use. The user should confirm that the clear cap is tight by grasping the cap and twisting. If the cap is loose, the device should be discarded. Do not attempt to use the device as the loose clear cap may have allowed the buffer to leak and the buffer volume may be insufficient | 2024-11-10 15:16
DiaSorin Inc. | LIAISON Q.S.E.T. Device Plus | 1 | - | Internal investigation has determined that of the tubes in the lots identified may have loose clear caps The loose clear caps may allow leakage of the buffer from the tube | Devices should be inspected for loose clear caps prior to use. The user should confirm that the clear cap is tight by grasping the cap and twisting. If the cap is loose, the device should be discarded. Do not attempt to use the device as the loose clear cap may have allowed the buffer to leak and the buffer volume may be insufficient | 2024-11-10 15:24
Boston Scientific | ACURATE neo2™ Aortic Valve System | 455 | - | Boston Scientific has become aware of new information related to valve under expansion that emerged from review of the year clinical trial data from the ACURATE IDE x D The Main Randomized Cohort represents randomization of ACURATE neo as the test article versus the Control which is the commercially available SAPIEN Edwards or Evolut Medtronic At year the ACURATE IDE missed its primary endpoint ACURATE neo non inferiority to the Control group for the composite of death stroke and rehospitalization | As ACURATE neo2 is a product with mandatory physician training, Boston Scientific is also updating the ACURATE neo2 global physician training program for the risk of valve under expansion and the practices that may reduce this risk. This is consistent with ACURATE PRIME training already available and deployed to new ACURATE Prime users. | 2024-11-11 15:55
Quantel Medical | ME0000010185SFDAA00004 | 8 | - | One case of overdosing power during ThermoCyclo treatment Continuous mode was reported and after investigation shows that the device s internal power monitoring was involved resulting in a delivered power which was over the nominal value without triggering a power control safety error which should occurs when internal power measurement is over the limit | software update required | 2024-11-12 12:00
Philips Medical Systems Nederland B.V. | Allura Xper FD10 | 35 | - | Philips has identified potential safety issues with delayed LTE kit installations for Allura R R systems exceeding years of life The kit includes a propeller motor clamping bolt see Figure and the X ray tube locking bolt see Figure If one of these bolts breaks part of the system will not function as expected x D x D If the installation of the LTE kit is delayed breakage of the propeller motor clamping bolt and or the X ray tube locking bolt may be expected If this occurs x D x D The broken propeller motor clamping bolt may cause erratic C arm movements and potentially colliding with patients or bystanders see Figure These unbalanced movements may also cause the system to falsely detect a collision consequently blocking motorized C arm movements x D The broken X ray tube locking bolt may cause the tube to hit the cover causing unexpected noise and the appearance of a black area maximum up to on the X ray image Consequently the area of interest in the X ray image may be obscured leading to the need for an additional scan | Philips will contact all affected customers to arrange for a Field Service Engineer visit to install a Lifetime Extension Kit (reference FCO72200598). As of the date of this Urgent Field Safety Notice, Philips expects the solution to be available by Q1 2025. | 2024-11-13 17:17
Philips Medical Systems Nederland B.V. | Allura Xper FD10 OR Table | 35 | - | Philips has identified potential safety issues with delayed LTE kit installations for Allura R R systems exceeding years of life The kit includes a propeller motor clamping bolt see Figure and the X ray tube locking bolt see Figure If one of these bolts breaks part of the system will not function as expected x D x D If the installation of the LTE kit is delayed breakage of the propeller motor clamping bolt and or the X ray tube locking bolt may be expected If this occurs x D x D The broken propeller motor clamping bolt may cause erratic C arm movements and potentially colliding with patients or bystanders see Figure These unbalanced movements may also cause the system to falsely detect a collision consequently blocking motorized C arm movements x D The broken X ray tube locking bolt may cause the tube to hit the cover causing unexpected noise and the appearance of a black area maximum up to on the X ray image Consequently the area of interest in the X ray image may be obscured leading to the need for an additional scan | Philips will contact all affected customers to arrange for a Field Service Engineer visit to install a Lifetime Extension Kit (reference FCO72200598). As of the date of this Urgent Field Safety Notice, Philips expects the solution to be available by Q1 2025. | 2024-11-13 17:17
Philips Medical Systems Nederland B.V. | Allura Xper FD10/10 | 35 | - | Philips has identified potential safety issues with delayed LTE kit installations for Allura R R systems exceeding years of life The kit includes a propeller motor clamping bolt see Figure and the X ray tube locking bolt see Figure If one of these bolts breaks part of the system will not function as expected x D x D If the installation of the LTE kit is delayed breakage of the propeller motor clamping bolt and or the X ray tube locking bolt may be expected If this occurs x D x D The broken propeller motor clamping bolt may cause erratic C arm movements and potentially colliding with patients or bystanders see Figure These unbalanced movements may also cause the system to falsely detect a collision consequently blocking motorized C arm movements x D The broken X ray tube locking bolt may cause the tube to hit the cover causing unexpected noise and the appearance of a black area maximum up to on the X ray image Consequently the area of interest in the X ray image may be obscured leading to the need for an additional scan | Philips will contact all affected customers to arrange for a Field Service Engineer visit to install a Lifetime Extension Kit (reference FCO72200598). As of the date of this Urgent Field Safety Notice, Philips expects the solution to be available by Q1 2025. | 2024-11-13 17:17
Philips Medical Systems Nederland B.V. | Allura Xper FD20 | 35 | - | Philips has identified potential safety issues with delayed LTE kit installations for Allura R R systems exceeding years of life The kit includes a propeller motor clamping bolt see Figure and the X ray tube locking bolt see Figure If one of these bolts breaks part of the system will not function as expected x D x D If the installation of the LTE kit is delayed breakage of the propeller motor clamping bolt and or the X ray tube locking bolt may be expected If this occurs x D x D The broken propeller motor clamping bolt may cause erratic C arm movements and potentially colliding with patients or bystanders see Figure These unbalanced movements may also cause the system to falsely detect a collision consequently blocking motorized C arm movements x D The broken X ray tube locking bolt may cause the tube to hit the cover causing unexpected noise and the appearance of a black area maximum up to on the X ray image Consequently the area of interest in the X ray image may be obscured leading to the need for an additional scan | Philips will contact all affected customers to arrange for a Field Service Engineer visit to install a Lifetime Extension Kit (reference FCO72200598). As of the date of this Urgent Field Safety Notice, Philips expects the solution to be available by Q1 2025. | 2024-11-13 17:17
Philips Medical Systems Nederland B.V. | Allura Xper FD20 OR Table | 35 | - | Philips has identified potential safety issues with delayed LTE kit installations for Allura R R systems exceeding years of life The kit includes a propeller motor clamping bolt see Figure and the X ray tube locking bolt see Figure If one of these bolts breaks part of the system will not function as expected x D x D If the installation of the LTE kit is delayed breakage of the propeller motor clamping bolt and or the X ray tube locking bolt may be expected If this occurs x D x D The broken propeller motor clamping bolt may cause erratic C arm movements and potentially colliding with patients or bystanders see Figure These unbalanced movements may also cause the system to falsely detect a collision consequently blocking motorized C arm movements x D The broken X ray tube locking bolt may cause the tube to hit the cover causing unexpected noise and the appearance of a black area maximum up to on the X ray image Consequently the area of interest in the X ray image may be obscured leading to the need for an additional scan | Philips will contact all affected customers to arrange for a Field Service Engineer visit to install a Lifetime Extension Kit (reference FCO72200598). As of the date of this Urgent Field Safety Notice, Philips expects the solution to be available by Q1 2025. | 2024-11-13 17:17
Philips | Allura Xper FD20 Biplane | 35 | - | Philips has identified potential safety issues with delayed LTE kit installations for Allura R R systems exceeding years of life The kit includes a propeller motor clamping bolt see Figure and the X ray tube locking bolt see Figure If one of these bolts breaks part of the system will not function as expected x D x D If the installation of the LTE kit is delayed breakage of the propeller motor clamping bolt and or the X ray tube locking bolt may be expected If this occurs x D x D The broken propeller motor clamping bolt may cause erratic C arm movements and potentially colliding with patients or bystanders see Figure These unbalanced movements may also cause the system to falsely detect a collision consequently blocking motorized C arm movements x D The broken X ray tube locking bolt may cause the tube to hit the cover causing unexpected noise and the appearance of a black area maximum up to on the X ray image Consequently the area of interest in the X ray image may be obscured leading to the need for an additional scan | Philips will contact all affected customers to arrange for a Field Service Engineer visit to install a Lifetime Extension Kit (reference FCO72200598). As of the date of this Urgent Field Safety Notice, Philips expects the solution to be available by Q1 2025. | 2024-11-13 17:17
Philips | Allura Xper FD20 Biplane OR Table | 35 | - | Philips has identified potential safety issues with delayed LTE kit installations for Allura R R systems exceeding years of life The kit includes a propeller motor clamping bolt see Figure and the X ray tube locking bolt see Figure If one of these bolts breaks part of the system will not function as expected x D x D If the installation of the LTE kit is delayed breakage of the propeller motor clamping bolt and or the X ray tube locking bolt may be expected If this occurs x D x D The broken propeller motor clamping bolt may cause erratic C arm movements and potentially colliding with patients or bystanders see Figure These unbalanced movements may also cause the system to falsely detect a collision consequently blocking motorized C arm movements x D The broken X ray tube locking bolt may cause the tube to hit the cover causing unexpected noise and the appearance of a black area maximum up to on the X ray image Consequently the area of interest in the X ray image may be obscured leading to the need for an additional scan | Philips will contact all affected customers to arrange for a Field Service Engineer visit to install a Lifetime Extension Kit (reference FCO72200598). As of the date of this Urgent Field Safety Notice, Philips expects the solution to be available by Q1 2025. | 2024-11-13 17:17
.Siemens Healthcare Diagnostics Inc | Atellica® CI 1900 Analyzer | 20 | - | Issue Description x D Siemens Healthineers has confirmed through the internal investigation of customer complaints that the IMT Diluent volume remaining remaining does not decrease as expected on the Atellica CI Analyzer The IMT module runs a Standard A refresh sequence before processing samples if at least two minutes of inactivity have passed IMT Diluent is used by this sequence however the calculated diluent volume is not accurate potentially leading to the IMT Diluent being empty while still displaying that volume is remaining In this case Sodium Na Potassium K and Chloride Cl test results may be falsely elevated Quality Control materials demonstrate the same behavior Refer to Appendix x D Several factors influence the probability of occurrence including the sample process workflow and the number of times the IMT module is inactive for greater than two minutes between samples x D x D Kindly review the attached Urgent Field Safety Notice | Customer Actions Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. Perform the instructions provided below to adjust the IMT Fluid Volume threshold: 1. Login to Atellica CI instrument with Lab Manager credentials. 2. Change the IMT fluid volume alert threshold %. a. From the Home screen of the user interface (UI) click on System Navigator. b. On the System Navigator screen choose Supply Thresholds. c. Click on the IMT Fluid Volume Threshold (%) dropdown arrow and select 50%. d. Select Save. e. Check the Supply Needs on the home screen. 3. Replace the appropriate IMT Supply fluid when the threshold is below 50% as indicated (or prompted) by the software. Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached Urgent Field Safety Notice | 2024-11-13 17:26
Siemens Healthcare Diagnostics Inc. | Atellica IM T4 | 622 | - | Issue Description x D Siemens Healthineers has confirmed through the investigation of customer complaints the Atellica IM Total Thyroxine T assay does not meet the Pack Calibration Interval of days as stated in the Instructions for Use x D When valid calibration is achieved and Quality Control QC meet defined ranges patient results are accurate and acceptable for reporting x D Please follow the Customer Actions section to reduce the Pack Calibration Interval in combination with standard QC practice to mitigate the issue x D Siemens Healthineers is currently investigating the root cause of this issue x D x D Kindly review the attached Customer Notification | Customer Actions For the products listed in the table above, perform the actions provided in the attached Addendum to reduce the Pack Calibration Interval from 21 days to 4 days. Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached Customer Notification | 2024-11-14 07:41
Siemens Healthcare Diagnostics Inc. | Atellica IM T4 +A | 622 | - | Issue Description x D Siemens Healthineers has confirmed through the investigation of customer complaints the Atellica IM Total Thyroxine T assay does not meet the Pack Calibration Interval of days as stated in the Instructions for Use x D When valid calibration is achieved and Quality Control QC meet defined ranges patient results are accurate and acceptable for reporting x D Please follow the Customer Actions section to reduce the Pack Calibration Interval in combination with standard QC practice to mitigate the issue x D Siemens Healthineers is currently investigating the root cause of this issue x D x D Kindly review the attached Customer Notification | Customer Actions For the products listed in the table above, perform the actions provided in the attached Addendum to reduce the Pack Calibration Interval from 21 days to 4 days. Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached Customer Notification | 2024-11-14 07:41
Diagnostic Grifols, S.A | Erytra Eflexis | 35 | - | The functionality Duplicate the Recipient Donor profile assignment feature is accessible from within the Donor assignment button This manual feature allows the operator to duplicate the recipient donor s assignment for the purpose of re assigning the same donor profiles to a different recipient Under the following specific conditions the potential exists for the crossmatch results to be assigned to the incorrect recipient x D A crossmatch test has been previously executed for a recipient and at least one donor x D The Duplicate the Recipient Donor Profile Assignments button was used to repopulate the crossmatch programming window and change the recipient information x D The crossmatch test result s for the previous recipient was were in Pending Validation Validated or Exported status at the time the crossmatch test s was were ordered for the new recipient x D When these specific conditions are met the crossmatch test ordered for the latest recipient donor s pairing will not be queued Instead the results from that first recipient donor will be removed and assigned to the latest recipient donor s pairing | install version 3.2.0 | 2024-11-14 09:59
Beckman Coulter Inc | Access hsTnI (reagent pack) | 377 | - | In remote circumstances if a sample containing high cardiac troponin cTnI pg mL is tested using an Access assay other than Access hsTnI cTnI carryover may occur if the next test performed using the same probe on that instrument is Access hsTnI on a different sample |  Follow your established laboratory protocols for analyzing and retesting discrepant samples if an observed Access hsTnI test result does not align with the patient’s clinical presentation.  Beckman Coulter recommends reviewing the content of this letter with your laboratory and/or medical director to determine appropriate next steps. | 2024-11-14 23:37
MedSource | 150 ML infusion set | 44790 | - | impurities inside the infusion set | replacement of the affected lots | 2024-11-17 11:15
Philips Medical Systems Nederland B.V. | Allura Xper FD10 | 4 | - | Philips has identified that due to the forces applied during the movement of the C Arc of the Azurion and Allura systems the bolts supporting the FlexMove Carriage may become loose and or break and cracks may appear in the FlexMove Carriage see Figures and x D x D If FlexMove rail fixation bolts become loose or broken or cracks appear in the carriage the following issues may occur x D x D Transversal movements of the C Arc stop due to false collision detections because of additional friction x D Manual movement of the FlexMove Carriage is not possible due to added friction x D Abnormal noise during transversal movement of the C Arc x D Unstable C Arc suspension x D Fall of the C Arc Assembly kg if all bolts in the X axis are broken loose x D Drop of the C Arc Assembly up to cm if C Arc is at one side of the rail up to cm if C Arc is in the center of the rails and cm if the C Arc is in the Anterior Posterior position if all bolts in the Y axis are broken loose x D x D As of May Philips has received fourteen complaints related to eleven systems reporting loose and or broken bolts In three cases cracks were also identified In none of these cases the C Arc Assembly fell or dropped No harm to patients or bystanders was reported | Philips expects to release the permanent solution by December 2024 (reference FCO72200581). Philips will contact you to schedule a visit to install this permanent solution. Philips will schedule these visits ensuring that the permanent solution is implemented within 3 years after the initial inspectional activities or after installation. It is estimated that installation of this solution will require 3 working days. We kindly ask your collaboration in this scheduling. In case your system will not be corrected with the permanent technical solution within 3 years following the first inspection/installation, Philips will contact you to schedule an additional inspection(s) of the system(s). | 2024-11-18 16:28
Philips Medical Systems Nederland B.V. | Allura Xper FD20 | 4 | - | Philips has identified that due to the forces applied during the movement of the C Arc of the Azurion and Allura systems the bolts supporting the FlexMove Carriage may become loose and or break and cracks may appear in the FlexMove Carriage see Figures and x D x D If FlexMove rail fixation bolts become loose or broken or cracks appear in the carriage the following issues may occur x D x D Transversal movements of the C Arc stop due to false collision detections because of additional friction x D Manual movement of the FlexMove Carriage is not possible due to added friction x D Abnormal noise during transversal movement of the C Arc x D Unstable C Arc suspension x D Fall of the C Arc Assembly kg if all bolts in the X axis are broken loose x D Drop of the C Arc Assembly up to cm if C Arc is at one side of the rail up to cm if C Arc is in the center of the rails and cm if the C Arc is in the Anterior Posterior position if all bolts in the Y axis are broken loose x D x D As of May Philips has received fourteen complaints related to eleven systems reporting loose and or broken bolts In three cases cracks were also identified In none of these cases the C Arc Assembly fell or dropped No harm to patients or bystanders was reported | Philips expects to release the permanent solution by December 2024 (reference FCO72200581). Philips will contact you to schedule a visit to install this permanent solution. Philips will schedule these visits ensuring that the permanent solution is implemented within 3 years after the initial inspectional activities or after installation. It is estimated that installation of this solution will require 3 working days. We kindly ask your collaboration in this scheduling. In case your system will not be corrected with the permanent technical solution within 3 years following the first inspection/installation, Philips will contact you to schedule an additional inspection(s) of the system(s). | 2024-11-18 16:28
Philips Medical Systems Nederland B.V. | Allura Xper FD10 OR Table | 4 | - | Philips has identified that due to the forces applied during the movement of the C Arc of the Azurion and Allura systems the bolts supporting the FlexMove Carriage may become loose and or break and cracks may appear in the FlexMove Carriage see Figures and x D x D If FlexMove rail fixation bolts become loose or broken or cracks appear in the carriage the following issues may occur x D x D Transversal movements of the C Arc stop due to false collision detections because of additional friction x D Manual movement of the FlexMove Carriage is not possible due to added friction x D Abnormal noise during transversal movement of the C Arc x D Unstable C Arc suspension x D Fall of the C Arc Assembly kg if all bolts in the X axis are broken loose x D Drop of the C Arc Assembly up to cm if C Arc is at one side of the rail up to cm if C Arc is in the center of the rails and cm if the C Arc is in the Anterior Posterior position if all bolts in the Y axis are broken loose x D x D As of May Philips has received fourteen complaints related to eleven systems reporting loose and or broken bolts In three cases cracks were also identified In none of these cases the C Arc Assembly fell or dropped No harm to patients or bystanders was reported | Philips expects to release the permanent solution by December 2024 (reference FCO72200581). Philips will contact you to schedule a visit to install this permanent solution. Philips will schedule these visits ensuring that the permanent solution is implemented within 3 years after the initial inspectional activities or after installation. It is estimated that installation of this solution will require 3 working days. We kindly ask your collaboration in this scheduling. In case your system will not be corrected with the permanent technical solution within 3 years following the first inspection/installation, Philips will contact you to schedule an additional inspection(s) of the system(s). | 2024-11-18 16:28
Philips Medical Systems Nederland B.V. | Allura Xper FD20 OR Table | 4 | - | Philips has identified that due to the forces applied during the movement of the C Arc of the Azurion and Allura systems the bolts supporting the FlexMove Carriage may become loose and or break and cracks may appear in the FlexMove Carriage see Figures and x D x D If FlexMove rail fixation bolts become loose or broken or cracks appear in the carriage the following issues may occur x D x D Transversal movements of the C Arc stop due to false collision detections because of additional friction x D Manual movement of the FlexMove Carriage is not possible due to added friction x D Abnormal noise during transversal movement of the C Arc x D Unstable C Arc suspension x D Fall of the C Arc Assembly kg if all bolts in the X axis are broken loose x D Drop of the C Arc Assembly up to cm if C Arc is at one side of the rail up to cm if C Arc is in the center of the rails and cm if the C Arc is in the Anterior Posterior position if all bolts in the Y axis are broken loose x D x D As of May Philips has received fourteen complaints related to eleven systems reporting loose and or broken bolts In three cases cracks were also identified In none of these cases the C Arc Assembly fell or dropped No harm to patients or bystanders was reported | Philips expects to release the permanent solution by December 2024 (reference FCO72200581). Philips will contact you to schedule a visit to install this permanent solution. Philips will schedule these visits ensuring that the permanent solution is implemented within 3 years after the initial inspectional activities or after installation. It is estimated that installation of this solution will require 3 working days. We kindly ask your collaboration in this scheduling. In case your system will not be corrected with the permanent technical solution within 3 years following the first inspection/installation, Philips will contact you to schedule an additional inspection(s) of the system(s). | 2024-11-18 16:28
Philips Medical Systems Nederland B.V. | AZURION 7 M20 | 4 | - | Philips has identified that due to the forces applied during the movement of the C Arc of the Azurion and Allura systems the bolts supporting the FlexMove Carriage may become loose and or break and cracks may appear in the FlexMove Carriage see Figures and x D x D If FlexMove rail fixation bolts become loose or broken or cracks appear in the carriage the following issues may occur x D x D Transversal movements of the C Arc stop due to false collision detections because of additional friction x D Manual movement of the FlexMove Carriage is not possible due to added friction x D Abnormal noise during transversal movement of the C Arc x D Unstable C Arc suspension x D Fall of the C Arc Assembly kg if all bolts in the X axis are broken loose x D Drop of the C Arc Assembly up to cm if C Arc is at one side of the rail up to cm if C Arc is in the center of the rails and cm if the C Arc is in the Anterior Posterior position if all bolts in the Y axis are broken loose x D x D As of May Philips has received fourteen complaints related to eleven systems reporting loose and or broken bolts In three cases cracks were also identified In none of these cases the C Arc Assembly fell or dropped No harm to patients or bystanders was reported | Philips expects to release the permanent solution by December 2024 (reference FCO72200581). Philips will contact you to schedule a visit to install this permanent solution. Philips will schedule these visits ensuring that the permanent solution is implemented within 3 years after the initial inspectional activities or after installation. It is estimated that installation of this solution will require 3 working days. We kindly ask your collaboration in this scheduling. In case your system will not be corrected with the permanent technical solution within 3 years following the first inspection/installation, Philips will contact you to schedule an additional inspection(s) of the system(s). | 2024-11-18 16:28
Randox Laboratories Ltd | Evidence Multistat | 107 | - | After importing a batch and running a calibration error appears It was theoretically possible to install the electronic kit supplied with a batch and after a system restart the files were missing to perform kit calibration Error would be presented during calibration This issue has been fixed | • Review the software release notes provided for further information. • Please complete the mandatory software upgrade. Please contact technical.services@randox.com or support@randoxtoxicology.com or support@randoxfood.com for the OneDrive link to the software update files. • Complete and return the response form 12187-QA technical.services@randox.com or support@randoxtoxicology.com or support@randoxfood.com within five working days. Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. | 2024-11-20 10:36
Draeger Medical, Inc. | Ponta C/H/E/S | 314 | - | two different error patterns have been identified that may also occur in combination due to incorrect installation | Our customers will be informed of the issue in a customer letter The Ponta must be checked for both error patterns via a service appointment within the prescribed 2-year maintenance interval (in accordance with the instructions for use). | 2024-11-21 13:53
ICU MEDICAL | ParaPAC plus™ | 38 | - | Smiths Medical became aware of an issue related to a potential for the patient outlet connector to loosen detach from the paraPAC Plus P and P ventilators impacting the active ventilation function | Actions to be taken by the User/Customer: 1. Please identify all affected paraPAC plusTM units in your possession. 2. Perform an inspection to determine if your devices are affected, per the instructions below: a. First perform a visual inspection to determine if the outlet connector is disconnected. b. Second evaluate the connector physically to determine if the outlet connector is loose or moves when placing a patient circuit on the connector or when removing it. c. If outlet connector remains tight after physical inspection, you can continue use of the device with heightened awareness and following all pre-use checks as per the user manual. d. If the outlet connector moves or feels loose, the device must be removed from use and repaired by Smiths Medical. Report the event to Global Complaint Management at globalcomplaints@icumed.com. . 3. Every use thereafter of every device, pre-use checks must be completed as described in the user manual and extra caution must be taken in inspecting the outlet connector prior to use and placing the patient circuit on the connector and during removal. a. When using the device, all instructions, including warnings and cautions in the User Manual must be followed with heightened awareness. b. This is inclusive, but not limited to the following: • All pre-use checks must be performed before each use. • Constant monitoring of the patient • Blood oxygenation and end tidal carbon dioxide levels should be monitored independently using pulse oximetry and capnography. • Alternative means of ventilation such as bag mask ventilation, must be available in the event of ventilator failure or malfunction. c. The paraPAC plusTM is also equipped with the following design mitigations & alarms: • Low Pressure/Disconnect Alarm: Disconnection/Detachment of the Patient Outlet connector would trigger low pressure/disconnect alarm. • Pressure Monitor: Disconnection/Detachment of the Patient Outlet connector would be indicated by no movement of the manometer needle 4. Share this FSN notification with all potential users of the devices to ensure they are aware of this issue and proposed mitigations. If the devices are used at another location, please ensure that this communication is delivered to these locations. 5. Complete and return the attached Customer Response Form to EMEA-FSN@icumed.com within ten days of receipt to acknowledge your understanding of this notification. | 2024-11-21 15:23
ICU MEDICAL | ParaPAC plus™ | 38 | - | Smiths Medical became aware of an issue involving a paraPAC Plus P and P ventilator where there is potential for inadvertent tidal volume knob movement from the original setting when set at high mL and low mL settings | Actions to be taken by the User/Customer: 1. Please identify all affected paraPAC plusTM units in your possession. 2. Perform an inspection to determine if your devices are affected, per the instructions below: a. Rotate the tidal volume knob to the lowest position of the control and evaluate whether or not the knob stays in position as set or if it moves to a higher position. b. Repeat this step, rotate tidal volume knob to the highest position of the control and evaluate whether the knob stays in position as set or if it moves to a lower position. c. If the knob stays in the position as set when tested, then your product is not affected, and you can continue use of the device as normal. Complete the attached Response Form confirming that you have no affected product. d. If the knob moves or changes position from the set position, then your product is affected, and it will need to be removed from use and repaired. Do not attempt to use or repair the affected product. Report the event to Global Complaint Management at globalcomplaints@icumed.com. Complete and return the attached Response Form. . 3. Share this FSN notification with all potential users of the devices to ensure they are aware of this issue and proposed mitigations. If the devices are used at another location, please ensure that this communication is delivered to these locations. 4. Complete and return the attached Customer Response Form to EMEA-FSN@icumed.com within ten days of receipt to acknowledge your understanding of this notification. | 2024-11-21 16:53
Fritz Stephan GmbH | SOPHIE | 1 | - | Two cases were reported to Fritz Stephan GmbH in which a loss of mean airway pressure occurred under the x D combined ventilation mode HFO IMV x D x D This loss of mean airway pressure can be triggered by an extensive loss of pressure in the breathing circuit e g x D due to a loss of gas supply disconnection of the breathing circuit or an open emergency valve | You can continue to use the Sophie device, but you should avoid HFO-IMV mode if possible. If you absolutely want to use HFO-IMV mode, you should check the measured value display for peep and mean pressure immediately after each technical alarm. If a lowered peep and/or mean pressure has occurred, you can stabilize the pressure by temporarily switching from HFO-IMV to HFO-CPAP and back. To avoid an undesired drop in ventilation pressure under HFO-IMV, you should never disconnect the device from gas supply. Need to install software configuration PC V2.9.4 | 2024-11-23 11:16
Immunotech S.A.S a Beckman Coulter Company | Access hsTnI | 377 | - | In remote circumstances if a sample containing high cardiac x D troponin cTnI pg mL is tested using an Access assay x D other than Access hsTnI cTnI carryover may occur if the next test x D performed using the same probe on that instrument is Access x D hsTnI on a different sample x D | Beckman Coulter is currently investigating options to resolve this issue. | 2024-11-24 10:14
Becton Dickinson and Company 7 Loveton Circle , Sparks | BD FACSLyric™ Flow Cytometer | 15 | - | BD has identified through complaints an increase in field failures replacements of the power supply module on BD FACSLyric flow cytometers The capacitor on the power supply can split which has the potential for the instruments to fail to power on and or stay on x D x D | Clinical User Actions 1. Continue normal operation of your BD FACSLyric™ Flow Cytometer according to the Instructions For Use. 2. The health care team should manage the patient risks according to their institutional policies and procedures. BD Actions: BD has identified the root cause and is taking action to prevent recurrence of this product issue. Actions to be taken by BD: BD Field Service Engineers will replace the power supply module of affected BD FACSLyric™ Flow Cytometers. Customer Actions: • Review the information in Table 1 to determine if instruments in your possession are impacted. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 18th December 2024. • Circulate this notice to all those who need to be aware within your organisation or to any organisation where the potentially affected product has been transferred. • If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: • Review the information in Table 1 and determine if instruments in your possession are impacted. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 18th December 2024. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. | 2024-11-24 14:03
Abbott | e AVEIR™ DR Leadless Pacemaker | 26 | - | Abbott is informing customers of a programmer software anomaly that may be encountered in a very specific x D circumstance during finalization of the AVEIR DR Leadless Pacemaker LP System This can occur during an x D upgrade from an AVEIR VR to an AVEIR DR system or during a new AVEIR DR system implant A specific sequence of x D programmer actions combined with a loss of telemetry during a small window seconds may cause the finalization x D step to fail If this occurs the Merlin PCS Programmer displays a Loss of telemetry detected window and x D there is no pacing Atrial sense markers coincident to intrinsic QRS complexes may be observed Subsequent attempts x D to finalize the device will also result in a Loss of telemetry detected window Full functionality including ability to x D finalize the device and restore pacing is reinstated through Abbott Technical Services intervention x D Abbott has observed this issue at a rate of approximately for implanted AVEIR DR systems globally There have x D been no reports of any harm to patients resulting from this issue In each event devices were subsequently restored to x D standard functionality | Patient Management Recommendations: Updated Merlin™ PCS 3650 programmer software (version 28.3.2 rev 1 or higher) corrects this issue. Your Abbott Representative will upgrade programmer software beginning in January 2025. In consultation with Abbott CRM’s Medical Advisory Board (MAB), Abbott is providing the following recommendations for implanting AVEIR DR LPs prior to the programmer software upgrade: • Prior to finalization, ensure robust telemetry connection. • For pacemaker dependent patients: o Consider backup pacing support through implant finalization. o During a new AVEIR DR system implant, consider implanting and finalizing the RA LP first before the RV LP if appropriate for the patient, as this order of finalization does not allow the issue to occur. • If this issue occurs during the finalization step of an AVEIR DR system, contact Abbott Technical Services at +4684744147 as device restoration will be required. | 2024-11-24 21:14
Reach Surgical INC. | IREACH Magnum Staplers | 100 | - | One batch of Magnum Staplers may potentially malfunction as staples may not have formed as the B shape perfectly which does not align with internal product inspection specifications | replace the affected items | 2024-11-25 16:53
GS Elektromedizinische Geräte G. Stemple GmbH | Corpuls AED | 20 | - | We are observing in the field a certain number of devices that perform an unplanned restart with self test during operation The cause for this is a loss of contact between the battery and the device When there is pressure on one side of the device the inserted battery is tilted to the maximum and the brief loss of contact leads to a restart of the device with self test | To identify conspicuous corpuls aed batteries, please perform a pressure point test. This test checks the system stability of the device and makes sure that manipulation of the battery withing the housing does not lead to a restart of the device. | 2024-11-26 11:11
Fritz Stephan GmbH | SOPHIE | 1 | - | In the course of market surveillance Firtz Stephan GmbH became aware that unexpected behavior of the Device can occur in the event of large leaks or incorrect flow measurement when using the SIMV ventilation Form | Need to install software configuration PC V4.5.0, MC V3.3.2 | 2024-11-26 11:52
Datex-Ohmeda, Inc. | CARESCAPE R860 | 73 | - | This letter is to inform you that we have identified additional CARESCAPE R Ventilators affected by GE HealthCare s Urgent Medical Device Correction notification Ref A copy of the previous letter Ref is attached for your reference x D These additional systems identified in Appendix did not receive the main system software update and the User Reference Manual URM addendum with updated maintenance instructions | Please perform the instructions described in the “Actions to be taken by Customer” section of the Urgent Medical Device Correction Notice included in this mailing. GE HealthCare will correct all affected products at no cost to you. A GE HealthCare representative will contact you to arrange for the software correction. Please place the enclosed URM addendum with the CARESCAPE R860 product. If additional addendums are needed, please contact GE HealthCare Service or your local Service Representative. | 2024-11-26 12:44
Siemens | PSA IMMULITE 2000 PSA assay | 8 | عيب مصنعي | - | IMMULITE 2000 and IMMULITE 2000 XPI kit lots D168 and above meet the IFU high-dose hook effect claim Siemens continues to investigate root cause of the issue observed with the impacted kit lots | 2024-11-28 14:46
Carl Zeiss Meditec AG | AT 030 | 2 | - | The AT identification label for SL at the AT shows an incorrect manufacturing date | You can paste over the label of the AT 030 with the label attached to this letter. Kindly fill in the “Acknowledgement and Receipt Form” which is attached along with this letter and confirm completion. Return the “Acknowledgement and Receipt Form” to your local ZEISS organization. If you have any questions, please contact your local representative, and ask for support. | 2024-11-28 19:42
Datex-Ohmeda, Inc. | Carestation 620 A1 | 31 | - | GE HealthCare has become aware that Carestation c Anesthesia Delivery Systems with Software Revision SP SP SP or SP will change ventilation settings without user confirmation when the display brightness level is adjusted The ventilation mode parameter settings and the recruitment maneuver settings will change from those set by the user to default settings for the case profile x D This issue will only occur after the system starts up if screen brightness is changed It will not occur on subsequent cases unless the system is rebooted x D A non directed change in ventilation parameters to case profile default settings if unnoticed and not corrected can result in patient injury x D x D There have been no injuries reported as a result of this issue x D x D | The field action will be to do software upgrade on all affected devices. | 2024-12-01 14:12
Datex-Ohmeda, Inc. | Carestation 650 A1 | 31 | - | GE HealthCare has become aware that Carestation c Anesthesia Delivery Systems with Software Revision SP SP SP or SP will change ventilation settings without user confirmation when the display brightness level is adjusted The ventilation mode parameter settings and the recruitment maneuver settings will change from those set by the user to default settings for the case profile x D This issue will only occur after the system starts up if screen brightness is changed It will not occur on subsequent cases unless the system is rebooted x D A non directed change in ventilation parameters to case profile default settings if unnoticed and not corrected can result in patient injury x D x D There have been no injuries reported as a result of this issue x D x D | The field action will be to do software upgrade on all affected devices. | 2024-12-01 14:12
Datex-Ohmeda, Inc. | Tec 6 Plus | 65 | - | GE HealthCare has recently become aware that due to a manufacturing issue certain Tec Plus and Tec series vaporizers can deliver less anesthetic agent than the vaporizer dial setting If the vaporizer is not used in conjunction with an anesthetic agent monitor this can result in insufficient anesthetic effect which could result in patient movement during surgery and or intraoperative awareness x D There have been no injuries reported as a result of this issue | Response Form. Strategy for replacement of vaporizer will be detailed in phase 2. | 2024-12-02 11:51
Datex-Ohmeda, Inc. | Tec 820 | 65 | - | GE HealthCare has recently become aware that due to a manufacturing issue certain Tec Plus and Tec series vaporizers can deliver less anesthetic agent than the vaporizer dial setting If the vaporizer is not used in conjunction with an anesthetic agent monitor this can result in insufficient anesthetic effect which could result in patient movement during surgery and or intraoperative awareness x D There have been no injuries reported as a result of this issue | Response Form. Strategy for replacement of vaporizer will be detailed in phase 2. | 2024-12-02 11:51
Datex-Ohmeda, Inc. | Tec 850 | 65 | - | GE HealthCare has recently become aware that due to a manufacturing issue certain Tec Plus and Tec series vaporizers can deliver less anesthetic agent than the vaporizer dial setting If the vaporizer is not used in conjunction with an anesthetic agent monitor this can result in insufficient anesthetic effect which could result in patient movement during surgery and or intraoperative awareness x D There have been no injuries reported as a result of this issue | Response Form. Strategy for replacement of vaporizer will be detailed in phase 2. | 2024-12-02 11:51
Datex-Ohmeda, Inc. | Tec 6 Plus | 65 | - | GE HealthCare has recently become aware that due to a manufacturing issue certain Tec Plus and Tec series vaporizers can deliver less anesthetic agent than the vaporizer dial setting If the vaporizer is not used in conjunction with an anesthetic agent monitor this can result in insufficient anesthetic effect which could result in patient movement during surgery and or intraoperative awareness x D x D There have been no injuries reported as a result of this issue x D | Customer Letter and Customer Response Form. Strategy for replacement of vaporizer will be detailed in phase 2. | 2024-12-02 12:00
Datex-Ohmeda, Inc. | Tec 820 | 65 | - | GE HealthCare has recently become aware that due to a manufacturing issue certain Tec Plus and Tec series vaporizers can deliver less anesthetic agent than the vaporizer dial setting If the vaporizer is not used in conjunction with an anesthetic agent monitor this can result in insufficient anesthetic effect which could result in patient movement during surgery and or intraoperative awareness x D x D There have been no injuries reported as a result of this issue x D | Customer Letter and Customer Response Form. Strategy for replacement of vaporizer will be detailed in phase 2. | 2024-12-02 12:00
Datex-Ohmeda, Inc. | Tec 850 | 65 | - | GE HealthCare has recently become aware that due to a manufacturing issue certain Tec Plus and Tec series vaporizers can deliver less anesthetic agent than the vaporizer dial setting If the vaporizer is not used in conjunction with an anesthetic agent monitor this can result in insufficient anesthetic effect which could result in patient movement during surgery and or intraoperative awareness x D x D There have been no injuries reported as a result of this issue x D | Customer Letter and Customer Response Form. Strategy for replacement of vaporizer will be detailed in phase 2. | 2024-12-02 12:00
GE Medical Systems, LLC | Optima XR200amx | 5 | - | GE HealthCare has become aware that in certain Optima XR amx systems that have had a column replacement performed some bolts used in the Column Safety mechanism can be missing If these bolts are missing and a failure of the internal counterpoise cable occurs the Column Safety mechanism will not work properly causing the arm with X ray tube and collimator to drop downward on the column track This could result in patient injury during an exam x D x D There have been no injuries reported as a result of this issue x D | GEHC FE will perform the FMI instruction to remove and replace the hardware for the safety pawl springs applying Loctite and appropriate torque and torque mark. FEs will then perform a functional test of the safety pawl. | 2024-12-02 15:51
GE Medical Systems, LLC | Optima XR220amx | 5 | - | GE HealthCare has become aware that in certain Optima XR amx systems that have had a column replacement performed some bolts used in the Column Safety mechanism can be missing If these bolts are missing and a failure of the internal counterpoise cable occurs the Column Safety mechanism will not work properly causing the arm with X ray tube and collimator to drop downward on the column track This could result in patient injury during an exam x D x D There have been no injuries reported as a result of this issue x D | GEHC FE will perform the FMI instruction to remove and replace the hardware for the safety pawl springs applying Loctite and appropriate torque and torque mark. FEs will then perform a functional test of the safety pawl. | 2024-12-02 15:51
GE Medical Systems, LLC | Optima XR240amx | 5 | - | GE HealthCare has become aware that in certain Optima XR amx systems that have had a column replacement performed some bolts used in the Column Safety mechanism can be missing If these bolts are missing and a failure of the internal counterpoise cable occurs the Column Safety mechanism will not work properly causing the arm with X ray tube and collimator to drop downward on the column track This could result in patient injury during an exam x D x D There have been no injuries reported as a result of this issue x D | GEHC FE will perform the FMI instruction to remove and replace the hardware for the safety pawl springs applying Loctite and appropriate torque and torque mark. FEs will then perform a functional test of the safety pawl. | 2024-12-02 15:51
Siemens Healthcare Diagnostics Inc. | Atellica IM T3 | 382 | - | Issue Description x D Siemens Healthineers has confirmed through the investigation of customer complaints the Atellica IM Total Triiodothyronine T assay does not meet the Pack Calibration Interval of days as stated in the Instructions for Use for the kit lot numbers listed in the table below x D Please follow the Customer Actions section to reduce the Pack Calibration Interval in order to mitigate the issue x D When valid calibration is achieved and Quality Control QC meet defined ranges patient results are acceptable for reporting x D Siemens Healthineers is currently investigating the root cause of this issue x D x D Kindly review the attached | Customer Actions For the products listed in the table above, perform the actions provided in the attached Addendum to reduce the Pack Calibration Interval from 21 days to 3 days. Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached | 2024-12-02 17:06
Siemens Healthcare Diagnostics Inc. | Atellica IM T3 +A | 382 | - | Issue Description x D Siemens Healthineers has confirmed through the investigation of customer complaints the Atellica IM Total Triiodothyronine T assay does not meet the Pack Calibration Interval of days as stated in the Instructions for Use for the kit lot numbers listed in the table below x D Please follow the Customer Actions section to reduce the Pack Calibration Interval in order to mitigate the issue x D When valid calibration is achieved and Quality Control QC meet defined ranges patient results are acceptable for reporting x D Siemens Healthineers is currently investigating the root cause of this issue x D x D Kindly review the attached | Customer Actions For the products listed in the table above, perform the actions provided in the attached Addendum to reduce the Pack Calibration Interval from 21 days to 3 days. Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached | 2024-12-02 17:06
Medline Industries, LP | NEURO SPINE | 14 | - | At the recommendation of Integra LifeSciences Medline Industries conducted an assessment on non sterile Codman Surgical Patties and Strips produced by Integra Lifesciences due to potential endotoxin contamination After completing the assessment Medline Industries is issuing a voluntary recall for specific item s and lot s of Custom Convenience Kits which contain Non sterile Codman Surgical Patties Surgical Strips During an internal investigation Integra LifeSciences identified higher than expected levels of endotoxin within the raw material used to produce sterile Codman Surgical Patties and Strips that may have resulted in out of specification levels of endotoxin in those finished goods | 1. Immediately check your stock for the specific kit number and the affected kit lot number (attached list), quarantine any kits with the listed lot numbers. If you don’t have any of the lots listed above skip to step 3. 2. Upon receipt of your response form, over-labels will be provided, which are to be placed on kits which contain the recalled component and will state that the affected component shall be removed and discarded from further use 3. Please fill out and return the enclosed Verification Form via email to MD.Reg@thimar.com Note: By signing and returning the Verification Form you acknowledge that you received and understand this notification. | 2024-12-03 16:15
Medline Industries, LP | TRANSPHENOIDAL KIT | 14 | - | At the recommendation of Integra LifeSciences Medline Industries conducted an assessment on non sterile Codman Surgical Patties and Strips produced by Integra Lifesciences due to potential endotoxin contamination After completing the assessment Medline Industries is issuing a voluntary recall for specific item s and lot s of Custom Convenience Kits which contain Non sterile Codman Surgical Patties Surgical Strips During an internal investigation Integra LifeSciences identified higher than expected levels of endotoxin within the raw material used to produce sterile Codman Surgical Patties and Strips that may have resulted in out of specification levels of endotoxin in those finished goods | 1. Immediately check your stock for the specific kit number and the affected kit lot number (attached list), quarantine any kits with the listed lot numbers. If you don’t have any of the lots listed above skip to step 3. 2. Upon receipt of your response form, over-labels will be provided, which are to be placed on kits which contain the recalled component and will state that the affected component shall be removed and discarded from further use 3. Please fill out and return the enclosed Verification Form via email to MD.Reg@thimar.com Note: By signing and returning the Verification Form you acknowledge that you received and understand this notification. | 2024-12-03 16:15
Medtronic Navigation, Inc. | Pin, Sterile, Percutaneous | 18 | - | Medtronic has become aware that certain percutaneous pin lots see Attachment A Table have been identified as having the potential for an out of round diameter that may render the percutaneous pin unable to fit into the Patient Reference Frame x D or Percutaneous Pin Adapter This issue is associated with specific lots of the Percutaneous Pin used during spinal surgeries | Medtronic requests that you immediately take the following actions: 1. Immediately locate and quarantine all unused impacted product(s). See Attachment A for affected lot numbers and product identification. 2. Return unused impacted product(s) to Medtronic following the instructions in the enclosed Customer Acknowledgment Form. 3. Complete the Customer Acknowledgment Form enclosed with this letter, acknowledging that you have received this information 4. This notice needs to be passed on to all those who need to be aware within your organization or to any organization where the potentially affected devices have been transferred. Please maintain a copy of this notice in your records. | 2024-12-05 11:38
Accuray Incorporated | CyberKnife | 3 | - | Accuray Incorporated Accuray has discovered an issue related to the CyberKnife Standard Treatment Couch STC When an internal mechanism becomes disconnected during clinical use uncontrolled rotation can occur up to degrees in roll or until blocked by the floor or couch column | Accuray is committed to providing our customers and their patients with products that deliver safe and effective radiation treatments. Accuray will proactively provide a full inspection and regardless of condition replace the existing snap rings on the existing retention mechanism of the STC. | 2024-12-07 21:28
Ortho Clinical Diagnostics Inc. | DGXN Slides | 34 | - | A bias of up to will result in a clinically significant negatively biased digoxin result A negatively biased digoxin result may lead to erroneous dose adjustment which could result in a higher than expected serum digoxin concentration Due to digoxin s narrow therapeutic index inappropriate dose adjustments may result in chronic digoxin toxicity Clinical factors x D such as repeat digoxin monitoring routine EKGs and clinical assessment of early signs of toxicity increases the detectability of erroneous results and mitigates serious patient consequences The likelihood of occurrence of severe medical harm is considered remote x D x D QuidelOrtho does not recommend a review of previous results due to the difficulty in identifying patient results which may have been affected by hemoglobin interference and assessing potential patient impact based solely on digoxin test results Discuss any concerns regarding previously reported results with your Laboratory Medical Director to determine the appropriate course of action The results from any diagnostic test should be evaluated in conjunction with a patient s history risk factors clinical presentations signs and symptoms as well as the results of other tests x D x D As of Dec QuidelOrtho has received customer complaint related to this issue with no reports of adverse effects x D | • For customers who have VITROS 5,1 FS/4600/5600/XT 3400/XT 7600 Systems: Load ADD DRV 6308 upon receipt. • For customers who have VITROS 250/350 Systems: Please be aware of the lowering of the Interferent Concentration limit from 300 mg/dL to 100 mg/dL at a digoxin concentration of 2.0 ng/mL (2.6 nmol/L). • Complete and return the enclosed Confirmation of Receipt form no later than Month DD, 2024. • Save this notification with your User Documentation or post this notification by each VITROS 250/350/5,1 FS/4600/5600/XT 3400/XT 7600 System in your laboratory until the issue has been resolved. • Please forward this notification if the affected product was distributed outside of your facility. • If your laboratory has experienced the issue with this product and you have not already done so, please report the occurrence to your local Global Services Organization. | 2024-12-09 08:36
Biomerieux | TEMPO® AC | - | - | A manufacturing issue has led to a certain number of empty vials of the specific lot of x D TEMPO AC ref We apologize for the inconvenience that this matter can cause at your site x D For your information all vials containing the powder media have the correct volume and there is no x D risk of underfilled vials Either the vial is filled with the good amount of media or is totally empty x D The investigation and corrective actions aimed at preventing this type of defect are currently x D underway | We request you to take the following actions: - Check the presence of powder inside the vials of the lot: 1010748510 of TEMPO® AC. If the vial is empty, please discard it. - If empty vials are found, contact your Local Customer Service to ask for a replacement. - Distribute this information to all appropriate personnel in your laboratory, retain a copy in your files, and forward this information to all parties that may use this product, including others to whom you may have transferred the product. | 2024-12-09 14:02
ABBOTT S.r.l. | AlinIQ AMS | 2 | - | Please find attached letter | Please find attached letter | 2024-12-11 10:26
Getinge | Vaporizer Sevoflurane, QUIK FIL,Vaporizer Isoflur, Maquet filling | 51 | - | The Sevoflurane vaporizers are used with the following brands AbbVie Sevorane Ultane Maruishi Sevofrane Piramal Sevoflurane and Sevoflurane Baxter x D Item No Vaporizer Sevoflurane Maquet filling is used with AbbVie Sevorane Ultane Maruishi Sevofrane Piramal Sevoflurane and Sevoflurane Baxter x D Item No Vaporizer Sevoflurane SAFE T SEAL is used with Sevoflurane Baxter x D Item No Vaporizer Sevoflurane Quik Fil is used with Abbvie Sevorane x D Maquet Critical Care AB has received complaints where a discoloration corrosion inside the vaporizer have been observed under normal operating conditions Analysis has revealed chemical degradation of sevoflurane resulting in formation of hydrogen fluoride within the vaporizer x D The vaporizer is a component in the Flow family anesthesia systems used to contain and vaporize the anesthetic agent Sevoflurane is a sweet smelling nonflammable highly fluorinated methyl isopropyl ether used as an inhalational anesthetic for induction and maintenance of general anesthesia x D There is an inherent risk for chemical degradation of sevoflurane to hydrogen fluoride acid This may result in inhalation and or skin exposure of hydrogen fluoride when the vaporizer is used with low water content sevoflurane x D This issue has only been observed when using low water content Sevoflurane Piramal Sevoflurane and Sevoflurane Baxter No complaints have been reported with the use of Abbvie Sevorane Ultane or Maruishi Sevofrane x D To date no patient or operator injuries have been reported x D The item number REF and serial number SN are found on the label located underneath each vaporizer see image below | Please refer to the FSN | 2024-12-15 08:57
GeneOhm Sciences Canada, Inc. | BD MAX™ PCR Cartridges | 4416 | - | BD has identified through customer complaint investigations that signal drift has been observed in specific lots of BD PCR Cartridges BD has identified the following impacts x D x D For the BD MAX RVP assay product codes signal drift could cause false positive results for influenza A x D x D For specific BD MAX molecular assays including Enteric Bacterial Panel EBP product code Enteric Parasite Panel EPP product code StaphSR product code MDR TB product code CPO product code signal drift could potentially cause an increase in unresolved results UNR x D x D For third party assays used with the open system reagent kits BD MAX ExK DNA DNA DNA DNA MMK DNA MMK SPC TNA MMK TNA MMK SPC TNA and TNA signal drift could potentially cause an increase in unresolved results UNR or unexpected results x D | Clinical User Actions For the BD MAX™ RVP assay (product codes 445215 & 445373): 1. Do not use the affected BD PCR Cartridge lots listed in Appendix 1 with the RVP Assay. 2. Users can continue running the RVP Assay (product codes 445215 & 445373) with unaffected cartridge lots. 3. Customers may consider confirmatory testing for when a positive influenza A result has been obtained using the affected PCR cartridges with the RVP assay (product codes 445215 & 445373) but the patient has not shown clinical improvement. All other BD molecular assays: 1. Users could possibly observe an increase in UNR results for the other assays referenced above when used with the affected lot numbers of the BD PCR Cartridges. 2. Users should adhere to the guidance contained in product instructions for use for managing UNR results. 3. There are no recommendations for other results review or retesting, but customers are not precluded from doing so if they deem beneficial. BD Actions: BD has thoroughly investigated this issue and has identified the root cause. Therefore BD is taking actions both internally and with our supplier to prevent recurrence of this product issue in the future. Customer Actions: • Review the information in Appendix 1 to determine if the BD PCR Cartridges in your possession are impacted. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 21st January 2025. • Circulate this notice to all those who need to be aware within your organisation or to any organisation where the potentially affected product has been transferred. • If you experience any issues please report as a complaint as per your normal process. Distributor Actions: • Review the information in Appendix 1 and determine if the BD PCR Cartridges in your possession are impacted. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 21st January 2025. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues please report as a complaint as per your normal process. | 2024-12-15 14:39
Getinge | Vaporizer Sevoflurane, QUIK FIL,Vaporizer Isoflur, Maquet filling | 150 | - | Sevoflurane is susceptible to degradation which can occur through contamination with strong Lewis acids The aluminum container in the vaporizer is designed to ensure product stability during expected clinical practice but is not intended for long term storage of sevoflurane Furthermore product stability is only verified with the specific brand of anesthetic agent it is intended to be used with | we want to inform all users of Getinge/Maquet sevoflurane vaporizers, both existing and new, to pay attention to the following information. 1. Do not leave sevoflurane in the vaporizer for a longer period than 30 days or during external transportation. The vaporizers must be emptied if the sevoflurane is not planned to be used within the next 30 days or before external transportation. Empty and dry-run the vaporizer according to instructions in the User´s Manual and Appendix 1. 2. Prior to using a vaporizer containing sevoflurane, ensure that the vaporizer has been in use within the most recent 30 days. If the sevoflurane has not been used during the most recent 30 days, it must be discarded. Empty and dry-run the vaporizer according to instructions in the User´s Manual and Appendix 1. 3. As already stated in the User’s Manual, Sevoflurane vaporizers are only tested for use according to the table | 2024-12-16 14:42
Cardiac Pacemakers Inc. a subsidiary of BSC Corporation | ACCOLADE™ Family of Pacemakers and CRT-Ps | 14 | - | A subset of approximately of devices from the ACCOLADE family built before Sep are included in the advisory population and have an increased potential to initiate Safety Mode during telemetry or other normal higher power operations due to latent high battery impedance | Boston initialed field action noticed with recommendations: Recommendations: 1- Individual patient evaluation. Promptly identify patients within the advisory population who are at risk of harm due the non-programmable parameters in Safety Mode. 2- Replacement. Safety Mode. If a device enters Safety Mode, perform emergent replacement for patients who are at risk of harm. For other patients, non-emergent replacement is recommended. When choosing a replacement interval, do not rely on previously reported battery time remaining estimates which do not account for Safety Mode’s increased outputs nor the battery’s high impedance state. 3- Follow-up interval. Perform system follow-up in accordance with the instructions for use: • Perform a system follow-up via remote or in-office interrogation at least every 12 months; and • When remaining longevity reaches One-Year-Remaining, follow-up every three (3) months thereafter until replacement is indicated. 4- Medical records. For each patient with an affected device, append/update the patient’s medical record with this letter to maintain awareness to all follow-up physicians of this topic for the remaining service life of the device. | 2024-12-16 14:47
GE Hualun Medical Systems Co. Ltd | Brivo OEC 715 | 7 | - | GE HealthCare has become aware that certain Infinia VariCam VG VH Brivo NM Discovery NM Discovery NM CT and Optima NM CT Nuclear Medicine systems see Affected Products Details and Appendices could have been transported or relocated without proper detector support which can result in excessive load being applied to the detector mounting mechanisms If this occurs it can potentially result in a detector fall and lead to life threatening bodily injury x D x D There have been no reports of detector falls or injuries as a result of this issue x D | Inspection tool, ‘immediate repair’ screws kit, inspection & correction instructions, physical warning label referencing up-to-date de-install & install manual, FRU packaging available for order | 2024-12-17 12:39
GE Hualun Medical Systems Co. Ltd | Discovery NM 630 | 7 | - | GE HealthCare has become aware that certain Infinia VariCam VG VH Brivo NM Discovery NM Discovery NM CT and Optima NM CT Nuclear Medicine systems see Affected Products Details and Appendices could have been transported or relocated without proper detector support which can result in excessive load being applied to the detector mounting mechanisms If this occurs it can potentially result in a detector fall and lead to life threatening bodily injury x D x D There have been no reports of detector falls or injuries as a result of this issue x D | Inspection tool, ‘immediate repair’ screws kit, inspection & correction instructions, physical warning label referencing up-to-date de-install & install manual, FRU packaging available for order | 2024-12-17 12:39
GE Hualun Medical Systems Co. Ltd | Discovery NM/CT 670 | 7 | - | GE HealthCare has become aware that certain Infinia VariCam VG VH Brivo NM Discovery NM Discovery NM CT and Optima NM CT Nuclear Medicine systems see Affected Products Details and Appendices could have been transported or relocated without proper detector support which can result in excessive load being applied to the detector mounting mechanisms If this occurs it can potentially result in a detector fall and lead to life threatening bodily injury x D x D There have been no reports of detector falls or injuries as a result of this issue x D | Inspection tool, ‘immediate repair’ screws kit, inspection & correction instructions, physical warning label referencing up-to-date de-install & install manual, FRU packaging available for order | 2024-12-17 12:39
RDG | IRON2 | 6841 | - | Internal investigations confirmed the issue and revealed a systematic sample drift up to mol L absolute for x D IRON over the entire measuring range The bias increases with the number of tests performed from one cobas c x D pack without further calibration The first measurements are not affected while the last sample can exhibit the maximal x D observed bias x D The magnitude of the effect depends on multiple factors of the laboratory s routine time analyzer throughput IRON x D throughput calibration intervals The effect is not linked to the on board time x D Optimal hardware and maintenance status of the module can reduce the risk of the occurrence of the issue Optimizing x D piercer reagent probe reagent rotor adjustment as well as outside wash adjustment and gear pump pressure x D adjustment also mitigate the issue Iron abraded from the reagent probes caused by the screw caps of other cobas c x D packs used in parallel to IRON leads to iron contamination of the IRON reagents resulting in a positive bias x D Only IRON in the cobas c pack is affected | The availability of the new cassette IRON Gen.2 (100 tests) Mat. No. 10059605190 depends on local registration timelines. The mitigation is achieved by lowering the total throughput from a cassette while adjusting the filling volume of the assay's reagents. The workaround described below is no longer necessary after switching to the new cassette. Updated settings for cobas c 311, cobas c 501/502 and COBAS INTEGRA 400 plus and a new method sheet for cobas c and INTEGRA 400 plus are released. For the US two new method sheets for cobas c and COBAS INTEGRA 400 plus are released. An updated cassette definition for IRON2 for COBAS INTEGRA 400 plus is provided via TAS 36.10 by your local affiliate. Using the new material number no longer require the workaround of manual calibrations. Using the workaround in conjunction with IRON Gen.2, 200Tests remains safe and effective. The updated setting allows for the use of both cassettes. Required user actions for various instrument platforms ● On cobas c 311, cobas c 501, these settings cannot be overwritten by an updated e-barcode, but a complete deletion and re-installation of the application will update these user-editable settings. ● On cobas c 311, cobas c 501 and cobas c 502, c packs in use (including spare cobas c packs for the same parameter) must be removed before the updated e-barcode version can be downloaded, and cannot be registered again afterwards. ● On cobas c 502, the reagent c pack in use (inclusive stand by c packs for the same parameter) must first be unloaded before the updated e-barcode version can be downloaded. User actions on cobas c 501/311 systems: On cobas c 501 and c 311, the reagent c pack in use (inclusive stand by c packs for the same parameter) must first be unloaded and discarded before the increased e-barcode version can be downloaded. Delete the IRON2 ACN 661 application and install the updated application by choosing the Download button. User actions on cobas c 502 systems: Completely overwrite the IRON2 ACN 661 application with the updated e-barcode version by choosing the "Overwrite" button of the Confirmation window. In version 1 and 2 the customers were advised to implement the following workarounds depending on their throughput on the respective analyzer: ● Run batch measurements for IRON2 (this workaround is applicable regardless of number on the test determinations per day) or ● It is recommended to run a blank calibration with the zero standard using deionized water on the cobas c 311/501/502 analyzers or perform a full calibration on COBAS INTEGRA® 400 plus after at least every 50 IRON2 determinations out of one cobas c pack. Several workaround possibilities are described below separated by ● Customers performing < 50 IRON2 determinations per day out of one cobas c pack ● Customers performing ≥ 50 IRON2 determinations per day out of one cobas c pack For technical details with respect to different analyzers, please refer to the instructions attached to the FSN-RDS- CoreLab-2021-003 version 2. | 2024-12-19 10:18
Olympus | Single Use Mechanical Lithotriptor V | 6 | - | Olympus has identified an increase in complaints for BML V QR The complaint data analysis found that distal tip tearing see Figure of the Mechanical Lithotriptor V had increased beginning with the production of lot K Olympus has identified complaints for the BML V QR globally between June through July There were reportable malfunctions and there was one report of a serious injury in relation to this issue Olympus investigation confirmed that the issue is limited to the lots included in this letter and there is an ongoing investigation of this issue to prevent further occurrence x D | 1. Examine your inventory for the impacted Single Use Mechanical Lithotriptor V lot numbers (Table 1) and quarantine any affected devices. The lot number can be located on the package as follows: 2. Cease usage of the impacted lot numbers with immediate effect. 3. If you have affected products in your inventory, please contact Olympus with regard to return of affected products. Olympus will issue a credit to your facility upon return of your affected product. 4. Olympus requests that you acknowledge receipt of this letter by completing and returning the enclosed Reply Form to your local Olympus representative ra@olympus-mea.com 5. Please forward this notice to other users who may have the affected products if you have further distributed it. | 2024-12-19 16:32
Draeger Medical Systems Inc | Infinity CentralStation | 2 | - | While the new Edifier speaker configured with the ICS in analog mode the new Edifier speaker may not work as expected after a power loss or power USB cable being disconnected The new Edifier speaker configured with ICS in analog mode default back to USB mode after power being restored or power cable USB cable reconnected | Solution is through instructions provided in the ‘Important Safety Notice’. If the devices are affected, VG 04 SW will install | 2024-12-22 09:39
Olympus | Thunderbeat | 36 | - | Olympus has been made aware via customer feedback of reports where probe tips of the THUNDERBEAT Type S Hand instrument are becoming damaged or are breaking Reference example in Figure as well as instances of pad damage or detachment Reference example in Figure These issues can occur when the instructions and warnings in the IFU are not followed specifically taking very large bites of tissue contact with metal while activating and activation without tissue between the jaws x D During use if the device probe breaks or there is significant tissue pad damage an audible alarm tone will be generated and an error will be visibly displayed on the generator Follow the steps on the screen and in the IFU associated with the error message For probe damage and ultrasonic instrument damage related errors per the IFU immediately interrupt the procedure and replace the instrument Even if the damage cannot be confirmed visually do not attempt to reuse the instrument A very fine invisible crack may develop into a detached instrument tip Not replacing the instrument per the IFU can cause probe damage breakage and pad damage detachment which can result in patient harm | 1.Ensure all personnel are completely knowledgeable and thoroughly trained on the content of this FSN and the relevant IFU. 2. If you have further distributed this product, identify your customers, and forward this FSN to them. 3. Olympus requests that customers acknowledge receipt of this FSN. Olympus requests that customers report any complaints, ra@olympus-mea.com | 2024-12-22 14:20
Agilent Technologies, Inc. | PD-L1 IHC 22C3 pharmDx | 7 | - | The use of controls as directed in the instructions for use IFU monitors assay performance to help x D confirm the assay is working as intended the negative controls are used to identify nonspecific staining or x D background noise such that only PD L expression is detected The risk of harm is deemed improbable x D and there is no change to the validity of past test results | . If increased nonspecific background staining is observed, please contact Agilent technical support, your local sales representative, or Agilent Field Service Support. Please confirm that you have received this information by signing and returning the enclosed Acknowledgement Form to fieldactions@agilent.com | 2024-12-22 15:38
Intuitive Surgical | Reusable Instruments with Jaws | 1820 | - | An increase in complaints regarding frayed or broken cables on some da Vinci X and Xi reusable instruments These instruments can be used with da Vinci X da Vinci Xi and da Vinci systems | Please take the following Actions: 1. Customers can continue using the products in accordance with the user manual. 2. As a reminder, when using da Vinci X and Xi reusable instruments, follow the Inspection Before Use and Warnings listed in the manual provided with your system to inspect for any broken cables. Refer to Figures C and D for examples. Inaddition, please refer to Appendix B for additional images for detection of grip cable failures. 3. If you observe any failed (frayed or broken) grip cables prior to use, during procedure, or during reprocessing, please stop use of instrument, remove from use and inform Intuitive via the standard complaint process. 4. Inform Intuitive of any Adverse Events*/Serious Incidents** or quality problems concerning the use of the subject device via the standard complaint process. | 2024-12-23 00:16
Draeger Medical, Inc. | Oxylog 3000 plus | 483 | - | In the course of our worldwide market monitoring cases have become known to us where the Oxylog plus emergency and transport ventilators have stopped ventilation due to empty batteries despite being reconnected to the mains power supply after previous battery operation x D Due to a timing problem in the firmware of the charging circuit PCB the charging process may be interrupted making the switch to mains operation impossible x D If the Int Battery Not Charging alarm message is displayed and the charge status LED is red the battery must be removed and reinserted or the internal battery must be replaced Do not disconnect Oxylog plus from the external power supply | Customer notification on the supplement to the instructions for use via the updated safety information by 2025-02-28. Firmware update for the charging circuit PCB during next maintenance, at the latest by 2025-05-31. | 2024-12-23 09:25
Fresenius Medical Care AG | 4008B, 4008S (excluding S V10!) | 740 | - | Fresenius Medical Care received post market suveillance reports that the new BD Plastipak ml LL syringe article code F is not locking properly into the heparin syringe tray of the following no longer sold devices H B S excluding S V Art Standard Universal and Prometheus x D Currently the BD Plastipak ml LL syringe is the only approved ml heparin syringe for use with the affected devices following the phase out of its predecessor x D Due to a smaller outer diameter the syringe may fall out of the syringe pump tray during treatment by applying a small amount of force e g when adjusting the disposable consumables | The heparin line is placed after the blood pump on the arterial blood line. The positive pressure at this point may cause blood in the heparin line to rise and push the syringe plunger out of the syringe. This may result in an uncontrolled loss of blood and can lead to serious patient harm. Fresenius Medical Care recommends that healthcare providers do the following: Stop using the BD Plastipak 30ml LL syringe with the above-mentioned Fresenius Medical Care devices. Please apply adequate anticoagulation according to prescription by other means. Consult the responsible physician. Stop using the BD Plastipak 30ml LL syringe with the above-mentioned Fresenius Medical Care devices. Please apply adequate anticoagulation according to prescription by other means. Consult the responsible physician. | 2024-12-23 10:31
Intuitive Surgical | Tenaculum Forceps - Small Graptor | 55 | - | This Medical Device Notification is to inform customers that Intuitive has observed an increase in complaints regarding pitch cable failure on the Tenaculum Forceps PN and Small Graptor PN Pitch cable failure may also result in a fragment of the pitch cable and its end crimp becoming dislodged from the instrument The purpose of the communication is to inform users inspect for any fragments if a pitch cable failure occurs x D | 1. As a reminder, when using the Tenaculum Forceps and the Small Graptor instruments, please refer to and follow the instructions, warnings and cautions provided in the General Overview and EndoWrist Instrument chapters of the da Vinci X/Xi Instruments and Accessories User Manual and Reprocessing Instructions User Manual. a. In addition, please reference section titled “General Precautions for Intraoperative Use of Instruments” in the da Vinci X/Xi Instruments & Accessories User Manual and the section titled “General Cautions and Warnings” in the da Vinci X/Xi Reprocessing Instructions User Manual. b. Please refer to Appendix B for additional images for detection of pitch cable failures. 2. If you observe any failed (frayed or broken) pitch cables prior to use, during procedure, or during reprocessing, please stop use of instrument, remove from use and inform Intuitive via the standard complaint process. 3.Please display this communication with your systems; ensure it is in a place likely to be seen/viewed by operators. 4. Complete the attached Acknowledgement Form immediately and return it via fax or email to Intuitive as instructed on the form. 5. Please retain a copy of this letter and the acknowledgement form for your files. 6. Inform Intuitive of any Adverse Events*/Serious Incidents** or quality problems concerning the use of the subject devices via the standard complaint process | 2024-12-23 12:23
Werfen | HemosIL Heparin Calibrators | 49 | - | We have received customer complaints of heparin controls reporting low out of their assigned acceptance range with x D HemosIL Liquid Anti Xa leading to delays in patient testing Our internal investigation has identified that the value x D assignment of the HemosIL Heparin Calibrators has shown a bias against the WHO International Standards for UFH and x D LMWH potentially causing reduced recovery when measuring heparin on both controls and samples x D Note Apixaban and rivaroxaban measurements with HemosIL Liquid Anti Xa are unaffected as these tests use x D different calibrators x D To ensure accurate heparin reporting customers must follow the instructions below to update the calibrator values for x D Heparin Calibrator HEP CAL and Heparin Calibrator HEP CAL in the ACL TOP test parameters all instrument x D models ACL TOP Family ACL TOP Family Series and ACL TOP Family Series using the corrected values provided x D in Appendix A The Heparin Calibrator HEP CAL level remains unchanged as IU mL x D We will also verify the value assignments for on market lots of HemosIL UF Heparin Controls Part No and x D HemosIL Low Molecular Weight Heparin Controls Part No using a representative reagent lot recalibrated x D based on the corrected HEP CAL and HEP CAL values If needed we will provide new acceptance ranges for these x D controls to you Until then perform quality control testing per labeled instructions for use If the HemosIL UF Heparin x D Controls or HemosIL Low Molecular Weight Heparin Controls recover out of the lot specific acceptance ranges x D determine your own control ranges per your internal protocol | Based on the above, please take the following immediate actions:  Check inventory for lots of HemosIL Heparin Calibrators (Part No. 0020300600). See Appendix A for a list of in-date lots with corrected value assignments for the HEP CAL 2 and HEP CAL 3 levels.  Follow the instructions below to update the ACL TOP test parameters (same instructions for all models):  Select “Setup” from the menu bar and then select “Test List” from the dropdown menu.  From “Test List”, double click on the Test Code “Anti-Xa” to open “Test Definition – Anti-Xa”.  From “Test Definition – Anti-Xa, click on “Pre-diluted calibrators” under the header “Calibration”. An editable table titled “Calibration points” will appear with HEP CAL 1, HEP CAL 2, and HEP CAL 3 and their predefined “Target Values IU/mL” of 0.00, 0.80 and 2.00, respectively.  In the “Calibration points” table, change the target values for HEP CAL 2 and HEP CAL 3 to the lot-specific calibrator values provided in Appendix A.  After entering the corrected HEP CAL 2 and HEP CAL 3 values, click on the save icon (picture of floppy disk) to save the changes.  Recalibrate using the HemosIL Liquid Anti-Xa assay lot in use after making the test parameter update.  Run quality controls per labeled instructions for use. If needed, determine control ranges for HemosIL UF Heparin Controls (Part No. 0020300300) and HemosIL Low Molecular Weight Heparin Controls (Part No. 0020300200) in accordance with your internal protocol until the new acceptance ranges are available.  Repeat the above test parameter update, recalibration, and quality control for all ACL TOP instruments.  Share this information with your laboratory staff.  Forward this notification to all affected locations within your facility.  Retain a copy of this notification for your records. | 2024-12-23 13:30
Imaxeon Pty Ltd | MEDRAD CENTARGO CT INJECTION SYSTEM | -1 | - | In August Microsoft disclosed a cybersecurity vulnerability impacting systems with x D IPv enabled This vulnerability would allow an unauthenticated attacker in a worst case x D scenario exploit to obtain full system access Only the Centargo Injector System is x D vulnerable x D This recall action was not notified to the TGA prior to being undertaken x D Recall Actionix x D Product Defect Correctio | informed manufacturer to send software once received will go to site to perform correction | 2024-12-23 14:55
Baxter Healthcare SA | PrisMax | 34 | - | The values sent to the Electronic Medical Records EMR Patient Monitoring Systems history screen and the operation screens are properly displayed until therapy reaches hours At the hour mark the displayed Effluent Delivered values appear as half of what was actually delivered to the patient Although the displayed Effluent Delivered dose is incorrect the correct amount of effluent gets removed from patients by the device and therapy is performed correctly by the PrisMax system | Baxter Healthcare Corporation (Baxter) is issuing Correction for the PrisMax Systems that have been upgraded to software version 3.4. Baxter has identified an issue with this software version that causes some displayed Effluent dose values to be different than what was actually delivered by the PrisMax system. | 2024-12-23 15:10
KLS Martin SE & Co. KG | Intermediate ceiling constr. | 8 | - | Incorrect tightening torques are specified in the installation instructions in relation to the suspended ceiling construction In the tightening torques have been increased by mistake The correct torque is Nm for M bolts and Nm for M bolts | Identify Device Take note of amendment of Assembly Instructions and Service Manual/ Forward this FSN to all customers who have been supplied with the affected products. The customer must check his system on site in accordance with the specifications in the service manual (Chapter 13.3.1 Inspection instructions). If damage is detected, the lighting system must be blocked or dismantled and a KLS Martin technician must be informed immediately. | 2024-12-23 17:34
Spacelabs Healthcare | Eclipse Pro , Eclipse Mini | 4 | - | x D Some Holter devices Eclipse Pro and Eclipse Mini referred to as devices in the rest of this document may have mismatched serial numbers between the label and what is programmed in the devices This is due to a process error on the manufacturing line which has resulted in an incorrect serial number being set in a small number of the devices We are notifying all customers using the devices and providing simple guidelines to ensure this mismatch does not result in any unreasonable risk x D | ANY serial numbers do NOT match Action required: please contact your service representative and they will update the serial number loaded in the device remotely for you. If ALL device serial numbers match No further action is required. Spacelabs has automatically been notified of the result of all the checks carried out. | 2024-12-24 10:05
Medtronic, Inc. | EOPA™ Elongated One-Piece Arterial Cannula | 20 | - | During a review of complaints for devices manufactured at the Medtronic Grand Rapids formerly Viant facility five additional complaint events were identified as mislabeled and warranted assessment via CAPA Products for the impacted lot numbers were incorrectly labeled with an incorrect size Based on complaints received Medtronic cannot determine the exact number of mislabeled units but it is confirmed that at least one cannula from each of the affected lot numbers has been mislabeled | • Review your inventory for listed product. • Immediately identify and quarantine all unused, listed product in your inventory. • Return unused, listed product in your inventory to Medtronic. Your Medtronic representative can assist you in the return of affected product as necessary. • Complete the enclosed Customer Acknowledgement Form and email to Medtronic This form must be returned even if you do not have any affected product in your possession. • This notice needs to be passed on to all those who need to be aware within your organization or to any organization where the potentially affected devices have been transferred. • Please maintain a copy of this notice in your records. | 2024-12-25 09:43
Boston Scientific Corporation | AXIOS™ Stent and Electrocautery Enhanced Delivery System | 18 | - | Boston Scientific is initiating a removal of specific lots of the AXIOS Stent and Electrocautery Enhanced Delivery System due to reports of the outer sheath distal black tip detaching from the device x D | Removal of specific lots of the AXIOS Stent | 2024-12-25 16:17
GE OEC Medical Systems, Inc. | OEC Elite | 2 | - | GE HealthCare has become aware that recently manufactured OEC Elite and OEC D mobile C arms as well as recently replaced X ray tubes on certain OEC Elite OEC D and OEC OEC mobile C arms potentially have insufficient sealing that can result in an oil leak This issue could result in loss of fluoroscopic imaging capability x D x D There have been no injuries reported as a result of this issue x D x D x D | Phase 1 – Inspection Only; Phase 2 – If defect is applicable, replacement of x-ray tube is required. | 2024-12-26 16:14
GE OEC Medical Systems, Inc. | OEC 9900 Elite | 2 | - | GE HealthCare has become aware that recently manufactured OEC Elite and OEC D mobile C arms as well as recently replaced X ray tubes on certain OEC Elite OEC D and OEC OEC mobile C arms potentially have insufficient sealing that can result in an oil leak This issue could result in loss of fluoroscopic imaging capability x D x D There have been no injuries reported as a result of this issue x D x D x D | Phase 1 – Inspection Only; Phase 2 – If defect is applicable, replacement of x-ray tube is required. | 2024-12-26 16:14
GE OEC Medical Systems, Inc. | OEC 3D | 2 | - | GE HealthCare has become aware that recently manufactured OEC Elite and OEC D mobile C arms as well as recently replaced X ray tubes on certain OEC Elite OEC D and OEC OEC mobile C arms potentially have insufficient sealing that can result in an oil leak This issue could result in loss of fluoroscopic imaging capability x D x D There have been no injuries reported as a result of this issue x D x D x D | Phase 1 – Inspection Only; Phase 2 – If defect is applicable, replacement of x-ray tube is required. | 2024-12-26 16:14
Imaxeon Pty Ltd | MEDRAD CENTARGO CT INJECTION SYSTEM | 1 | - | x D installed without one being present This situation would then trigger the injection system to x D auto prime saline which in this case would be dispensed onto the floor and it would also x D change the light display to Fluid color Blue Subsequent Patient Line installation would not x D be detected by the system and system auto prime would not occur creating the potential for x D air injection in extremely rare circumstances | checked the unit and verified the problem performance the corrective action as per the manufacturer procedure instruct also the user regarding this procedure unit tested and working properly | 2025-01-01 16:49
CooperSurgical | Medical Device IdentifierK‐Systems™ T47 470‐ x 400‐mm Warming Plate, K‐Systems™ T45 356‐ x 220‐mm Warming Plate, K‐Systems™ G73 Block Warmer, K‐Systems™ G85 Mini Incubator, K‐Systems™ R65 Trolley | 33 | - | Excessive temperatures could result in a x D potential loss of gametes and embryos or burns to the user if skin were to come in contact with the warming x D surfaces or incubator wells Itisimportantto note that the visual and audible alarmsfunction properly when there x D is a variance of degree from the set temperature notifying the user of the variance | G85 Mini Incubators with serial numbers listed in the Appendix are impacted and require repair. 1. Inspect your inventory, identify, and discontinue use. 2. Complete the attached Customer Information Letter Response Form and return the form to CooperSurgical asindicated atthe top of the form. Note: Please complete and return the formeven if you do not have any affected product in your inventory so that we may document receipt of this Customer Information Letter. 3. Contact yourlocal Service Representative to arrange a repair at no costto you. T45, T47, R65, and G73 with serial numbers listed in the Appendix may be impacted and require inspection to determine the need for repair. 1. Inspect yourinventory, identify, and discontinue use. 2. Complete the attached Customer Information Letter Response Form and return the form to CooperSurgical asindicated atthe top of the form. Note: Please complete and return the formeven if you do not have any affected product in your inventory so that we may document receipt of this Customer Information Letter. 3. Contact your local Service Representative to arrange an inspection of the affected product. If your product is affected, the product will be repaired at no cost to you. | 2025-01-05 11:41
Ortho Clinical Diagnostics Inc. | VITROS 5600 Integrated System | 119 | - | If an expired pack is used on the system given that shelf life claims are established by generating data beyond the labelled claim the gradual signal loss overtime and typical customer usage patterns we do not expect any significant biased patient results or quality control result failures In the worst case and unlikely scenario that an expired SR pack results in calibration failures there is the potential for a delayed patient result depending on how long it takes to investigate and troubleshoot the failure Delayed results could result in suboptimal management of critically ill patients x D x D Condition code PVG correctly appears when Signal Reagent expires either on board or shelf expiration x D x D QuidelOrtho does not recommend a review of previous results in identifying patient results which may have been affected by this issue given the gradual signal decline and low likelihood of any significant bias Discuss any concerns regarding previously reported results with your Laboratory Medical Director to determine the appropriate course of action The results from any diagnostic test should be evaluated in conjunction with a patient s history risk factors clinical presentations signs and symptoms as well as the results of other tests x D x D Root Cause x D QuidelOrtho has identified the cause as a software anomaly This issue will be corrected with the next software release estimated for Q of | • Do not use expired Signal Reagent packs. See the Questions & Answers section for more details. • Complete the enclosed Confirmation of Receipt form no later than Month DD, 2024. • Save this notification with your User Documentation or post this notification by each VITROS 3600/5600/XT 7600 System in your laboratory until the issue has been resolved. | 2025-01-06 16:34
Ortho Clinical Diagnostics Inc. | VITROS 3600 Immunodiagnostic System | 119 | - | If an expired pack is used on the system given that shelf life claims are established by generating data beyond the labelled claim the gradual signal loss overtime and typical customer usage patterns we do not expect any significant biased patient results or quality control result failures In the worst case and unlikely scenario that an expired SR pack results in calibration failures there is the potential for a delayed patient result depending on how long it takes to investigate and troubleshoot the failure Delayed results could result in suboptimal management of critically ill patients x D x D Condition code PVG correctly appears when Signal Reagent expires either on board or shelf expiration x D x D QuidelOrtho does not recommend a review of previous results in identifying patient results which may have been affected by this issue given the gradual signal decline and low likelihood of any significant bias Discuss any concerns regarding previously reported results with your Laboratory Medical Director to determine the appropriate course of action The results from any diagnostic test should be evaluated in conjunction with a patient s history risk factors clinical presentations signs and symptoms as well as the results of other tests x D x D Root Cause x D QuidelOrtho has identified the cause as a software anomaly This issue will be corrected with the next software release estimated for Q of | • Do not use expired Signal Reagent packs. See the Questions & Answers section for more details. • Complete the enclosed Confirmation of Receipt form no later than Month DD, 2024. • Save this notification with your User Documentation or post this notification by each VITROS 3600/5600/XT 7600 System in your laboratory until the issue has been resolved. | 2025-01-06 16:34
Ortho Clinical Diagnostics Inc. | VITROS XT 7600 Integrated System | 119 | - | If an expired pack is used on the system given that shelf life claims are established by generating data beyond the labelled claim the gradual signal loss overtime and typical customer usage patterns we do not expect any significant biased patient results or quality control result failures In the worst case and unlikely scenario that an expired SR pack results in calibration failures there is the potential for a delayed patient result depending on how long it takes to investigate and troubleshoot the failure Delayed results could result in suboptimal management of critically ill patients x D x D Condition code PVG correctly appears when Signal Reagent expires either on board or shelf expiration x D x D QuidelOrtho does not recommend a review of previous results in identifying patient results which may have been affected by this issue given the gradual signal decline and low likelihood of any significant bias Discuss any concerns regarding previously reported results with your Laboratory Medical Director to determine the appropriate course of action The results from any diagnostic test should be evaluated in conjunction with a patient s history risk factors clinical presentations signs and symptoms as well as the results of other tests x D x D Root Cause x D QuidelOrtho has identified the cause as a software anomaly This issue will be corrected with the next software release estimated for Q of | • Do not use expired Signal Reagent packs. See the Questions & Answers section for more details. • Complete the enclosed Confirmation of Receipt form no later than Month DD, 2024. • Save this notification with your User Documentation or post this notification by each VITROS 3600/5600/XT 7600 System in your laboratory until the issue has been resolved. | 2025-01-06 16:34
Imaxeon Pty Ltd | MEDRAD Intego 200 PET Infusion System | 3 | - | Impact is to Intego or Intego PET Infusion Systems with a POC or POC x D display These displays are only compatible with Windows only a POC display is x D compatible with Windows | Product Defect Correction | 2025-01-06 17:15
LANDANGER | INSUFFLATION NEEDLE | 1950 | - | The Landanger company has been informed of complications following a surgical procedure for which a single use instrument from the Landanger range was used This was an insufflation needle reference E lot expiry date The user has reported a perforation of an organ intestine during a laparoscopic procedure x D The user informs us that he did not feel the jump that must occur to cross the different planes and that he she pushed his needle in too far and it went through the intestine | 1- Identify the products in the batch concerned. 2- Quarantine the products in the batch concerned. 3- Send this letter containing the information to be added to the instruction leaflet to users. | 2025-01-07 10:07
Randox Laboratories | HbA1c II CAL | 4 | - | Randox Laboratories has identified that the HbAlc Il Calibrator Series HA is running with a negative bias on QC s and Patient samples We have realigned the target values in the above lot numbers to the NGSP standardisation and provided instrument specific targets for the RX Imola and RX Modena You may experience a positive shift in Quality Control and patient sample recovery by approximately | Action to be taken: Review results generated in line with the clinical profile of the patient. Discuss the contents of this notice with your Medical Director. Complete and return the response form 12187-QA to technical.services@randox.com within five working days. | 2025-01-09 11:11
Biocartis NV | Idylla™ Instrument | - | - | Biocar s iden fied that false posi ve MSI H test results could be or could have been generated by the Idylla MSI Test at your laboratory for the Idylla Instruments listed in Appendix and | 1. Local Regulatory Authori)es have been informed of this no)ce. 2. A service interven)on for the Instrument(s) listed in Appendix 1 will be performed. 3. Following an assessment of Idylla™ MSI Test run data for Appendix 2 Instrument(s), a service interven)on may be performed. 4. A Biocar)s representa)ve will contact you and provide details on when and how the (poten)al) service interven)on will be performed. | 2025-01-12 07:18
Coloplast | Foley catheters, Prostatic catheters, Urinary diversion devices, Neoplex® urethral catheters without balloon, Urethral bougies, Urodynamic catheters, Percutaneous nephrostomy catheters, Supra-pubic drainage set and Surgical drainage devices | 101217 | - | A sterility issue was detected | 1) Identify the device 2) Destory the device | 2025-01-13 07:42
Anhui Intco Medical Products Co., Ltd. | YASHFEEN VINYL EXAMINATION GLOVES | 11000 | عدم الالتزام بمتطلبات النقل والتخزين | - | All affected quantities have been withdrawn. | 2025-01-13 11:01
Tem Innovations GmbH (Werfen) | ROTEM sigma complete - ROTEM sigma complete + hep | 243 | - | It has come to our attention through a customer complaint that the FIBTEM C A results obtained with the x D current IVDR version exhibited a bias from measurements made with the previous IVDD version for a x D population within the Obstetric setting leading to an adjustment in a bleeding management algorithm | - Check if you have established a bleeding management algorithm for ROTEM sigma - If yes, verify the established bleeding management algorithm for a potential impact due to the bias between the IVDD and the IVDR cartridges and take appropriate validation actions for your algorithm following your internal procedures. - Document the acknowledgement on the Customer Reply Form and return the completed and signed form to the fax number or e-mail address listed on the next page. - Share this information with your laboratory staff and follow your internal procedures. - Forward this notification to all affected locations within your facility. - Retain a copy of this notification for your records | 2025-01-14 13:38
Tem Innovations GmbH (Werfen) | ROTEM sigma complete - ROTEM sigma complete + hep | 243 | - | It has come to our attention through a customer complaint that the FIBTEM C A results obtained with the x D current IVDR version exhibited a bias from measurements made with the previous IVDD version for a x D population within the Obstetric setting leading to an adjustment in a bleeding management algorithm x D | - Check if you have established a bleeding management algorithm for ROTEM sigma - If yes, verify the established bleeding management algorithm for a potential impact due to the bias between the IVDD and the IVDR cartridges and take appropriate validation actions for your algorithm following your internal procedures. - Document the acknowledgement on the Customer Reply Form and return the completed and signed form to the fax number or e-mail address listed on the next page. - Share this information with your laboratory staff and follow your internal procedures. - Forward this notification to all affected locations within your facility. - Retain a copy of this notification for your records. | 2025-01-14 14:26
(Solventum company (3M | 3M™ Prevena™ Plus 125 | 346 | - | This medical device recall is to inform customers who purchased the affected products of this issue to x D cease use and quarantine the impacted devices and to provide instructions on how to return devices for x D replacement x D Product complaints for this issue have been received but zero reports of injury have been reported x D There are no risks to patients The devices continue to provide negative pressure wound therapy but due x D to the lower leak alarm threshold nuisance alarms may be triggered and continue during therapy x D although there are no visible leaks x D Our records indicate that you have purchased one or more products affected by this voluntary field x D corrective action | Actions Required by Your Facility: 1. Screen your facility or other storage locations for the impacted products and lot numbers listed in Attachment 1. 2. Immediately cease use and quarantine any identified product. 3. Complete the Field Notice Receipt and Customer Response Form acknowledging you have received a copy of this notice and select one of two options: a. Acknowledge you do not have any product impacted by the field action. b. Acknowledge you have product impacted by the field action and complete the catalog number, lot numbers and quantities you have identified and quarantined. 4. Return the form via email to kci3mfieldactionresponse@solventum.com. 5. Call Solventum customer service at 1-800-275-4524 to help coordinate the return shipment and receive return labels. 6. Products returned to Solventum will be assessed, reprogrammed, and returned or replaced. a. Product skus and quantities shipped to Solventum will be returned or replaced with the same quantities. b. Returned or replaced product will have a visual identifier confirming the assessment and reprogramming. c. Returned or replaced product may have different lot numbers and expiry dates; Solventum plans to ship available product to help reduce wait times. | 2025-01-14 15:03
Randox laboratories | Liquid Assayed Specific Protein Controls | - | - | Randox have confirmed that the levels of Free Lambda Light Chains in our Liquid Assayed Specific Protein Controls will increase over the shelf life of the product While this analyte will remain in the Liquid Assayed Specific Protein Controls values and stability claims will no longer be quoted for this analyte for all lots The IFUs have been updated and are available on www randox com Please discard the old IFUs and download the updated IFUs | Action to be taken: Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA} to technical.services@randox.com within five working days. | 2025-01-14 15:56
Randox laboratories | Liquid Assayed Specific Protein Controls | - | - | Randox have confirmed that the levels of Free Lambda Light Chains in our Liquid Assayed Specific Protein Controls will increase over the shelf life of the product While this analyte will remain in the Liquid Assayed Specific Protein Controls values and stability claims will no longer be quoted for this analyte for all lots The IFUs have been updated and are available on www randox com Please discard the old IFUs and download the updated IFUs | Action to be taken: Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA} to technical.services@randox.com within five working days. | 2025-01-14 15:56
Randox laboratories | Liquid Assayed Specific Protein Controls | 2311 | - | The IFURandox have confirmed that the levels of Free Lambda Light Chains in our Liquid Assayed Specific Protein Controls will increase over the shelf life of the product While this analyte will remain in the Liquid Assayed Specific Protein Controls values and stability claims will no longer be quoted for this analyte for all lots s have been updated and are available on x D www randox com Please discard the old IFUs and download the updated IFUs | Action to be taken: Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA} to technical.services@randox.com within five working days. | 2025-01-14 16:01
Spacelabs Healthcare Inc. | Xhibit Telemetry Receiver | 1 | - | This recall is to inform users of an update to the XTRs user manual to include a routine restart of the device every months to maintain optimal performance and to reduce the risk of the device restating unexpectantly x D | "Xhibit Telemetry Receiver (96280) Service Manual, P/N 070-2409-07 has been updated by adding a section for “Required periodic system maintenance of Spacelabs Telemetry Farm”, pages 5-30 through 5-35.Download current versions of the Service Manual and PM Instructions via the Spacelabs user manual site at: https://manuals.spacelabshealthcare.com/ If you do not wish to create an account, you may obtain a copy of the latest manual by contacting Spacelabs Healthcare at the number listed below. .Discard previous copies of these documents .Complete and submit the customer acknowledgement form by mailing the included PDF copy of the response form." | 2025-01-14 17:27
BIOTRONIC AG | Pulsar-35 | 10 | - | It was determined that one lot of the Pulsar consisting of units was x D wrongly labeled The Pulsar LOT t was wrongly labelled x D as pulsar LOO | please note that your country "saudi Arabia" is NOT affected by this FSCA. This FSCA is only implemented in France. | 2025-01-15 07:22
Ortho Clinical Diagnostics | ORTHO VISION Analyzer | 129 | - | The purpose of this notification is to inform you that QuidelOrtho identified an issue on the ORTHO VISION and ORTHO VISION Max Analyzers that is configured to a Laboratory Information System LIS In certain cases an unexpected file directory has been configured for the LIS Images folder and after the installation of software version v the clean up feature removes file s older than days from the unexpected directory | • Prior to updating to Application Software v5.16 (MOD 90), users must verify a dedicated directory for LIS image files is configured. • Complete and return the enclosed Confirmation of Receipt form no later than Month DD, 2024. • Save this notification with your User Documentation or post this notification by each VISION or VISION Max in your laboratory until the issue has been resolved. • Please forward this notification if the affected product was distributed outside of your facility. | 2025-01-15 12:54
Ortho Clinical Diagnostics | ORTHO VISION Max Analyzer | 129 | - | The purpose of this notification is to inform you that QuidelOrtho identified an issue on the ORTHO VISION and ORTHO VISION Max Analyzers that is configured to a Laboratory Information System LIS In certain cases an unexpected file directory has been configured for the LIS Images folder and after the installation of software version v the clean up feature removes file s older than days from the unexpected directory | • Prior to updating to Application Software v5.16 (MOD 90), users must verify a dedicated directory for LIS image files is configured. • Complete and return the enclosed Confirmation of Receipt form no later than Month DD, 2024. • Save this notification with your User Documentation or post this notification by each VISION or VISION Max in your laboratory until the issue has been resolved. • Please forward this notification if the affected product was distributed outside of your facility. | 2025-01-15 12:54
Butterfly Network | Butterfly iQ+ Ultrasound System / Butterfly iQ App | 1 | - | Butterfly Network has identified a software bug in the Butterfly iQ application where the OB presets OB GYN and OB intermittently disappear on iOS mobile devices when a user goes offline As a result the user is temporarily unable to access the OB specific presets and applicable tools including OB Calculations and Fetal Heart Sounds When the user reconnects to the internet the presets and applicable tools become available All other presets imaging modes and features remain available This bug impacts Butterfly iQ app versions through when used with a Butterfly iQ probe | Healthcare professionals may use another preset, imaging modality, or other clinical assessment tools if an ultrasound exam with the Butterfly iQ+ device was not possible with the OB presets. A permanent fix for this software bug will be implemented in software version 2.38.1. To ensure resolution, please upgrade your Butterfly iQ app through the Apple App Store. | 2025-01-16 10:26
Siemens Healthcare Diagnostics Inc. | Atellica® CH RCRP | 84000 | - | Issue Description x D Siemens Healthineers has confirmed the occurrence of No Calculation flags when using the Atellica CH RCRP assay on the Atellica CH Analyzer This flag is only triggered for some samples when the CRP concentration is less than the low end of the measuring interval of mg dL mg L When this flag is triggered the concentration index is reported as Error Results within or above the measuring interval of mg dL mg L are not impacted For the RCRP assay no other conditions will trigger this flag x D This issue does not affect Atellica CH RCRP results on the Atellica CI Analyzer x D x D Kindly review the attached | Customer Actions • For the product listed in the table above, please perform the following steps: 1. For customers using ADM (Atellica Data Manager), contact your local Siemens support representative to create a rule for the RCRP assay to convert a “No Calculation” flag to a <0.05 mg/dL (<0.5 mg/L) result. 2. For customers not using ADM, RCRP test results that produce a “No Calculation” flag can be reported as <0.05 mg/dL (<0.5 mg/L). • Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached | 2025-01-16 14:57
HOYA Corporation | PENTAX Medical VIDEO PROCESSOR | 12 | - | Please find attached report | Please find attached report | 2025-01-19 12:21
HOYA Corporation | Pentax Video Colonscope | 12 | - | Please find attached report | Please find attached report | 2025-01-19 12:21
HOYA Corporation | PENTAX Medical Video Upper GI Scope | 12 | - | Please find attached report | Please find attached report | 2025-01-19 12:21
Sysmex Corporation | Sample Preparation System PS-10 | 1 | - | Sysmex was made aware of a report of false results because of carry over issue by omission of the probe wash step during antibody pipetting on sample preparation unit PS The lack of probe washing may cause carry over of small portion of antibody reagent to another reagent a primary sample or secondary daughter tubes x D This may cause incorrect sample preparation in the secondary daughter tubes and incorrect results when the sample is analyzed on a flow cytometer We would like to emphasize that there is only potential carry over of antibody reagent and no sample carry over x D The investigation reproduced the failure and confirmed the lack of necessary washing process during antibody pipetting This omission of the probe wash step could occur when the number of potential dispenses per antibody in a run exceeds two This software failure was introduced into PS software Ver and affects Ver and only Even if Ver or Ver is used the defect does not affect the cocktail creation | 1. Until a Sysmex representative visits you, please limit use only for cocktail creation or please do not use the instrument running PS-10 software Ver 1.5 or 1.6. 2. Review the content of this communication with your facility's physician and/or pathologist and retain this letter for any future reference. 3-Sysmex advises you to consult your facility's physician and/or pathologist to determine any implications (including retrospective review and/or re-testing) specific to your patients. | 2025-01-19 18:51
Biomet France | Optipac Refobacin Plus Bone Cement | 7 | - | the issue can be linked to an incomplete tightening of the mixing rod when it is assembled with the cylinder during use | Review this Field Safety Notice for awareness of the contents. This is a correction only, IFU will be updated, not a recall. | 2025-01-19 18:58
ICU medical | Medima P100, P200, and P300 Volumetric infusion Pumps | 25 | - | Our stepper motor supplier informed us of a potential issue due to an error in their manufacturing process with a specific lot of the stepper motors used to load and release the Free Flow Protection Clamp loading and releasing the pump set line This manufacturing issue can lead to the pump runway shown in Figure below not opening to allow the infusion set to be inserted into the pump | SW update | 2025-01-20 10:52
ICU medical | Medima Infusion Pump | 85 | - | A software issue was detected during internal testing The issue only impacts infusions that meet both conditions x D below x D The infusion is programmed using the Profile infusion mode with units of mL min or mL h x D The user titrates the dose rate after the infusion is started | SW Update | 2025-01-20 11:00
Simens | IMMULITE 2000 XPi Analyzer | 5 | - | When the IMMULITE EBV VCM IgM Instructions For Use are followed and the kitted Quality Controls are in range patient results are valid and acceptable for reporting x D If the kitted Quality Controls are out of the target value ranges specified in the QC Instructions For Use patient results cannot be reported potentially leading to an apparent delay in patient sample testing x D Assay shelf life performance will be restored with kit lots above with an expected release date at the end of January | If in-range Quality Control results cannot be obtained, you may request no-charge replacement product from your local Siemens Healthineers or distributor office. . • Complete and return the Product Replacement Form attached to this letter if replacement kits are required. • Please retain this letter with your laboratory records and forward this letter to those who may have received the product. | 2025-01-21 11:55
Aesculap AG |  AESCULAP Aicon® Sterile Container System / | 18223 | - | The Device is made of anodized aluminum providing significant functional and safety benefits when compared to other materials If anodized aluminum is subject to cleaning disinfection solutions outside the defined pH range and inappropriate re processing parameters water quality temperature and duration of the drying phase white residues oxygen containing aluminum compounds may form on the Device s inner and or outer surface The white residues do not influence form fit and function and the safety and efficacy of the Device | Every user is instructed to take note of the updated IFU and, in the event of occurrence of white residues on its Sterilization Containers, to review the internal reprocessing procedures accordingly. | 2025-01-21 12:40
Alshifa Medical Products Co.(CJS) | Oral Dispensing Syringe White Plunger with Tip Cap | 61776 | - | ml oral syringe supplied dark color instead of amber transparent color | To replace all the quantities delivered to the hospitals | 2025-01-21 14:32
Philips Healthcare (Suzhou) Co., Ltd. | Multiva 1.5T | 1 | - | Philips has identified a thermostability issue with the magnet bore cover used on Multiva T MR Systems The cover may partly melt where the area is not covered by the Quadrature Body Coil QBC inside the bore if an unexpected arcing failure occurs on the Gradient Coil The magnet bore cover acts as a barrier between the heat and the patient x D x D If the arcing failure occurs the system will shut down and the user may observe the following x D Smoke and or fire from within the system x D If Gradient Coil arcing occurs in a region outside of the QBC or bridge support particles may melt through the magnet bore cover and contact the patient x D x D As of December Philips has received one complaint associated with the magnet bore cover partially melting on a Multiva T System There was no report of injury or serious harm to hospital staff or patients x D | Philips will contact you to schedule time for a Field Service Engineer (FSE) to visit your site to install a hardware solution to improve the thermostability of the magnet cover (Reference FCO78100598). | 2025-01-21 15:41
Philips Healthcare (Suzhou) Co., Ltd. | Incisive CT | 1 | - | Philips received the attached Riello UPS product recall letter see Appendix A regarding specific Riello Multi Sentry Uninterruptible Power Supply UPS that Philips supplied to you as a distributor in combination with certain Philips CT system models see Figure According to the attached letter there is a potential for the UPS capacitor to fail which may lead to overheating and that it is necessary to replace the capacitor | Together with the manufacturer, Riello UPS, Philips will schedule a time for a Field Service Engineer (FSE) and a Riello representative to replace the affected products. (ref FCO72800831) | 2025-01-21 16:50
Cardinal Health | Presource Kits | 643 | - | under assessment x D | "If you locate affected product within your facility, please complete the following: o 1. REVIEW your inventory immediately, stop using the affected product code and lots attached. o 2. COMMUNICATE with all personnel who utilize the Presource® Kits containing CODMAN® Patties & Strips© component produced by Integra LifeSciences. o 3. NOTIFY any customers to whom you may have distributed/forwarded affected product or to whom you plan to send the product about this medical device product correction and share a copy of this notice. o 4. REVIEW, complete, sign and return the enclosed Acknowledgement Form in accordance with the directions on the form. o 5. RETURN all affected product to your distributor )" | 2025-01-22 09:10
Philips | Allura CV20 Allura Xper FD10 AlluraXper FD10/10 Allura Xper FD10C AlluraXper FD20 AlluraXper FD20 Biplane AlluraXper FD20 Biplane OR Table AlluraXper FD20 OR Table AlluraXperFD20/10 Allura Xper FD20/20 AlluraXperFD20/15 | 174 | - | Philips has identified that during manual repositioning of the patient tabletop a finger can get entrapped between the longitudinal guiding rails and the tabletop possibly resulting in an injury of the finger see Figure This risk is applicable for the operators and service personnel | Philips is releasing an addendum to the instructions for use and a Quick Reference Card that will be available for customers at the end of March 2025. Philips is working on the design of additional brackets for the AD7 tables, and a new redesigned table cover for AD7X tables (ref. FCOs: FCO72200516, FCO72200552, FCO72200553, FCO72200554, FCO72200555, FCO72200556). Philips will contact you to schedule a visit once the new redesigns are available. As of the date of this URGENT Field Safety Notice letter, Philips expects the solution to be available by Q1 2026. | 2025-01-22 09:42
Philips | Azurion 3 M12 Azurion 3 M15 Azurion 5 M12 Azurion 5 M20 Azurion 7 B12 Azurion 7 B20 Azurion 7 M12 Azurion 7 M20 | 174 | - | Philips has identified that during manual repositioning of the patient tabletop a finger can get entrapped between the longitudinal guiding rails and the tabletop possibly resulting in an injury of the finger see Figure This risk is applicable for the operators and service personnel | Philips is releasing an addendum to the instructions for use and a Quick Reference Card that will be available for customers at the end of March 2025. Philips is working on the design of additional brackets for the AD7 tables, and a new redesigned table cover for AD7X tables (ref. FCOs: FCO72200516, FCO72200552, FCO72200553, FCO72200554, FCO72200555, FCO72200556). Philips will contact you to schedule a visit once the new redesigns are available. As of the date of this URGENT Field Safety Notice letter, Philips expects the solution to be available by Q1 2026. | 2025-01-22 09:42
Philips | Allura CV20 Allura Xper FD10 AlluraXper FD10/10 Allura Xper FD10C AlluraXper FD20 AlluraXper FD20 Biplane AlluraXper FD20 Biplane OR Table AlluraXper FD20 OR Table AlluraXperFD20/10 Allura Xper FD20/20 AlluraXperFD20/15 | 188 | - | Philips has identified some situations that may result in the patient falling from the table related to the mattress used on the patient table of the Philips Allura and Azurion systems x D Mattress slipping from the table While the patient is being transferred from the patient table to a cart strecher hospital bed and vice versa the mattress could move and slip potentially resulting in the patient falling from the table x D Incorrect positioning of the neuro mattress on the table If the neuro mattress is positioned at the top of the table the mattress will cover the neuro table head leaving the mattress unsupported in this area The patient could fall if while positioning making comfortable the patient places their hand on the side of their head where the mattress is not supported by the tabletop see figure x D Incorrect mattress being used on table If a cardiac long mattress is used on a neuro table the mattress will cover the neuro table head leaving the mattress unsupported in this area The patient could fall if while positioning making comfortable the patient places their hand on the side of their head where the mattress is not supported by the tabletop see figure | Philips is informing customers of this issue through this Urgent Field Safety Notice. Additionally, Philips is releasing an Addendum to the Instructions for Use and a Quick Reference Card that will be available for customers at the end of March 2025. | 2025-01-22 09:52
Philips | Azurion 3 M12 Azurion 3 M15 Azurion 5 M12 Azurion 5 M20 Azurion 7 B12 Azurion 7 B20 Azurion 7 M12 Azurion 7 M20 | 188 | - | Philips has identified some situations that may result in the patient falling from the table related to the mattress used on the patient table of the Philips Allura and Azurion systems x D Mattress slipping from the table While the patient is being transferred from the patient table to a cart strecher hospital bed and vice versa the mattress could move and slip potentially resulting in the patient falling from the table x D Incorrect positioning of the neuro mattress on the table If the neuro mattress is positioned at the top of the table the mattress will cover the neuro table head leaving the mattress unsupported in this area The patient could fall if while positioning making comfortable the patient places their hand on the side of their head where the mattress is not supported by the tabletop see figure x D Incorrect mattress being used on table If a cardiac long mattress is used on a neuro table the mattress will cover the neuro table head leaving the mattress unsupported in this area The patient could fall if while positioning making comfortable the patient places their hand on the side of their head where the mattress is not supported by the tabletop see figure | Philips is informing customers of this issue through this Urgent Field Safety Notice. Additionally, Philips is releasing an Addendum to the Instructions for Use and a Quick Reference Card that will be available for customers at the end of March 2025. | 2025-01-22 09:52
Philips | Cardiovascular-Allura Centron | 188 | - | Philips has identified some situations that may result in the patient falling from the table related to the mattress used on the patient table of the Philips Allura and Azurion systems x D Mattress slipping from the table While the patient is being transferred from the patient table to a cart strecher hospital bed and vice versa the mattress could move and slip potentially resulting in the patient falling from the table x D Incorrect positioning of the neuro mattress on the table If the neuro mattress is positioned at the top of the table the mattress will cover the neuro table head leaving the mattress unsupported in this area The patient could fall if while positioning making comfortable the patient places their hand on the side of their head where the mattress is not supported by the tabletop see figure x D Incorrect mattress being used on table If a cardiac long mattress is used on a neuro table the mattress will cover the neuro table head leaving the mattress unsupported in this area The patient could fall if while positioning making comfortable the patient places their hand on the side of their head where the mattress is not supported by the tabletop see figure | Philips is informing customers of this issue through this Urgent Field Safety Notice. Additionally, Philips is releasing an Addendum to the Instructions for Use and a Quick Reference Card that will be available for customers at the end of March 2025. | 2025-01-22 09:52
Siemens Healthcare Diagnostics Inc. | Atellica CH IgM_2 | 20160 | - | Issue Description x D In December Siemens Healthineers issued Urgent Field Safety Notice ACHC A OUS communicating the potential for a negative bias with quality control QC and patient sample results when using the Atellica CH Immunoglobulin M IgM reagent Customers were instructed to batch test Atellica CH IgM samples x D Siemens is pleased to inform you that the root cause for the reagent instability has been identified and the manufacturing control system has been updated to ensure that there is no impact to future lots The issue has been resolved and beginning with lot batch testing of Atellica CH IgM samples will no longer be required Refer to the Customer Actions section below x D x D Kindly review the attached UFSN | Customer Actions • Please review this letter with your Medical Director. • If your laboratory is using a lot prior to lot 242315, you must continue to follow instructions in ACHC24-01.A.OUS. • If your laboratory has lot 242315 or above, stop following the instructions in ACHC24-01.A.OUS. • Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. • Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached UFSN | 2025-01-23 07:25
Lowenstein Medical GmbH &Co.KG | Software update for the intensive care ventilator elisa 300/500/600/800/800VIT | 201 | - | Software error with automatic patient detection APD | Operators and users can continue to use the ventilators but must ensure that alarms are responded to appropriately. Automatic Patient Detection (APD) must be deactivated until a software update is available. This will result in the disconnection alarm being triggered, but ventilation will continue. Please confirm receipt of this safety notice by returning Annex A. | 2025-01-23 09:25
Abbott Molecular Division Inc. | Alinity m SYSTEM | 1 | - | Please find attached FSN report | Please find attached FSN report | 2025-01-23 12:06
HUMAN Gesellschaft fȕr Biochemica & Diagnostica mbH | Glucose Liquicolor | 17 | - | It was estimated that only for such samples a falsely low result below the real concentration may be displayed For samples with glucose concentrations above the upper limit of the measuring range mg dL this can lead to samples being not correctly displayed as above measuring range These discrepancies were only observed in a very few single instruments | To fix these discrepancies, HUMAN has already developed an optimized setting, which should be installed immediately on all HumaStar 300SR instruments on which GLUCOSE liquicolor is measured | 2025-01-24 15:21
Abbott GmbH. & Co. KG | Alinity s System | 3 | - | Please find attached FSN letter | Please find attached FSN letter | 2025-01-26 13:48
Maquet Cardiopulmonary GmbH | Heater-Cooler Unit HCU 40 High Voltage | 12 | - | The manufacturer became aware of this issue in course of complaints While earlier complaints were x D detected by visual detection during maintenance or repair one recently reported low flow during cleaning x D process another one too high water pressure due to particles in the oxygenator during treatment Upon x D inspection white crystalline powder was confirmed on all Ice Bracket indicating progressive corrosion Figure x D x D An internal investigation came to the conclusion that the failure rate of the Ice Bracket becomes unacceptably x D high after two years | For devices where the Ice Bracket was replaced more than 2 years ago: • The Ice Bracket shall be replaced as soon as possible free of charge. For devices where the Ice Bracket was replaced less than 2 years ago: • The Ice Bracket shall be replaced during the next preventive maintenance free of charge. In the future (not part of this Field Action): • The Ice Bracket will be replaced biennially during preventive maintenance. | 2025-01-26 14:41
LivaNova USA, Inc. | VNS Therapy® Model 1000 SenTiva™ and Model 1000-D SenTiva Duo™ | 25 | - | Premature End of Service x D LivaNova issuing this voluntary safety notification to hospitals and treating physicians that implanted Model and Model D generators may stop delivering therapy due to an internal component issue The microcontroller is a component inside the VNS Therapy implantable pulse generator that is responsible for controlling the functions of the generator LivaNova has identified that generators produced using microcontrollers from certain supplier manufacturing lots may experience shorter than expected device longevity Please refer to attached customer letters for more details | Please ensure the following actions are completed urgently: 1. Read attached customer letter for complete understanding of the issue. 2. Report to Health Authority: Evaluate if this field action required reporting to Health authority in your respective region/ country. If yes, complete the submission and share a copy of communication as evidence with LivaNova representative for our records. 3. Return of Unused Units: • Identify and isolate any affected unused units in your inventory. • Return all affected units to LivaNova according to the return instructions attached to this email. 4. Inform End Users/Healthcare Providers: Hospital Communication: • Notify hospitals and healthcare providers that may have received affected devices and provide them with the necessary instructions for reporting and returning unused units. • Reassure healthcare providers that implanted devices do not require explanation and can remain in use. Informing Treating Physician: • Communicate with the attending physicians to ensure they are aware of the recall and the safety notice information provided. • Provide doctors with any support or information they might need to communicate the situation to their patients effectively. 5. Document the Communication: • Confirm receipt and understanding of this notice through the attached Distributor Response Form. • Every communication with treating physicians, hospitals, or end users must be documented thoroughly. This includes the date, time, method of contact, and outcome of the communication. The signed Customer Response Forms should also be collected and securely stored for future audits or follow-ups • Keep a record of the communication sent to end users and healthcare providers for your own records and compliance purposes. • Please return all signed customer responses forms from the hospital and treating physicians, as soon as action is closed for all customers. 6. Monitor and Report: • Continue to monitor the use of the device and report any adverse events or incidents related to the affected product • Should any new information arise, that would necessitate further action, we will update you accordingly. 7. Adverse Event Reporting: • Please remind hospitals, healthcare providers, and distributors to report any adverse events or concerns regarding the affected implantable generators to LivaNova as soon as possible. | 2025-01-26 15:31
LivaNova USA, Inc. | VNS Therapy® Model 1000 SenTiva™ and Model 1000-D SenTiva Duo™ | 463 | - | After a thorough evaluation we have identified a potential safety concern related to small percentage of implanted Model and Model D generators may stop delivering therapy due to a reed switch issue The reed switch is a component inside the VNS Therapy implantable pulse generator that senses changes in magnetic field in order to provide on demand stimulation Epilepsy indication only or to stop stimulation However based on our risk analysis there is no immediate risk to patient safety and implanted devices do not need to be explanted They can continue to be used as intended x D x D Furthermore there are no restrictions on the continued marketing and distribution of this device The affected devices can still be marketed and used in compliance with all applicable safety standards x D | We request that you take the following actions in response to this Field Safety Notice: 1. Read attached customer letter for complete understanding of the issue. 2. Report to Health Authority: Evaluate if this field action required reporting to Health authority in your respective region/ country. If yes, complete the submission and share a copy of communication as evidence with LivaNova representative for our records. 3. Identify the list of consignees, hospital, end user as well as treating physicians 4. Inform End Users/Healthcare Providers: • Communicate this safety notice to all consignee hospital, healthcare providers and end users who may be using or receiving the affected device(s). • Ensure that they are aware that the device does not pose an immediate risk to patients when used as intended. 5. Document the Communication: • Confirm receipt and understanding of this notice through the attached Distributor Response Form. • Every communication with treating physicians, hospitals, or end users must be documented thoroughly. This includes the date, time, method of contact, and outcome of the communication. The signed Customer Response Forms should also be collected and securely stored for future audits or follow-ups. • Keep a record of the communication sent to end users and healthcare providers for your own records and compliance purposes. • Please return all signed customer responses forms from all hospital and treating physicians, as soon as action is closed for all customers. 6. Monitor and Report: • Continue to monitor the use of the device and report any adverse events or incidents related to the affected product. • Should any new information arise, that would necessitate further action, we will update you accordingly. 7. No Need for Device Removal or Explanation: • As mentioned, implanted devices do not require explanation or removal. There is no need for any further action related to devices already implanted. Please follow instructions and details given in the attached customer letters to Hospital and treating physician/ doctor. | 2025-01-26 15:44
Biosense Webster | CARTO® OCTARAY™ Mapping Catheter with TRUEref™ Technology | 31 | - | Between September and October Biosense Webster received two complaints reporting that the OCTARAY Mapping Catheters showed signs of charring on the splines and electrodes after being used at the same time as ablation catheters The procedures associated with these two complaints were completed successfully with no patient consequences Biosense Webster investigated these complaints and determined that the most probable cause for the char in the mapping catheter is that the ablation catheter used in the procedure was in contact with the OCTARAY Mapping Catheter rings during radio frequency RF ablation x D Char formation is not expected in diagnostic mapping catheter rings because char is a physical phenomenon of RF and mapping catheters do not deliver RF energy To avoid char formation on the rings of this mapping catheter do not apply RF energy when the ablation catheter is in contact with one or more rings of the mapping catheter x D You can continue using CARTO OCTARAY Mapping Catheter with TRUEref Technology as directed | Please review this letter carefully and share it with any of your staff involved in such procedures. If you have additional questions about this letter, please contact your Biosense Webster representative. | 2025-01-26 16:41
Randox Laboratories Ltd | ABAXIS CHEMISTRY CONTROL LEVEL 1 | 8360 | - | Randox Laboratories can confirm the analyte Uric Acid in the mentioned lots of Abaxis Chemistry Control Leve and has decreased over the shelf life of the product and requires the targets and ranges to be reassigned | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Discard of current value sheet for 1627UNCM and 1300UECM and refer to updated value sheet provided. | 2025-01-27 09:53
Randox Laboratories Ltd | ABAXIS CHEMISTRY CONTROL LEVEL 2 | 8360 | - | Randox Laboratories can confirm the analyte Uric Acid in the mentioned lots of Abaxis Chemistry Control Leve and has decreased over the shelf life of the product and requires the targets and ranges to be reassigned | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Discard of current value sheet for 1627UNCM and 1300UECM and refer to updated value sheet provided. | 2025-01-27 09:53
Dexcom, Inc. | Dexcom G6 Receiver | 2855 | - | Product needs software update | Recommended Customer Action: If you are or plan to be an active user of an affected Dexcom G6 Receiver, please complete the following actions related to an important software update: 1. Please navigate to the following webpage on your internet browser using your tablet or computer: https://www.dexcom.com/g6-receiver-update 2. Please read the information provided on the webpage and any recommended actions regarding a software update for your receiver to permanently resolve this issue. If you are not and do not plan to be an active user of an affected Dexcom G6 Receiver, no further additional action is needed. Examples of where this may apply include if you have switched to a non-G6 CGM product or if you have permanently discontinued use of your G6 receiver in favor of a different display device. | 2025-02-02 10:08
Beckman Coulter Inc. | Access EPO | 13 | - | Beckman Coulter has determined that the Access Erythropoetin EPO reagent lot number listed in the table above exhibits an approximate concentration bias of with patient samples when compared to alternate EPO reagent lots |  Discontinue using and discard all remaining reagent packs of the Access EPO reagent lot listed above.  Beckman Coulter recommends sharing the content of this letter with your laboratory and/or medical director regarding the need to review previous patient test results.  Contact your local Beckman Coulter representative for replacement product requests. | 2025-02-02 12:20
medimaps group SA | TBS iNsight | 3 | - | The issue was recently discovered by conducting a preventive internal investigation The preventive action was triggered by a suspicious result obtained in off label and non clinical use of the TBS iNsight device The variability in calculations occurs due to the distinct scan acquisition method used in fast array mode compared to array mode This method can introduce differences in captured data particularly in patients with extreme BMI or significant weight changes and is specific to the fast array process on these devices The issue does not impact scans acquired in array mode or the Hologic Horizon device itself Further the issue does not impact GE or Osteosys DXA devices | Action To Be Taken by the User To reduce variability in calculations from fast array scans on Hologic Horizon machines and mitigate the newly identified risk, clinicians are advised to consider using the standard array mode for TBS calculations, where feasible, until the update becomes available. The forthcoming software maintenance release will incorporate an enhanced correction factor for fast array scans, ensuring a more tailored correction between the two scan modes and greater accuracy across the BMI range. Action Being Taken by the Manufacturer Medimaps will issue a maintenance release of the software by the end of February in the worst case and deploy it progressively in the field. The exact plan will be further communicated to customers. | 2025-02-02 13:50
Beckman Coulter Inc. | Access 2 Immunoassay System | 19 | - | Beckman Coulter has determined that the sample motor on the Access Immunoassay Analyzers listed in the appendix may have an improper alignment to the sample wheel that may result in irregular motion as the sample carousel turns | • If the error occurs: 1. Confirm the system is in the Ready or Not Ready mode. 2. Go to the Initialize System window. To get to this window from the Main Menu, select Diagnostics F7 to display the Diagnostics menu, then select Initialize System F1. The Initialize System window is displayed with a confirmation message. 3. Select Yes F1 to initialize the instrument. • Contact your Beckman Coulter representative if your laboratory has experienced sample carousel motion errors that are not resolved by initializing the analyzer | 2025-02-02 21:33
Siemens Shenzhen Magnetic Resonance Ltd.(SSMR) | Artis one | 2 | - | In very rare situations the first x ray release after patient registration might be performed with a wrong x D copper prefiltration | A software update for the affected systems is planned to correct the issue and ensure the correct function of the prefiltration. | 2025-02-03 14:06
GE Healthcare (Tianjin) Company Limited | SIGNA CREATOR | 13 | - | GE HealthCare has become aware that for certain MR systems a scan can be resumed following a Power Monitor trip when the minute average SAR is above the limit indicated on the SAR Display x D If this error occurs the Start Scan button on the Scan Control Interface Module SCIM may become inappropriately illuminated and the user interface screen may incorrectly display Please press start scan button enabling a scan to proceed earlier than intended x D If a scan is resumed by the operator under these conditions it can result in potential excessive tissue heating particularly if a patient were to be in contact with a conductive material or have an implant x D There have been no injuries reported to GE HealthCare as a result of this issue | Provide and install updated Application Software versions. | 2025-02-03 14:07
GE Healthcare (Tianjin) Company Limited | Signa Explorer | 13 | - | GE HealthCare has become aware that for certain MR systems a scan can be resumed following a Power Monitor trip when the minute average SAR is above the limit indicated on the SAR Display x D If this error occurs the Start Scan button on the Scan Control Interface Module SCIM may become inappropriately illuminated and the user interface screen may incorrectly display Please press start scan button enabling a scan to proceed earlier than intended x D If a scan is resumed by the operator under these conditions it can result in potential excessive tissue heating particularly if a patient were to be in contact with a conductive material or have an implant x D There have been no injuries reported to GE HealthCare as a result of this issue | Provide and install updated Application Software versions. | 2025-02-03 14:07
GE Healthcare (Tianjin) Company Limited | Brivo MR355 | 13 | - | GE HealthCare has become aware that for certain MR systems a scan can be resumed following a Power Monitor trip when the minute average SAR is above the limit indicated on the SAR Display x D If this error occurs the Start Scan button on the Scan Control Interface Module SCIM may become inappropriately illuminated and the user interface screen may incorrectly display Please press start scan button enabling a scan to proceed earlier than intended x D If a scan is resumed by the operator under these conditions it can result in potential excessive tissue heating particularly if a patient were to be in contact with a conductive material or have an implant x D There have been no injuries reported to GE HealthCare as a result of this issue | Provide and install updated Application Software versions. | 2025-02-03 14:07
GE Healthcare (Tianjin) Company Limited | Optima MR360 | 13 | - | GE HealthCare has become aware that for certain MR systems a scan can be resumed following a Power Monitor trip when the minute average SAR is above the limit indicated on the SAR Display x D If this error occurs the Start Scan button on the Scan Control Interface Module SCIM may become inappropriately illuminated and the user interface screen may incorrectly display Please press start scan button enabling a scan to proceed earlier than intended x D If a scan is resumed by the operator under these conditions it can result in potential excessive tissue heating particularly if a patient were to be in contact with a conductive material or have an implant x D There have been no injuries reported to GE HealthCare as a result of this issue | Provide and install updated Application Software versions. | 2025-02-03 14:07
Micro Therapeutics, Inc. | Pipeline Vantage Embolization Device with Shield Technology | 463 | - | Description of Issue x D Field reports of serious adverse events with braid deformation triggered a quality investigation on Pipeline Vantage This investigation is focused on braid deformation and or incomplete wall apposition issues including incomplete failure to open fishmouth collapse twist flatten ribbon hump and migration The hazards under investigation include x D Hazard Intra procedure braid deformation incomplete wall apposition x D Hazard Post procedure braid deformation incomplete wall apposition x D x D A review of the available data from the INSPIRE A registry and post market surveillance data i e complaints was conducted Performance of Pipeline Vantage was compared against the previous generation Pipeline Shield flow divertor as appropriate to assess relative performance The rate of occurrence and complaint rate for Post procedure braid deformation incomplete wall apposition is statistically significantly higher with Pipeline Vantage as compared to Pipeline Shield The Pipeline Vantage compatible sizes are similar to the Pipeline Shield product family in design characteristics As observed in INSPIRE A registry the rate of braid deformation for the Pipeline Vantage is lower than that observed for Pipeline Vantage x D As observed in the INSPIRE A registry Pipeline Vantage devices diameters mm appear to exhibit a higher incidence stent braid deformation compared to the Pipeline Shield For Pipeline Vantage devices fewer reports were received for incomplete wall apposition and or braid deformation Based on this information the retrieval is only isolated to unused inventory of the Pipeline Vantage devices x D | Advice on actions to be taken by the distributor and the user: Patient Management Considerations: The need for follow up imaging and/or changes in medical management should be made by the treating physician according to accepted guidelines, taking into consideration the patient’s overall health. The risks of dual antiplatelet therapy (DAPT) should be weighed against the potential risk posed by braid deformation. Required Actions for Impacted Product: Our records show that your facility has received one or more lots of the impacted products which includes all units with model number PED3-027-XXX-XX and PED3-021-XXX-XX. Consequently, Medtronic requires that you immediately take the following actions: 1. Do NOT use any impacted Pipeline Vantage 027 product listed above. Remove and quarantine all unused impacted products listed in Appendix C from your inventory. 2. Return the impacted products to Medtronic as per the instructions outlined in the Customer Acknowledgment form. Your Medtronic representative can assist in facilitating the return of product as necessary. If alternative product is needed, your Medtronic representative can assist you with identifying suitable replacement product. Medtronic has taken the necessary steps to prevent future shipment of the impacted Pipeline Vantage 027 product. 3. Medtronic is implementing changes to the IFU for the Pipeline Vantage 021 with part numbers PED3-021-XXX-XX. The purpose of these changes is to help achieve optimal size selection and stent braid deployment to reduce the risk of complications and patient harms. The key updates are: • Appropriate device diameter and length selection to account for complex anatomy. • Techniques to deploy Pipeline Vantage compared to Pipeline Shield using a balance of device tension and compression to achieve adequate wall apposition and landing around curves. • Warnings about the consequences of incomplete wall apposition and suboptimal deployment and the increased risk of braid deformation in females, especially in females ≤45 years. A copy of the IFU changes is enclosed with this letter. We strongly recommend following the redlined/proposed IFU language while we work through the regulatory approval process to make changes to IFU permanently. Please ensure the updated IFU is used when completing any future procedure with the Pipeline Vantage device. • Complete the enclosed Customer Acknowledgement Form and email to Medtronic . | 2025-02-03 16:50
Ortho Clinical Diagnostics | ORTHO Optix™ Reader | 2 | - | The Mixed Field MF Grade Error Code applies to assays where Mixed Field detection is enabled i e red cell antigen typing including User Defined Assays UDA The Mixed Field Grade Error Code appears instead of a reaction grade when the Image Processing System IPS detects a Mixed Field reaction pattern which an operator can review and modify to a grade or result interpretation as applicable based on their visual review of the cassette image x D x D QuidelOrtho received a customer complaint concerning the MF Grade Error Code not generating while another Grade Error Code Fibrin FIB was posted instead x D x D | • Disable the Automatically Accept Results Feature, following the directions provided above, until Software Version 2.0.2 (MOD 8) has been installed on your ORTHO Optix Reader. • Complete and return the enclosed Confirmation of Receipt form no later than February XX,2025. • Please forward this notification if the affected product was distributed outside of your facility | 2025-02-06 09:01
Medtronic, Inc. | DLP® Aortic Root Cannula with Flow-Guard™ | 3140 | - | During the manufacturing process at a Medtronic location an operator identified defects appearing to be gouges and loose plastic material similar to flash inside the inner diameter of the male luer used in the aortic root cannula While the loose plastic material has been determined to not be a result of a molding defect it exceeds the flash specification Medtronic has identified this issue with devices from supplier lots | Medtronic requests that you take the following actions: • Review your inventory for listed product using attachment A. • Immediately identify and quarantine all unused, listed product in your inventory. • Return unused, listed product in your inventory to Medtronic. Your Medtronic representative can assist you in the return of affected product as necessary. • Complete the enclosed Customer Acknowledgement Form and email to Medtronic . This form must be returned even if you do not have any affected product in your possession. • Please share this notification with others in your organization as appropriate. If product listed above has been forwarded to another facility, please notify the facility of this Medtronic Urgent Field Safety Notice. • Please maintain a copy of this communication in your records. | 2025-02-06 09:19
Medtronic Minimed | MiniMed 740G | 10564 | - | Medtronic Research and Development R D Team initiated a delivery accuracy dry run testing of the G NGP infusion pump under dynamic environmental use conditions temperature and pressure in preparation for the MiniMed G system Alternate Controller Enabled ACE submission x D During the dry run dynamic pressure condition testing it was observed that when air pressure decreased during take off simulation kPa to kPa in minutes there was an average of U of delivery volume observed i e U more than expected Based on the U h delivery rate the expected delivery volume in minutes at ambient pressure kPa is U Conversely when air pressure was increased during landing simulation i e kPa to kPa in minutes it was observed that the average volume delivered was U i e potential no delivery Based on the U h delivery rate the expected delivered volume in minutes at ambient pressure kPa is U x D Additionally it was found that unintended insulin delivery during take off simulation is independent of delivery rate i e a pump with zero U h programmed delivery would still deliver insulin during take off The probable root cause of this issue has been attributed to the physical phenomena i e the decrease in air pressure leading to the formation of air bubbles in the reservoir which in turn affects the delivery volume by displacing the insulin out of the reservoir x D | • Monitor Your Glucose Levels: Check your glucose frequently during activities like air travel, amusement park rides, or other situations where sudden changes or extremes of air pressure, altitude, or gravity may occur. • Discuss how to prepare for situations like this with your healthcare professional. Keep an emergency kit with rapid-acting glucose and backup insulin therapy available at all times. • Respond to Alerts and Symptoms: Pay attention to any alerts from your pump as well as symptoms of hypoglycemia or hyperglycemia. Follow your healthcare professional's treatment instructions in these situations. | 2025-02-06 17:07
Aesculap AG | AESCULAP® Haemostatic forceps - Haemostatic clamps | 142654 | - | According to market feedback the Haemostatic forceps have been used according to indications which are currently not covered by the design e g larger blood vessels or fixation of medical materials leading to customer complaint cases related to the breakage of the device s during use caused by the failure mode breakage of the box lock | The user is requested to promptly implement the new product-specific Instructions for Use (IFU), which is attached to this FSN, and to ensure the compliant use of the device(s) at patients. Additionally, the user must immediately implement guidelines for verifying the safety and performance of the device(s) within the context of lifecycle monitoring during the reprocessing routine | 2025-02-09 11:35
Olympus | Forceps/Irrigation Plug (Isolated Type) | 86 | - | Olympus investigation into reported infection and contamination complaints associated with the MAJ revealed a potential risk of patient infection due to inadvertent retention of biomaterial when the reprocessing procedure for the MAJ is not performed as instructed such as not disconnecting the MAJ from the endoscope and disassembling it before reprocessing x D | 1. Carefully read the content of this letter. 2. Use alternative devices to the MAJ-891 according to the compatibility table above. 3. Ensure all personnel are completely knowledgeable and thoroughly trained on the content of this letter. Specifically, please ensure that all reprocessing personnel are completely knowledgeable and thoroughly trained on the MAJ-891 reprocessing instructions. 4. Olympus will contact you to schedule a mutually beneficial time for an Olympus Field Representative to visit your facility and review the contents of this letter and the MAJ-891 reprocessing instructions. Refer to the “Training and Resources” section above for additional resources. 5. If you have further distributed this product, identify your customers, and forward them this letter. 6. Olympus requests that you acknowledge receipt of this letter. Indicate on the Reply Form that you have received and understood this letter by filling out and returning the completed enclosed Reply Form back to your local Olympus representative. | 2025-02-09 14:46
Advanced Medical Solutions Limited | SILVERCEL® Non-Adherent Hydro-Alginate Antimicrobial Dressing | 1100 | - | AMS has become aware of a defect on primary packaging pouches in which minor missing patches of Polyethylene have x D been detected The defect on these pouches could compromise the device s ability to maintain a sterile barrier In x D addition patches of burnt or cracked polyethylene have been identified on the inside face of the primary packaging x D | FSN requesting customers to inspect their inventory, destroy any devices where the packaging presents defect and provide a destruction certificate to distributor/manufacturer. | 2025-02-10 12:57
Saudi Mais Medical Products | Oral Care set | 9177 | - | product does not have Medical Device Marketing Authorization MDMA and Suspected contamination in some of the Antiseptic Oral Rinse | Stop the distribution of stock of the affected product and Quarantine immediately, After getting the MDMA certificate We will be replacing the remaining effected quantities with the product with a new Lot. | 2025-02-10 16:30
ABBOTT S.r.l. | AliniQ AMS | 43 | - | Please find attached the FSN letter | Please find attached the FSN letter | 2025-02-11 10:57
Beckman Coulter Inc. | DxC 500 AU clinical Chemistry Analyzer | 1 | - | Beckman Coulter has determined that the system cannot recalculate a calculated result when results are rerun within the calculated test |  The calculated results can be recalculated manually offline by users based on the formula defined and rerun results from constituent tests.  Configure calculated tests within LIS/Middleware, if available.  Beckman Coulter recommends sharing the content of this letter with your laboratory and/or Medical Director. | 2025-02-12 23:09
GE Healthcare | Centricity Universal Viewer | 62 | - | GE HealthCare has become aware of a potential security vulnerability in Centricity Universal Viewer UV Centricity PACS IW PACS IW when products are configured on the public internet e g without a secure private internet connection such as a virtual private network VPN which could allow a malicious actor to access the system and potentially manipulate patient data x D There have been no injuries or unauthorized access to patient data reported to GE HealthCare as a result of this issue | Unsafe configuration – Software fix Safe configuration – Prevention fix A qualified GE HealthCare employee/distributor will perform a site inspection to confirm if they require internet access without a secured network (VPN). If so, a software correction for Centricity Universal Viewer (UV) and PACS-IW products will be applied. For all other customers with affected software, a qualified GE employee/distributor will confirm product configuration to prevent outside access to their systems. | 2025-02-13 10:21
Sedana Medical Ltd | Sedaconda Syringe | 5 | - | The identified issue with the syringe to line connection may result in leakage of Isoflurane x D leading to insufficient delivery of the anaesthetic This could cause inadequate sedation resulting x D in the patient waking prematurely If not promptly addressed this may pose a risk of patient x D discomfort psychological stress or in severe cases unintended movement or self extubation x D potentially compromising patient safety | Affected customers are provided with immediate access to the updated IFU and e-guide. Instructions For Use Sedaconda ACD Rev 5, specifically 5.3 Point 6 and Instructions For Use Sedaconda Syringe Rev 1, specifically Picture 3. | 2025-02-13 16:54
Abbott | AssurityTM and EndurityTM pacemakers | - | - | Abbott is informing clinicians of the potential for a device malfunction affecting a subset AssurityTM and EndurityTM x D pacemakers This issue may result in incomplete mixing of epoxy during manufacturing and with time may permit x D moisture ingress into the pulse generator header introducing a risk of interrupting device functionality Affected x D devices were manufactured between August and June The specific manufacturing equipment associated x D with this issue is no longer in use No affected devices remain available for implant x D There have been no reports of permanent harm to patients resulting from this issue x D The observed rate through Abbott s post market surveillance is The mean implant duration at time of failure is x D currently years with a standard deviation of years Reported clinical impact has included loss of telemetry x D communication reduced battery longevity and or premature battery depletion and or loss of pacing x D As Abbott records indicate you are following one or more patients implanted with a potentially affected device noted in x D the enclosed Device List please reference the patient management recommendations below | Patient Management Recommendations: Recognizing that each patient requires individual consideration by their physician, in consultation with Abbott Cardiac Rhythm Management (CRM’s) Medical Advisory Board (MAB), Abbott provides the following guidelines: • Prophylactic generator replacement is not recommended due to the low rate of occurrence of this issue. • Routine follow-up should remain as per standard of care. Enroll patients on Merlin.net when possible, and consider increasing the frequency of scheduled interrogations in patients with non-RF enabled pacemakers. Review device function, including measured battery voltage, any unexpected change in battery consumption, and connectivity status on Merlin.net where available. Also, evaluate the potential for risk in patients who are pacemaker dependent, particularly if they are unable to be reliably followed using remote monitoring. • Prompt replacement for devices that demonstrate unexpected depletion to ERI/EOS, trigger an EPI notification, or demonstrate one of the clinical impacts listed above. As always, timing of replacements should be appropriate for the patient’s underlying clinical condition | 2025-02-16 16:17
Smiths Medical ASD, Inc. | Portex Tracheal Tube | 43874 | - | Smiths Medical became aware that certain sizes of Intubation ORAL NASAL Endotracheal Tube products may have a smaller diameter than expected | Customer Required Actions There is no need to remove an affected Endotracheal tube if it is already in use. However, please review the Instructions for Use with heightened awareness prior to continued use 1. Check all inventory locations within your institution for the affected products listed in the notification and discontinue use. Destroy all affected products following your institution’s process for destruction. If destroying is not immediately possible at your facility, then the product should be quarantined until disposal. 2. Share this notification with all potential users of the device, to ensure they are aware of this notification. If the devices are used at another location, please ensure this communication is delivered there. 3. Complete and return the attached Customer Response Form to EMEA-FSN@ | 2025-02-16 17:00
Abbott Molecular Division Inc. | Alinity m HIV-1 AMP Kit | 13 | - | Please find the attached FSN letter | Please find the attached FSN letter | 2025-02-17 11:28
GE Medical Systems, LLC | Revolution CT | 3 | - | GE HealthCare has become aware of a potential issue on certain Revolution CT and Revolution Apex series CT systems please see affected product details and the Appendix that can result in coolant fluid glycol leaking onto the floor near the base of the gantry The fluid is visible as it has a blue color It is not corrosive and is not at a harmful temperature However there is a potential for injury if the fluid is not noticed and a person slips and or falls x D x D There have been no injuries reported as a result of this issue x D | Issue a customer letter and correct the affected DTS hardware for impacted systems in the installed base by adding a hose restraint bracket and associated fastening hardware. | 2025-02-18 11:29
GE Medical Systems, LLC | Revolution Apex | 3 | - | GE HealthCare has become aware of a potential issue on certain Revolution CT and Revolution Apex series CT systems please see affected product details and the Appendix that can result in coolant fluid glycol leaking onto the floor near the base of the gantry The fluid is visible as it has a blue color It is not corrosive and is not at a harmful temperature However there is a potential for injury if the fluid is not noticed and a person slips and or falls x D x D There have been no injuries reported as a result of this issue x D | Issue a customer letter and correct the affected DTS hardware for impacted systems in the installed base by adding a hose restraint bracket and associated fastening hardware. | 2025-02-18 11:29
ORGENTEC Diagnostika GmbH | Anti-TPO | 21 | - | Following an internal performance evaluation it has been confirmed that performance claims in the current x D Instructions for Use IFU are not met x D The outcome of an internal evaluation showed a clinical Sensitivity of instead of as claimed in the x D IFU x D The Overall Agreement is instead of as claimed in the IFU x D For this reason ORGENTEC Diagnostika GmbH has decided to discontinue the product | we notified our customers about FSN and waiting for their corrective action. | 2025-02-18 15:04
Covidien llc | Microstream Advance Adult-Pediatric Intubated CO2 Filter Line | 4600 | - | In March during internal complaint review it was found that complaints were reported from one hospital in the United Kingdom regarding Neonatal Infant Intubated CO Filter Lines unable to be disconnected from endotracheal tube ETT when the neonatal patient needed suctioning using open suction system Due to the immediate need to suction the providers emergently extubated performed suctioning and re intubated the patients when the procedure was complete No further complications were reported x D The Neonatal Infant Intubated CO Filter Line interfaces with the endotracheal tube ETT via the Neonatal Airway Adapter NAA at the end of the CO Filter Line x D The observed severity for the complaints reported is higher than what is currently predicted in the risk file Upon initial review of the risk file it was determined the risk file had only previously considered non emergent scenarios The facility reporting the complaints indicated they were unable to disconnect the adapter from the ETT therefore the care team made the decision to urgently extubate the patient to allow for needed suctioning Because the patient is unable to maintain their own airway independently reintubation was required to prevent permanent impairment x D The product families affected by the investigation are Advance Luer intubated filter lines in production in field Legacy intubated filter lines no longer in production but still available in inventory and offered for sale and Legacy intubated filter lines no longer in production inventory nor offered for sale but still may be available for use in the field Within the intubated Filter Lines the two populations affected by the issue are the Neonatal Infant Intubated and the Adult Pediatric Intubated x D | • Continue using the Microstream™ Advance Neonatal-Infant and Adult Pediatric Intubated CO2 Filter Line, Microstream™ Luer Adult-Pediatric Intubated CO2 Filter Line and Microstream™ Luer Adult-Pediatric Airway Adapter Sampling Line devices. • Pass and post this notice for all those who need to be aware within your organization or to any organization where the potentially affected product has been transferred or distributed. • Please complete and return the enclosed Customer Acknowledgement Form to Medtronic. | 2025-02-20 10:45
HAMILTON MEDICAL AG | HAMILTON-C3 | 67 | - | Affected HAMILTON C C devices which are used for long term ventilation in x D neonatal patient group will switch to the Ambient State if the ventilation time is x D exceeding days | For HAMILTON-C3: Please conduct the Software Update and upgrade the affected devices to SW version 2.0.11 | 2025-02-25 16:47
Immucor GTI Diagnostics , Inc. | LIFECODES PF4 Enhanced | 6 | - | Please find attached FSN | Please find attached FSN | 2025-02-26 11:46
Covidien llc | Shiley Adult Flexible Tracheostomy Tube with TaperGuard Cuff, Reusable Inner Cannula | 407 | - | Medtronic has received an increase in the number of complaints for flange disconnection from outer cannula against the Shiley Adult Flexible Tracheostomy Tube with TaperGuard Cuff Reusable Inner Cannula devices were confirmed for reports of flange detachment from their outer cannulas through the evaluation of physical samples and or photographic evidence Currently the root cause of the reported issue is undetermined and will require further investigation | Patient Management: For patients with the affected lot of Shiley™ Adult Flexible Tracheostomy Tubes with TaperGuard™ Cuff Reusable Inner Cannula (lot 202405258X) currently in place, a replacement is required. The patient’s medical team should assess the overall patient risk when considering the timing of replacement. Clinicians should continue to follow current product Instructions For Use (IFU) along with facility specific policies and procedures. Additional actions you should take: • Quarantine all unused product from the affected lot of Shiley™ Adult Flexible Tracheostomy Tube with TaperGuard™ Cuff Reusable Inner Cannula • Return all unused product from the affected lot in your inventory to Medtronic as described on the Customer Acknowledgment Form. • Pass on this notice to all those who need to be aware within your organization or to any organization where the potentially affected product from the specified lot has been transferred or distributed. • Please complete and return the enclosed Customer Acknowledgment Form even if you do not have unused inventory. | 2025-02-26 13:56
Smiths Medical ASD, Inc. | ProPort | 17 | - | Potential Risk x D To date Smiths Medical has received two reports of serious injury and zero deaths associated with this issue If the defect is not identified during preparation identification of the separated port during implantation may result in delay of procedure or air introduced during implantation of the device The IFU provides instructions to reduce air embolisms and includes a flow check to examine for leaks during x D Nathan Lane North x D Minneapolis MN x D FA EN Page of x D implantation After the portal and catheter are implanted catheter tip location should be verified using fluoroscopy or x ray to assure stability of the system If the separation of the port housing occurs during use the separation of the port housing from the reservoir could lead to a fluid leakage during therapy This may result in extravasation potentially causing localized tissue irritation or tissue damage The IFU provides instruction for examining the portal pocket for swelling tenderness or infection which might indicate system leakage If system leakage is suspected radiologic imaging is recommended to determine if there are problems with the system x D Affected Product x D Our records indicate that your facility may have received potentially affected products Refer to Table below for a list of affected devices and lot numbers | Smiths Medical Actions: Smiths Medical is sending this notification to all customers who received affected product as listed above. Smiths Medical will provide full credit to affected customers. NOTE: Credits for product purchased through a distributor will be credited by the distributor. Recommended Actions for Healthcare Providers: a) As instructed in the IFU, continue to monitor patients who have an implanted ProPort™ Implantable Ports for signs of any adverse events. b) As instructed in the IFU, ensure the housing and reservoir feel secure and stable when palpating the portal. Symptoms such as swelling, redness, or discomfort at the implant site may indicate leakage or system failure. Customer Required Actions: 1) Check all inventory locations within your institution for the impacted lot numbers listed in the notification (Appendix A) and discontinue use. Destroy all affected products following your institution’s process for destruction. If destroying is not immediately possible at your facility, then the product should be quarantined until disposal. 2) Share this notification with all potential users of the device, to ensure they are aware of this notification. If the devices are used at another location, please ensure this communication is delivered there. 6000 Nathan Lane North Minneapolis, MN 55442 FA2501-02 (EN) Page 3 of 6 3) Complete and return the attached Customer Response Form to EMEA-FSN@icumed.com within 10 days of receipt to acknowledge your understanding of this notification, even if you do not have the affected product. | 2025-02-27 17:04
Beckman Coulter Inc. | DxC 500i Clinical Analyzer | 3 | - | Beckman Coulter has determined that the system cannot recalculate a calculated result when results are rerun within the calculated test |  The calculated results can be recalculated manually offline by users based on the formula defined and rerun results from constituent tests.  Configure calculated tests within LIS/Middleware, if available.  Beckman Coulter recommends sharing the content of this letter with your laboratory and/or Medical Director. | 2025-03-02 11:08
Beckman Coulter Inc. | DxC 500 AU Clinical Chemistry Analyzer | 1 | - | Beckman Coulter has determined that the DxC AU Clinical Chemistry Analyzer might assign a test result to an incorrect sample ID under the following workflow conditions x D After loading a barcoded sample tube Sample A onto the DxC AU Clinical Chemistry Analyzer all tests ordered for that barcoded sample tube are locked into the specific position of that barcoded tube on the analyzer for example the position on the DxLab Rack or position in the STAT table until all tests are completed or cancelled for that sample tube x D A problem might occur if for any reason all tests ordered for that barcoded sample tube Sample A are not completed or remain in pending and the operator attempts to analyze a second sample tube Sample B in that same position on the rack or STAT table If the bar code for Sample B is not successfully scanned by the analyzer or if the sample tube is not labelled the analyzer assumes the sample that previously occupied that position has been reloaded Sample A The analyzer will complete and report the pending tests associated with Sample A using Sample B No results will be reported for Sample B x D This issue does not arise if all barcoded sample tubes are successfully scanned or if testing is complete for the first barcoded sample Sample A |  Beckman Coulter recommends reviewing your laboratory workflow, sample labeling/label integrity and positioning to determine impact and sharing this with your laboratory and/or Medical Director.  The latest available system software (software V1.4.3) provides improved sample event information and real time reporting, identifying the specific location of the sample with tests that are not completed. This software will help in the management of this issue.  For current software version is V1.3, V1.4, V1.4.1, or V1.4.2 customers, contact your Beckman Coulter service representative at your earliest convenience to upgrade to software V1.4.3. To check the software version, follow the instructions in the Viewing Analyzer Version Information topic in the IFU:  Follow the attached appendix A Job Aid, which will prevent the operator workflow described in the ISSUE section. The Job Aid will outline how to use the sample list to manage the samples which have not completed. There is no requirement to reverify applications, for example, precision studies. | 2025-03-02 15:22
Welch Allyn, Inc | DuraShock | 10845 | - | Welch Allyn non automated blood pressure gauges may not meet leak and accuracy specifications following exposure to high temperature storage Baxter recommends storing the devices at a temperature not exceeding the operating temperature of C F | 1. Customers may continue to use Welch Allyn non-automated blood pressure gauges if the device is calibrated. To check that the device is calibrated, please perform the following quick check of calibration as described in the product Instructions for Use. At zero pressure, make certain the pointer is within the oval surrounding the zero-pressure gradation on the dial. If the pointer is fully outside of the tolerant zone (the darkly shaded area in the illustration) the device may need calibration. Although an unpressurized reading of zero does not guarantee accuracy at all scale points, failure of the pointer to indicate zero (±3 mm Hg) is an obvious sign of error. | 2025-03-03 20:10
Abbott Molecular Inc. | Alinity m SYSTEM | 11 | - | Please find attached FSN | Please find attached FSN | 2025-03-04 13:00
Occlutech International AB Landskronavägen 2 SE-252 32 Helsingborg  Sweden | Procedure Packs containing Flex-Pusher II (51FP100) | 2 | - | Occlutech has identified that certain Procedure Packs containing Flex Pusher II x D FP affected by this corrective action have a size recommendation of F for x D ODSIII and ODSv on the label of the Procedure Pack box and the sterile pouch of the x D implant that might lead to incompatibility x D This could happen because the interface between the Flex Pusher II delivery cable x D and the ODS III or v Delivery Set is determined by the interaction between the inside x D diameter of the Delivery Set and the outside diameter of the Delivery cable When both x D diameters are at their nominal values compatibility is assured However as tolerances x D are considered the dimensions of each device may vary within specified limits in case x D of Flex Pusher II FP at the upper end tolerance interval and ODS F ODSIII x D F DS US and ODSv F US at the lower end tolerance x D interval potential compatibility issues may occur | Actions to be taken by the physicians, hospitals and distributors: 1. Check your inventory and determine if you have any of the affected procedure packs: 29ASD06F, 29ASD07F, 29ASD09F, 29ASD10F, 71VSD06F, 89VSD08SF, 84VSD08F, 90VSD08LF. 2. In this case, please quarantine them and contact your direct Occlutech representative to return the products. 3. Please inform clinicians and users who may have the affected procedure packs if you have further distributed them. 4. A written confirmation of the receipt, the reading and understanding of the Field Safety Notice should be provided to info.mdso@occlutech.com, compiling the attached confirmation form. Please also refer to the contact information mentioned at the end of this letter. If you have any questions or would like assistance with this Field Safety Notice, please contact your direct Occlutech representative. | 2025-03-06 11:02
Occlutech International AB  Landskronavägen 2  SE-252 32 Helsingborg  Sweden | Procedure Packs containing Flex-Pusher II (51FP100) | 2 | - | Occlutech has identified that certain Procedure Packs containing Flex Pusher II x D FP affected by this corrective action have a size recommendation of F for x D ODSIII and ODSv on the label of the Procedure Pack box and the sterile pouch of the x D implant that might lead to incompatibility x D This could happen because the interface between the Flex Pusher II delivery cable x D and the ODS III or v Delivery Set is determined by the interaction between the inside x D diameter of the Delivery Set and the outside diameter of the Delivery cable When both x D diameters are at their nominal values compatibility is assured However as tolerances x D are considered the dimensions of each device may vary within specified limits in case x D of Flex Pusher II FP at the upper end tolerance interval and ODS F ODSIII x D F DS US and ODSv F US at the lower end tolerance x D interval potential compatibility issues may occur | Actions to be taken by the physicians, hospitals and distributors: 1. Check your inventory and determine if you have any of the affected procedure packs: 29ASD06F, 29ASD07F, 29ASD09F, 29ASD10F, 71VSD06F, 89VSD08SF, 84VSD08F, 90VSD08LF. 2. In this case, please quarantine them and contact your direct Occlutech representative to return the products. 3. Please inform clinicians and users who may have the affected procedure packs if you have further distributed them. 4. A written confirmation of the receipt, the reading and understanding of the Field Safety Notice should be provided to info.mdso@occlutech.com, compiling the attached confirmation form. Please also refer to the contact information mentioned at the end of this letter. If you have any questions or would like assistance with this Field Safety Notice, please contact your direct Occlutech representative. | 2025-03-06 11:06
Beckman Coulter Inc. | DxC 500i Clinical Analyzer | 3 | - | Beckman Coulter has determined that the DxC i Clinical Analyzer might assign a test result to an incorrect sample ID under the following workflow conditions x D After loading a barcoded sample tube Sample A onto the DxC i Clinical Analyzer all tests ordered for that barcoded sample tube are locked into the specific position of that barcoded tube on the analyzer for example the position on the DxLab Rack Carousel Rack or position in the STAT table until all tests are completed or cancelled for that sample tube x D A problem might occur if for any reason all tests ordered for that barcoded sample tube Sample A are not completed or remain in pending and the operator attempts to analyze a second sample tube Sample B in that same position on the rack or STAT table If the barcode for Sample B is not successfully scanned by the analyzer or if the sample tube is not labelled the analyzer assumes the sample that previously occupied that position has been reloaded Sample A The analyzer will complete and report the pending tests associated with Sample A using Sample B No results will be reported for Sample B x D For chemistry test analysis on the DxC i analyzer there is no real time sample event that alerts the operator when tests have not completed Only when the system returns to standby is the notification generated on the Sample Events page x D x D For immunoassay test analysis on the DxC i analyzer a sample event is generated immediately with details of the test that cannot be run x D e g Rack ID Position Insufficient reagent PSA Hyb loaded x D This issue does not arise if all barcoded sample tubes are successfully scanned or if testing is complete for the first barcoded sample Sample A |  Beckman Coulter recommends reviewing your laboratory workflow, sample labeling/label integrity and positioning to determine impact and sharing this with your laboratory and/or Medical Director.  Follow the attached Job Aid, which will prevent the operator workflow described in the ISSUE section. The Job Aid will outline how to use the sample list to manage the samples which have not completed.  Your Beckman Coulter service representative will contact you to schedule a software upgrade which will help in the management of this issue. This software upgrade will provide improved sample event information and real time reporting, identifying the specific location of the sample with tests that are not completed.  There is no requirement to reverify applications, for example, precision studies. | 2025-03-06 14:00
Beckman Coulter Inc. | DxC 500 AU Clinical Chemistry Analyzer | 1 | - | Beckman Coulter has determined that the DxC AU Clinical Chemistry Analyzer may not run a requested calibration order for a two part reagent and subsequently prevents any new calibration order from being placed x D For two part reagents when reagent blank or calibration is ordered during sample processing and then any of the components R and or R depletes to zero tests the analyzer will not be able to complete the calibration request and the calibration order will remain pending The analyzer will not allow additional calibration orders to be requested for any assays |  Do not place calibration orders when the analyzer is processing samples.  Refer to the DxC 500 AU IFU Reagent Overview section for instructions on Monitoring the Status of Reagents.  Beckman Coulter recommends sharing the content of this letter with your laboratory and/or Medical Director.  If you experience this issue on software version 1.3, 1.4, 1.4.1 or 1.4.2, contact Beckman Coulter for service. | 2025-03-06 14:20
MAQUET Cardiovascular, LLC | VASOVIEW HEMOPRO 2 Endoscopic Vessel Harvesting System | 1570 | - | Silicone peeling is defined as silicone that has peeled away from the Harvesting Tool Jaw but remains attached to the Jaw Silicone detachment is defined as silicone that has peeled away from the Harvesting Tool Jaw s and detached from the Jaw assembly Silicone peeling and detachment can occur when thermal loads on the device are not allowed to adequately cool down between actuations This can happen when the device is actuated for long periods of time or in quick succession such as in repeated tissue dissection or application of energy without tissue between the jaws | Review the IFU information provided in Appendix A of the letter, specifically note the warnings to minimize over-delivery of energy mentioned in the letter(page 5). Consider mitigating the risks mentioned in the customer letter(page 5). Getinge has created a training video, accessible via the QR code or link, to guide users on the proper use of the device and reinforce essential precautions. Use the QR code given in the letter to access the training.(page 6) Please forward this information to all current and prospective users of the VasoView Hemopro 2 Endoscopic Vessel Harvesting System within your hospital/facility. If you are a distributor who has shipped any affected products to customers, please forward this document to their attention for appropriate action. Please complete and sign the attached Response form to acknowledge that you have received this notification | 2025-03-07 01:16
Bard Access Systems, Inc. 1049 North 5600 West,Salt Lake City, Utah,84 116 United States | Power Picc Catheter | 878 | - | BD has conducted a comprehensive investigation and has identified that three lots of resin used to extrude the tubing material exceeded our supplier s specification for a material property called Melt Flow Index MFI BD s investigation has shown that higher MFI could lead to increased material fatigue leaks All of the extruded catheters meet BD specifications for dimensions and mechanical properties however BD is removing all product from the field that was extruded from resin that exceeded the supplier specification x D x D Although the increase in complaints is limited to specific geographies the product with higher MFI has been distributed worldwide therefore all impacted lots are being removed from the global market x D x D Figure Example of transverse circumferential crack in the catheter body x D BD has identified two additional contributing factors to the increase in material fatigue leaks seen on the Fr Single Lumen PowerPICCs in some geographies Each factor must be controlled in order to minimize the failures x D | Clinical User Actions BD is not recommending to explant product in-situ from any lots (including those being recalled). Consider the patient's infusion needs, alternative access options and the risks and benefits of continued catheter usage. The following additional advice is recommended Actions if catheter damage is not suspected: 1. Carefully examine the visible portion of the catheter to assess for any sign of damage to the catheter shaft. 2. Monitor the patient closely for signs and symptoms of catheter damage, such as increased extremity circumference, infusate leakage, or reports of pain. 3. If using a compression style securement device, consider replacing it with an adhesive-backed securement system to reduce the risk of material fatigue. Ensure securement systems are appropriately sized to accommodate the increase in diameter of the catheter in the taper region. 4. Ensure the catheter is positioned correctly, with the kink-resistant, reverse tapered region at the insertion site, to minimize the risk of leaks. Note that if the catheter is currently positioned with the kink-resistant, reverse tapered region external to the patient, it is not advisable to reposition it internally. Instead, be aware that not using this region may increase the risk of leaks. Use clinical judgment to decide whether to continue using the device or to replace it. Actions if catheter damage is suspected: 1. Immediately stop any infusion if catheter damage is suspected. 2. Follow your institution's guidelines for catheters with suspected damage. 3. If the catheter is confirmed to have a fracture or break, the catheter should be removed and an alternative route for access should be obtained. BD Actions • BD is removing all unexpired lots of 4Fr Single Lumen PowerPICC which exceed the supplier’s specification for MFI. • BD is notifying users that adhesive back securement systems are recommended. • BD is notifying users that if they do use compression style securement systems, they must be sized appropriately to accommodate the reverse taper. • BD is notifying customers to insert the catheter to as close as possible to the 0 mark, per the IFU. • BD has been updating the product IFU with cautions related to the use of compression style securement systems around the taper region of the PICC. • BD will provide product replacement or credit for all destroyed products. • BD will continue investigating this issue and will follow up with any additional actions. Customer Actions: • Cease use of any unused affected lot numbers of 4Fr Single Lumen PowerPICC. • Identify and quarantine all unused affected lot numbers of 4Fr Single Lumen PowerPICC. • Make a note of the lot numbers and immediately destroy all unused affected units. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 11th April 2025. • Circulate this notice to all those who need to be aware within your organization or to any organization where the potentially affected products have been transferred. • Please report all device-related incidents to the manufacturer, distributor or local representative, and the national Competent Authority if appropriate, as this provides important feedback. Distributor Actions: • Cease distribution of any unused affected lot numbers of 4Fr Single Lumen PowerPICC. • Identify, quarantine, making a note of the lot numbers then destroy all undistributed affected lot numbers of 4Fr Single Lumen PowerPICC. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 11th April 2025. • Complete and return the Customer Response Form following completion of your reconciliation activities. • Please report all device-related incidents to the manufacturer, distributor or local representative, and the national Competent Authority if appropriate, as this provides important feedback. | 2025-03-09 11:45
Beckman Coulter Inc. | DxC 500i Clinical Analyzer DxC 500i Access 2 Immunoasssay Module | 3 | - | Beckman Coulter has determined that during Installation of the DxC i Clinical Analyzer by the Service Engineer if the analyzer was configured with regional settings where the region uses commas as decimal identifiers information within the APF Assay Protocol File will load decimal numerals as whole numbers example will load as |  Your laboratory should cease running Access Toxo IgM II assay (REF34470) on the DxC 500i analyzer until an assessment is completed by your Field Service Engineer.  Beckman Coulter recommends sharing the content of this letter with your laboratory and/or Medical Director to determine if a review of previous patient test results should be conducted. | 2025-03-09 13:59
Beckman Coulter Inc. | DxC 500i Clinical Analyzer | 3 | - | Beckman Coulter has determined that the DxC i Clinical Analyzer may not run a requested calibration order for a two part reagent and subsequently prevents any new calibration order from being placed x D For two part reagents when reagent blank or calibration is ordered during sample processing and then any of the components R and or R depletes to zero tests the analyzer will not be able to complete the calibration request and the calibration order will remain pending The analyzer will not allow additional calibration orders to be requested for any assays |  Do not place calibration orders when the analyzer is processing samples.  Refer to the DxC 500i IFU Reagent Overview section for instructions on Monitoring the Status of Reagents.  Beckman Coulter recommends sharing the content of this letter with your laboratory and/or Medical Director.  If you experience this issue on software version 1.3.0 or 1.3.2, contact Beckman Coulter for service. | 2025-03-09 23:32
UMENDRA EXPORTS PVT LTD | Lubricant Skin Jelly | - | - | The medical device does not have Medical Devices Marketing Authorization MDMA which violates the medical device law Article | Withdraw the affected product from the market. | 2025-03-11 10:24
GE Medical Systems Ultrasound & Primary Care Diagnostics LLC | Venue | 54 | - | GE HealthCare has become aware that the battery for certain Venue Go and Venue Fit ultrasound systems with software versions R R and R see Affected Products List in this letter can potentially develop an internal failure which could result in smoke or fire x D x D There have been no injuries reported to GE HealthCare as a result of this potential issue x D | SW update + battery to all affected IB | 2025-03-12 14:42
GE Medical Systems Ultrasound & Primary Care Diagnostics LLC | Venue Go | 54 | - | GE HealthCare has become aware that the battery for certain Venue Go and Venue Fit ultrasound systems with software versions R R and R see Affected Products List in this letter can potentially develop an internal failure which could result in smoke or fire x D x D There have been no injuries reported to GE HealthCare as a result of this potential issue x D | SW update + battery to all affected IB | 2025-03-12 14:42
Medtronic XOMED, Inc | NIM VITAL™ CONSOLE | 60 | - | Medtronic engineers observed stimulus stim artifacts during a thyroidectomy using a NIM Vital system with software version Investigation determined that change impacting a function with a user configurable parameter implemented in v and kept in v may have resulted in an increased rate of stimulus artifact compared to prior software versions | • Identify affected products within your inventory. The product is not required to be returned for this issue as Medtronic has deployed NIM Vital™ Nerve Monitoring System software version 1.7.5. which is readily available to fix this issue. • Your Medtronic representative will contact you to install the new software version 1.7.5 for correction of the impacted product in your possession. • For patients who are currently being monitored with the NIM Vital Nerve Monitoring System (software version 1.6.4. or earlier), be aware of the possibility of increased stimulus artifact. Refer to the system instructions for use for instructions on how the stimulus artifact may be reduced or exacerbated through the adjustment of system settings including event threshold, stimulation current, and rejection period. • Please share this communication within your organization, with other organizations where impacted devices have been transferred, and any other associated organizations that may be impacted by this action. Maintain a copy of this letter for your records. • Please complete and return the customer acknowledgement form enclosed with this letter acknowledging receipt of this information even if you no longer have possession, custody or control over the affected product. Note: Instructions on returning the acknowledgement form to Medtronic is located on the form itself. • [Only required for customers who have this SW thumb drive in their control] Please discard any of the items below in your inventory. a. SOFTWARE NIM4SWU143 UPGRADE v1.4.3 b. SOFTWARE NIM4SWU154 UPGRADE v1.5.4 c. SOFTWARE NIM4SWU164 UPGRADE v1.6.4 | 2025-03-12 15:37
Radiometer Medical ApS | ABL90 Series | 86 | - | The device plays a video on screen to help the user perform the measurement An Intel Graphics driver that was used in the Windows operating system could make Windows crash during video playback | Software upgrade | 2025-03-13 12:43
Radiometer Medical ApS | ABL90 FLEX PLUS | 86 | - | The device plays a video on screen to help the user perform the measurement An Intel Graphics driver that was used in the Windows operating system could make Windows crash during video playback | Software upgrade | 2025-03-13 12:43
MTF Musculoskeletal Transplant Foundation | Posterior Tibialis Tendon | - | - | This voluntary market withdrawal is being initiated out of an abundance of caution due to a notification received by the CDC We do not believe this is a safety concern All donor suitability testing and final sterility results of each graft were negative There have been no reported adverse events or infections associated with any tissues from this donor | Voluntary market withdrawal of the following serial numbers from KSA market: 00423090891045 00423090891050 00423090891049 | 2025-03-13 23:46
MTF Musculoskeletal Transplant Foundation | Anterior Tibialis Tendon | - | - | This voluntary market withdrawal is being initiated out of an abundance of caution due to a notification received by the CDC We do not believe this is a safety concern All donor suitability testing and final sterility results of each graft were negative There have been no reported adverse events or infections associated with any tissues from this donor | Voluntary market withdrawal of the following serial numbers from KSA market: 00423090891045 00423090891050 00423090891049 | 2025-03-13 23:46
MTF Musculoskeletal Transplant Foundation | Posterior Tibialis Tendon | 3 | - | This voluntary market withdrawal is being initiated out of an abundance of caution due to a notification received by the CDC We do not believe this is a safety concern All donor suitability testing and final sterility results of each graft were negative There have been no reported adverse events or infections associated with any tissues from this donor | Voluntary withdrawal of the following serial numbers from KSA market: 00423090891045 00423090891050 00423090891049 | 2025-03-13 23:52
MTF Musculoskeletal Transplant Foundation | Anterior Tibialis Tendon | 3 | - | This voluntary market withdrawal is being initiated out of an abundance of caution due to a notification received by the CDC We do not believe this is a safety concern All donor suitability testing and final sterility results of each graft were negative There have been no reported adverse events or infections associated with any tissues from this donor | Voluntary withdrawal of the following serial numbers from KSA market: 00423090891045 00423090891050 00423090891049 | 2025-03-13 23:52
Beckman Coulter, Inc. | DxI 9000 Access Immunoassay Analyzer | 1 | - | This issue affects DxI analyzers that are run with the system software versions listed above and are connected to a host system laboratory information system LIS or middleware x D If a DxI analyzer has accumulated canceled QC test results that were not sent to the host system the analyzer may lose communication with the host system The lost connection cannot be restored by routine troubleshooting x D Sample processing is affected when the DxI analyzer loses communication with the host system | To Prevent the Issue: Each time that the Cleaning Routine maintenance procedure is due, Beckman Coulter recommends that your laboratory perform the steps provided in Appendix A before performing the Cleaning Routine. Docusign Envelope ID: CF77E07B-DC8C-492B-A04F-647911EF2031 Beckman Coulter, Inc. 1000 Lake Hazeltine Drive Chaska, MN 55318, USA Phone: 952.448.4848 fax: 800.232.3828 www.beckmancoulter.com FA-25027 Page 2 of 3 If the Issue Occurs:  If the DxI 9000 analyzer loses communication with the host system and does not generate an event message, perform the steps provided in Appendix A to restore communication with the host system.  If the DxI 9000 analyzer loses communication and generates an event message, follow the system event troubleshooting steps on the user interface to restore the LIS connection. If the troubleshooting steps do not resolve the event, perform the steps provided in Appendix A to restore communication with the host system.  Samples can be run on a different analyzer to prevent further delay to reporting patient results. | 2025-03-16 10:51
Baxter Medical Systems GmbH + Co. KG | TruSystem 7500 | 36 | - | Software issue which causes the upper back section of Model series tabletops to become inoperable non adjustable when the emergency mode function is enabled The emergency mode of the operating table is used to ensure that the tabletop functions as needed in case of a functional failure and can only be manually activated by the user on the column keypad This issue only occurs when emergency mode is enabled The tabletop and upper back section function normally when used outside emergency mode | To correct the issue, Baxter will install a software upgrade in the impacted TruSystem 7500 surgical table systems. | 2025-03-16 14:51
Siemens Healthcare Diagnostics Inc. | Atellica® CH RCRP | 228000 | - | Issue Description x D Siemens Healthineers has confirmed that incorrect software flagging may occur for the Atellica CH RCRP assay that may potentially lead to an erroneous result The probability of occurrence for an erroneous result in the absence of a flag is less than The probability of occurrence for an erroneous result with an error flag is or less This incorrect flagging is mitigated through the customer actions listed in this letter This issue can present with serum or plasma and with all Atellica CH RCRP reagent lots x D See Appendix A for additional information regarding the observed scenarios x D x D Kindly review the attached Urgent Field Safety Notice | Customer Actions • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • For both Atellica CH and Atellica CI analyzers, perform the instructions in Appendix B to temporarily reduce the measuring interval. o Until the measuring interval is restored, track additional reagent consumption as a result of these actions to report to Siemens Healthineers for future reimbursement/credit. • Additionally, for Atellica CH Analyzers, perform the instructions in Appendix C to remove rules for flagging of “No Calculation” results and to install Atellica Solution Software version 1.29.0 or higher. • Complete and return the Field Correction Effectiveness Check form attached to this letter within 30 days. • Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached Urgent Field Safety Notice | 2025-03-17 11:55
UMENDRA EXPORTS PVT.LTD | Lubricat Skin Jelly | 7200 | عدم حصول المستلزم الطبي على شهادة إذن بالتسويق (MDMA) | MDMA | The product has been recalled from customers. | 2025-03-17 11:57
Brainlab AG | Origin Data Management | 5 | - | patient records from unrelated individuals are automatically merged into one patient record without user notification | Following instructions | 2025-03-17 14:36
Rocket Medical PLC | Rocket Thoracentesis Catheter - 6Fg | 0 | - | During the production of these specific lots the adhesive glue used to secure the drainage line to the stitch plate was not correctly applied As a result the drainage line may separate from the tap | Upon the acknowledgement from the supplier, all the quantity was held to make sure that Saudi market is not affected by this Safety Alert | 2025-03-18 12:06
Rocket Medical PLC | Rocket Thoracentesis Catheter | 0 | - | During the production of these specific lots the adhesive glue used to secure the drainage line to the stitch plate was not correctly applied As a result the drainage line may separate from the tap | Upon the acknowledgement from the supplier, all the quantity was held to make sure that Saudi market is not affected by this Safety Alert | 2025-03-18 12:06
Philips | Philips lumea bri958 9000 series ipl hair removal tool | - | - | Reporting violating medical devices x D x D Recently we have noticed that Beauty sook is selling online in Saudi non registered non medical IPL Lumea products with a claim official Philips retailer PHILIPS LUMEA BRI SERIES IPL HAIR REMOVAL TOOL WITH ATTACHME BeautySook x D x D We are selling in Saudi only MDMA approved Saudi compliant IPL models hence we want to report SFDA about these non complying products Please share with us the steps for reporting to SFDA | Recently we have noticed that “Beauty sook” is selling online in Saudi non registered non medical IPL Lumea products with a claim “official Philips retailer” PHILIPS LUMEA BRI958 9000 SERIES IPL HAIR REMOVAL TOOL WITH 4 ATTACHME – BeautySook We are selling in Saudi only MDMA approved Saudi compliant IPL models hence we want to report SFDA about these non complying products. Please share with us the steps for reporting to SFDA. | 2025-03-19 13:35
Materialise | OrthoView | 46 | - | x D A security issue was discovered in OrthoView and earlier versions where unauthenticated users may be able to under certain conditions execute arbitrary commands on the OrthoView server This issue has been rated high with a CVSS score of The vulnerability will be published as CVE once affected customers are informed x D x D We currently have no indications that this vulnerability is actively exploited or has led to any impact on either patient health or patient data However prompt action is required to ensure your OrthoView setup is adequately secured x D x D The likelihood of causing injury to a patient is negligible There is a high chance should the vulnerability be exploited that a surgical procedure will be delayed or postponed due to the absence of a suitable plan x D x D How to recognize that the device may fail the user may not be able to log into the OrthoView application their preferences may have changed since last log in or the PACS connection may be misconfigured unavailable x D | Complete “Step 1: Confirm you have a vulnerable configuration”, as detailed in Appendix A If confirmed, complete “Step2: Resolve vulnerability by following these resolution instructions” described in Appendix A. The configuration of OrthoView needs to be changed to permanently remove this vulnerability. This can be done by disabling HTTP Servlet sharing and switching to network folder sharing. There is no loss in functionality or performance by performing these actions. These actions are confirmed as the long-term solution. The OrthoView team is taking this vulnerability and your security very seriously. To ensure that the identified vulnerability is adequately mitigated, we have contracted an external security penetration testing team, and they confirm that the above solution resolves the vulnerability. As part of our continual improvement, we will review internal development processes to additionally strengthen OrthoView security. Should you have any additional questions or concerns, please reach out to OrthoView support via orthoviewsupport@materialise.co.uk. Due to the nature of this notification, we request formal acknowledgement of this letter and confirmation of your intent to take these actions to rectify this issue. Please complete and return the attached Customer Response form by (Month Date, 2025). Return completed forms by mail or email or by telephone to: • Mail to: Meridian Technique Ltd., Epsilon House, Enterprise Road, Southampton Science Park, Southampton, Hampshire, SO16 7NS, Great Britain • Email: orthoviewsupport@materialise.co.uk • Telephone: +44 (0)2380 762 500 | 2025-03-20 12:35
GE Healthcare Finland Oy, , located at Kuortaneenkatu 2, FI-00510 Helsinki, Finland. | Centricity High Acuity Critical Care 5.X | 5 | - | GE HealthCare has become aware of a potential issue within the CHA Order module s Edit Order window Users may unintentionally reschedule the next administration task to a much later time by inadvertently entering the order End time into the Next fields This error could lead to delays in medication administration x D x D There have been no injuries reported related to this issue x D x D | The FMI correction will be available in two sw patches: CHA 5.8.2 patch C for 5.8.2 installations and CHA 5.8.4 patch A for 5.8.4. CHA 5.8.3 installations shall be upgraded to 5.8.4A as part for field action. | 2025-03-20 12:49
GE Healthcare Finland Oy, , located at Kuortaneenkatu 2, FI-00510 Helsinki, Finland. | Centricity High Acuity Anesthesia 5.X | 5 | - | GE HealthCare has become aware of a potential issue within the CHA Order module s Edit Order window Users may unintentionally reschedule the next administration task to a much later time by inadvertently entering the order End time into the Next fields This error could lead to delays in medication administration x D x D There have been no injuries reported related to this issue x D x D | The FMI correction will be available in two sw patches: CHA 5.8.2 patch C for 5.8.2 installations and CHA 5.8.4 patch A for 5.8.4. CHA 5.8.3 installations shall be upgraded to 5.8.4A as part for field action. | 2025-03-20 12:49
Abbott Laboratories Diagnostics Division | CELL-DYN Ruby | 182 | - | Please find attached FSN letter | Please find attached FSN letter | 2025-03-20 13:52
Hillrom | Progressa Surface | - | - | Not Specified | On site correction of the affected items | 2025-03-20 15:40
Datex-Ohmeda, Inc. 3030 Ohmeda Drive, PO Box 7550, Madison, 53707-7550, USA | Carestation 620/650/650c and 750/750c | 98 | - | GE HealthCare has become aware that certain Carestation c and c Anesthesia Delivery Systems see Affected Products List in this letter will not provide effective ventilation in Volume Control Ventilation VCV mode In these systems effective ventilation can be achieved in Pressure Control Ventilation PCV or Pressure Control Ventilation Volume Guarantee PCV VG modes or with Manual ventilation x D x D If this issue occurs it will be obvious to the user via observation and multiple alarms The inflated bellows visible through transparent glass will stop moving and an audible alarm and visual Unable to Drive Bellows message will alert the user Additional alarms including Apnea EtCO low MVexp low RR low and x D TVexp low will also alert the user to inadequate ventilation x D x D In the unlikely event this issue occurs and it is not noticed it can result in hypoxia x D x D There have been no injuries reported as a result of this issue x D | Send customer letter with CRF(Customer Response Form) and provide customers with screening test to be performed to determine whether or not device is impacted. CRF will document the failed device serial numbers Send FE to rework all failed system in the field based on the CRF | 2025-03-24 11:49
Brainlab AG | Origin Data Management | 1 | - | Under specific circumstances patient records from unrelated individuals x D are automatically merged into one patient record | Advice regarding use of device | 2025-03-24 13:36
SYMATESE | Nevelia® | 8 | - | Manufacturing issue in bonding the two PET layers causing possible separation | destroy the affected items | 2025-03-24 14:13
Beckman Coulter Inc. | DxC 500i Clinical Analyzer | 3 | - | Beckman Coulter has determined that the DxC i Clinical Analyzer may unexpectedly transition to a stopped state as a result of specific loading workflows potentially causing a delay in reporting of results |  After selecting the “Load Carousel Rack” button on the analyzer, select the first “Load Carousel Rack” pop-up message and select “Done.”  If additional messages have been created, ignore all messages while the immunoassay tests complete processing from the Carousel Rack and the immunoassay analyzer goes to the “Standby” state, to avoid the workflow outlined in the Impact section.  After the immunoassay analyzer has been confirmed to be in the “Standby” state, cancel all additional “Load Carousel Rack” messages.  The clinical chemistry analyzer will continue to process samples.  Select the “Load Carousel Rack” button again, tapping only once to generate one message, then load the carousel rack.  If the workflow outlined in the Impact section occurs, where the analyzer proceeds to the stopped state, the operator must restart the DxC 500i Clinical Analyzer. Please refer to the IFU for restart instructions.  As a result of the system restart, pending orders for loaded racks will not be processed and results for samples in progress will be canceled. These samples must be reordered and analyzed again.  Beckman Coulter recommends sharing the content of this letter with your laboratory and/or Medical Director. | 2025-03-24 22:51
Radiometer Medical ApS | ABL800 BASIC with AutoCheck | 571 | - | The study concludes that there is a remote probability of reporting biased out of x D specification pH results on blood samples x D This may occur if the calibration solution s pH value decreases during the in use x D period due to bacterial growth in the calibration solution bottles CAL and or CAL | software V6.20 MR2 implements an automatic surveillance of pH results for 2-point calibrations. The software implementation basically automates Procedure 2 | 2025-03-25 12:17
WERFEN | HemosIL AcuStar ADAMTS13 Activity | 180 | - | IFU Results section x D A test result of will automatically be flagged by the instrument with the following notification Assay x D intended for diagnosis and monitoring of TTP Check sample integrity and verify results against other clinical and x D laboratory findings prior to reporting x D IFU Limitations interfering substances section x D HemosIL AcuStar ADAMTS Activity is intended to aid in the diagnosis and monitoring of TTP for adult patients x D only Internal testing has not been performed to establish performance claims in the following settings x D Pediatric population years of age Current performance claims were established with adult samples x D To guide patient management plans Once diagnosed TTP patient management and therapy strategies x D should be based on current guidelines and recommendations x D Predict TTP relapse and reoccurrence x D Important Per current guideline TTP diagnosis is not solely based on ADAMTS activity results ADAMTS x D activity results should be used in conjunction with other clinical and laboratory findings | Requested Customer Actions Please take the following actions: 􀁸 Review the HemosIL AcuStar ADAMTS13 Activity labeled intended use. 􀁸 Evaluate patient results based on the above clarifications and best practices. 􀁸 Assess whether additional testing is necessary before taking clinical action for results that are inconsistent with other clinical or laboratory findings. 􀁸 Share this information with your laboratory staff. 􀁸 Forward this closure notification to all affected locations within your facility. 􀁸 Retain a copy of this closure notification for your records. | 2025-03-26 12:53
ENDOGASTRIC SOLUTIONS, inc. / Merit Medical | ESOPHYX Z+ | - | - | Inadequate warnings in the original IFU regarding over rotation and multiple deployments of fasteners at the same site | Update IFU with additional warnings Discard old IFU and post the updated version near products Notify internal staff and downstream users Submit signed Customer Response | 2025-03-27 11:11
VIOL CO., Ltd | Bi-Polar Electrodes used with Scarlet S Device | - | - | Counterfeit medical products | • Withdraw all affected products. • Submit A Field Safety Corrective Action through the NCMDR website: https://ade.sfda.gov.sa/Home/NcmdrReport | 2025-03-27 12:52
Advanced Sterilization Products Division of Ethicon,Inc. A Johnson & Johnson Company | STERRAD® 100NX® Sterilizer Validation Kit, International | 0 | - | ASP records indicate your facility has a STERRADTM NX Sterilizer and or a STERRAD NXTM Sterilizer that x D may be impacted by this issue Users can still operate the sterilization systems in accordance with their x D respective user s guide pending the correction | Maintain a copy of this Customer Notification at your facility and ensure it remains available to Healthcare Workers. • Ensure Healthcare Workers who operate the sterilizers have read and understood this Customer Notification. • Complete and return the Customer Reply Form within five (5) business days of receipt of this notification. • If your facility observes this issue, contact ASP at 888-STERRAD (888-783-7723) with the STERRAD Sterilization System serial number. | 2025-03-28 00:17
ZOLL Medical Corporation | Powerheart G5 | 200 | - | The issue is likely caused by exposure to humidity outside the device s specified limits which results in failed automated self tests and generation of EC Affected AEDs become non functional and display this error preventing therapy delivery | Routine for All Customers: 1. Regularly monitor AED self-test results. 2. Immediately respond to any self-test failures. For Customers with At-Risk Devices: 1. Share this notification with all users. 2. Ensure programs are in place to routinely monitor device status. 3. Store devices within specified environmental limits: • Temperature: 0°C to 50°C (32°F to 122°F) • Humidity: 10% to 95% (non-condensing) 4. Acknowledge receipt of the notice using the customer notification form. | 2025-03-29 15:21
Covidien llc | Hugo™ RAS Arm Cart Assembly | 37 | - | An increase in complaints related to intermittent connectivity of instruments with the Hugo RAS System was observed This failure mode is defined as Instrument ID failure and occurs when the Hugo RAS System does not recognize an instrument being connected to the system The failure mode can be observed during setup of the Hugo RAS system and when in use intraoperatively Connectivity issues can result from contact interface with the gold plated pins of the Sterile Interface Module SIM | • Please immediately quarantine and discontinue the use of the affected model number with associated serial numbers listed in Attachment #1: List of Affected Serial Numbers. • Please return the affected product. All products from the affected model number and associated serial numbers must be returned. • Notify all personnel in all care environments in which the Hugo™ RAS system is used about this medical device recall. • If you experience this issue, replace the SIM and report any incidents related to this issue to your local Medtronic representative. • Complete the attached Customer Acknowledgement Form. • Please maintain a copy of this notice in your records. | 2025-03-31 08:53
Philips | Big Bore RT | 4 | - | Philips has become aware of two unintended motion issues related to Interventional Control that may lead to a collision of the Gantry or table with the operator or patient x D Interventional Control Touch Panel Motion Button not Released x D When releasing the motion button in touch panel of Interventional control IVC by sliding or dragging off the touch panel area the motion button of vertical table motion and or rocker tilt angle may not be released x D This issue could result in additional table up down and or Gantry tilt motion x D Plan Box not Updated during Interventional Procedure x D The Plan Box may not move to the desired Needle Position or Work Position under the following circumstances x D Operator manually edits the Needle Position or Work Position value and x D Operator moves Plan Box by dragging it or by changing its center value and plan box position and x D Operator uses Interventional Controls Move to Selected Scan Work Position buttons x D This issue could result in the table moving to a location the operator does not expect x D Philips has identified other issues affecting Philips Big Bore CT system that do not present a risk to health Detailed descriptions and advice to customers pertaining to these issues are provided in Appendix A x D Philips has not received any reports of an adverse event associated with these issues as of March | Philips will contact you to schedule a time for a Philips Field Service Engineer (FSE) to visit your site and install the solution to resolve the issue (reference FCO72800825). | 2025-04-03 11:42
Philips Medical Systems (Cleveland), Inc. | Brilliance CT Big Bore | 4 | - | Philips has become aware of two unintended motion issues related to Interventional Control that may lead to a collision of the Gantry or table with the operator or patient x D Interventional Control Touch Panel Motion Button not Released x D When releasing the motion button in touch panel of Interventional control IVC by sliding or dragging off the touch panel area the motion button of vertical table motion and or rocker tilt angle may not be released x D This issue could result in additional table up down and or Gantry tilt motion x D Plan Box not Updated during Interventional Procedure x D The Plan Box may not move to the desired Needle Position or Work Position under the following circumstances x D Operator manually edits the Needle Position or Work Position value and x D Operator moves Plan Box by dragging it or by changing its center value and plan box position and x D Operator uses Interventional Controls Move to Selected Scan Work Position buttons x D This issue could result in the table moving to a location the operator does not expect x D Philips has identified other issues affecting Philips Big Bore CT system that do not present a risk to health Detailed descriptions and advice to customers pertaining to these issues are provided in Appendix A x D Philips has not received any reports of an adverse event associated with these issues as of March | Philips will contact you to schedule a time for a Philips Field Service Engineer (FSE) to visit your site and install the solution to resolve the issue (reference FCO72800825). | 2025-04-03 11:42
Philips | Brilliance CT Big Bore Oncology | 4 | - | Philips has become aware of two unintended motion issues related to Interventional Control that may lead to a collision of the Gantry or table with the operator or patient x D Interventional Control Touch Panel Motion Button not Released x D When releasing the motion button in touch panel of Interventional control IVC by sliding or dragging off the touch panel area the motion button of vertical table motion and or rocker tilt angle may not be released x D This issue could result in additional table up down and or Gantry tilt motion x D Plan Box not Updated during Interventional Procedure x D The Plan Box may not move to the desired Needle Position or Work Position under the following circumstances x D Operator manually edits the Needle Position or Work Position value and x D Operator moves Plan Box by dragging it or by changing its center value and plan box position and x D Operator uses Interventional Controls Move to Selected Scan Work Position buttons x D This issue could result in the table moving to a location the operator does not expect x D Philips has identified other issues affecting Philips Big Bore CT system that do not present a risk to health Detailed descriptions and advice to customers pertaining to these issues are provided in Appendix A x D Philips has not received any reports of an adverse event associated with these issues as of March | Philips will contact you to schedule a time for a Philips Field Service Engineer (FSE) to visit your site and install the solution to resolve the issue (reference FCO72800825). | 2025-04-03 11:42
Beckman Coulter Inc. | REMISOL Advance | 34 | - | Beckman Coulter has identified an issue with REMISOL Advance when receiving a sample order from the Laboratory Information System LIS if the patient s Date of Birth DoB is not entered or is unknown x D If the LIS sends as the DoB REMISOL Advance using the ASTM host driver incorrectly interprets the age as days x D If the LIS sends an empty DoB REMISOL Advance using the HL host driver incorrectly interprets the age as days x D Please note that you will not experience any issues if DoB field is empty when your system uses the ASTM host driver or if the default date is displayed when your system uses the HL host driver | • Verify that your system is configured with the default validation rules and ranges. REMISOL Advance will utilize these default validation rules in situations where the patient's age or date of birth is unknown (please see the chapter “Reference and Validation Ranges” 1-233 in the IFU) OR • Ensure that the DoB field is not left empty and confirm with your LIS vendor that they do not use 01/01/1900 as a default date of birth. If they do, request that this default date be changed to 02-Jan-1900 (02/01/1900) to allow REMISOL Advance to correctly estimate the patient's age. • The IFU will be revised to specify that the DoB is required for generating and validating results. • Beckman Coulter will provide software updates upon customer request. For HL7 Host Driver Users: • Beckman Coulter has implemented the correction in software versions ≥ 2.3.09. For ASTM Host Driver Users: • Beckman Coulter will implement the correction in software version 2.4. | 2025-04-05 15:51
Beckman Coulter Inc. | REMISOL Advance (using HL7 host driver) | 1 | - | As a result of internal investigation Beckman Coulter has identified an issue concerning the sample order priority update with the product REMISOL Advance when it uses the HL host driver x D When a sample order is initially set to Routine and later updated to STAT by the Laboratory Information System LIS REMISOL Advance fails to reflect this change retaining the sample s priority as Routine x D All REMISOL Advance versions using the HL host driver are affected x D If you are not using HL host driver you are not affected | After updating the sample order status from "Routine" to "STAT" in the LIS, please ensure that you also manually update the status in REMISOL Advance (see section Working with Current Requests / Samples List Popup Menu 1-129 in the IFU). | 2025-04-05 16:01
Biomerieux | BIOFIRE® Blood Culture Identification 2 (BCID2) Panel | 31 | - | The purpose of this letter is to inform you that bioM rieux is releasing a mandatory Pouch Module software update for the BIOFIRE Blood Culture Identification BCID Panel part number RFIT ASY that will mitigate risk to previous field safety notices | - Update your BIOFIRE BCID2 Panel Pouch Module software | 2025-04-06 09:26
Fresenius Medical Care AG | 4008S | 22 | - | Fresenius Medical Care discovered during quality controls a production issue leading to a potentially leaking disinfection valve V in the S V hemodialysis machine A kinked hose on our manual disinfection valve workstation was causing a leak test for the valve to give false positive test results As a result potentially leaking valves were not correctly identified and replaced in production These were processed further and ended up in production for certain S V devices and as spare parts x D The faulty valves could lead to hazardous situations due to insufficient disinfection While it can be excluded that disinfectant agent are infused into the patient as the disinfection valve is positioned downstream of the dialyzer the backflow of fluid from the device through a leaky valve leads to dilution of disinfectant in the container This could lead to a hazardous situation to the patient due to insufficient disinfection of the device x D Fresenius Medical Care has not received any complaints or incidents on this topic x D x D Potentially faulty solenoid valves can be identified by a technician as they do not have a serial number printed on the label and were produced from calendar week to calendar week | Potentially affected devices and those that received the above-mentioned spare parts need to be checked and affected solenoid valves need to be replaced. Until a Fresenius Medical Care technician has performed the checks, Fresenius Medical Care recommends that healthcare providers do the following: The affected 4008 S V10 devices (see Appendix 1) and those that received the above-mentioned spare parts must not be used for treatments until the solenoid valve V84 has been replaced. Disinfectant canister used for those affected devices need to be discarded and replaced as they might be diluted due to the leaking valve. | 2025-04-06 14:15
Q’Apel Medical, Inc. | Hippo 072 Aspiration System | 3 | - | Q Apel recently received a Warning Letter from the US Food and Drug Administration FDA in which the agency raised concerns about certain of our internal processes as well as the scope of our clearance for the Hippo product as it relates to the distal tip Specifically FDA has raised questions about the features and characteristics of the distal tip of the aspiration catheter when removing a clot during aspiration x D We have therefore developed a comprehensive action plan to address the agency s concerns and are confident that our planned actions many of which have already been or are in the process of being implemented will fully address the FDA s observations x D Following ongoing discussions with the FDA it has become clear that further efforts to satisfy the agency s expectations do not align with our strategic priorities Consequently we have made the difficult decision to voluntarily discontinue distributing the Hippo Aspiration System effective immediately and have made a business decision to voluntarily remove any remaining inventory from your inventory | We ask that you quarantine any remaining Hippo 072 Aspiration Systems in your possession and return them to us. If you have provided this product to others, please let us know so that we can notify them of this action and retrieve any remaining product. We have included a customer Acknowledgement and Response Card, and we ask that you complete and return that card so that we can account for all product and ensure that Q’Apel has accounted for all device in this voluntary action. Please send completed cards to orders@qapelmedical.com. Once we receive your completed Acknowledgement and Response Card, and if you have units to return, we will provide a Returned Material Authorization (RMA) number and shipping instructions. You may request to have your Q’Apel sales representative pick up the devices and return them on your behalf. In both cases, we require your completed Acknowledgement and Response Card in order to issue the RMA. Please do not discard or destroy any Hippo 072 Aspiration Systems in your possession as our goal is to recover and account for all units; we will require receipt of the returned goods in order to issue a timely credit for purchased devices. | 2025-04-07 10:46
Abbott Molecular Division Inc. | Alinity m SYSTEM | 11 | - | Please find the attached FSN letter | Please find the attached FSN letter | 2025-04-08 11:16
Getinge | GETINGE AQUADIS 56 | 17 | - | It has come to our attention that there is a possibility that the device can have an unprogrammed Frequency Inverter x D The Frequency Inverter is facilitated in the electrical cabinet which is a part of the washer disinfector x D The Frequency inverter controls the rotational speed of the circulation pump When the Frequency inverter is not programmed this causes the circulation pump to start earlier than expected creating possibility for a short dry run before lubrication of axial seal is maintained by the waterflow in the pump housing As a consequence if the frequency inverter is not programmed this can potentially cause additional wear and tear on the axial seal in the pump housing x D The device is still safe to use and you do not have to undertake any preventive actions The machine is working as intended with the exception that in the long run this can cause wear and tear earlier than expected and the water consumption might be higher than intended | 1.Please make sure that all caregivers and users of the Getinge Washer Disinfector Aquadis 56 are made aware of this Field Corrective Notice. 2.Complete and sign the enclosed Customer Response Form and return this form to the local Getinge office. 3.A Getinge Sales or Service person will plan this service within next planned maintenance. Note: If no service agreement with Getinge, a Sales or Service person will contact the person you listed on the Customer Response Form to investigate if your device is affected by this failure and if so, replace the memory module in the Frequency Inverter, free of charge. 4.Make sure that all listed devices at your facility are available for the Field corrective action during the Getinge service technician visit. | 2025-04-09 12:36
Philips Medical Systems | FilterLine set | 62 | - | The supplier of the Microstream Advance Intubated CO Filter Line VitaLine Intubated CO Filter Line and FilterLine Intubated CO Filter Line has provided an update to the Instructions for Use IFU for affected product The IFU addendum is being issued to address customer reports of difficulty or inability to disconnect an adapter from a patient s endotracheal tube in order to perform a procedure The addendum includes information on potential adverse events that may result from a failure to follow instructions for safe use of the airway adapter as originally stated in the IFU Supplier did not identify any anomaly or non conformance with the product | • Following addendum to the IFU will be issued: Adverse events associated with failure to follow the Instructions for Use while attaching and detaching the airway adapter from the breathing circuit are listed in descending order of severity: unintended extubation, respiratory failure, hypoxia, low oxygen saturation, aspiration/inhalation and delay to treatment. Any serious incident related to device use that may occur should be reported immediately to the manufacturer, the local competent authority, and any other regulators as required. | 2025-04-09 16:41
Philips Medical Systems | FilterLine H set | 62 | - | The supplier of the Microstream Advance Intubated CO Filter Line VitaLine Intubated CO Filter Line and FilterLine Intubated CO Filter Line has provided an update to the Instructions for Use IFU for affected product The IFU addendum is being issued to address customer reports of difficulty or inability to disconnect an adapter from a patient s endotracheal tube in order to perform a procedure The addendum includes information on potential adverse events that may result from a failure to follow instructions for safe use of the airway adapter as originally stated in the IFU Supplier did not identify any anomaly or non conformance with the product | • Following addendum to the IFU will be issued: Adverse events associated with failure to follow the Instructions for Use while attaching and detaching the airway adapter from the breathing circuit are listed in descending order of severity: unintended extubation, respiratory failure, hypoxia, low oxygen saturation, aspiration/inhalation and delay to treatment. Any serious incident related to device use that may occur should be reported immediately to the manufacturer, the local competent authority, and any other regulators as required. | 2025-04-09 16:41
Philips Medical Systems | VitaLine H Set | 62 | - | The supplier of the Microstream Advance Intubated CO Filter Line VitaLine Intubated CO Filter Line and FilterLine Intubated CO Filter Line has provided an update to the Instructions for Use IFU for affected product The IFU addendum is being issued to address customer reports of difficulty or inability to disconnect an adapter from a patient s endotracheal tube in order to perform a procedure The addendum includes information on potential adverse events that may result from a failure to follow instructions for safe use of the airway adapter as originally stated in the IFU Supplier did not identify any anomaly or non conformance with the product | • Following addendum to the IFU will be issued: Adverse events associated with failure to follow the Instructions for Use while attaching and detaching the airway adapter from the breathing circuit are listed in descending order of severity: unintended extubation, respiratory failure, hypoxia, low oxygen saturation, aspiration/inhalation and delay to treatment. Any serious incident related to device use that may occur should be reported immediately to the manufacturer, the local competent authority, and any other regulators as required. | 2025-04-09 16:41
Philips Medical Systems Nederland B.V. | IntelliSpace Cardiovascular | 20 | - | Philips has identified a software issue affecting IntelliSpace Cardiovascular ISCV software versions x x and x When using the Echo Module of ISCV the issue may result in missing or incomplete information when Finding Codes are used to add information under the following circumstances x D a When the same Finding Code is added into multiple Finding Groups in the same section all of the expected Finding Codes do not appear in the report or x D b If the Finding Code is sent to the conclusion section of the report by double clicking the Finding Code the expected Finding Code does not appear in the conclusion of the report x D x D Please refer to Attachment A for a detailed description of the two workflows associated with the issue x D The issue was identified during internal testing there have been no customer complaints or reported adverse events associated with this issue | Once all customer responses have been received, a Philips representative will contact customers who use or may use the affected workflows to schedule a time to install the software solution on your system(s) to resolve the issue (reference FCO-TBD). | 2025-04-09 16:56
Alshifa Medical Syringe Mfg Co. | RAMY Hypodermic Syringe | 44000 | - | The scale printed on the syringe is not accurate for the smaller increments which may lead to incorrect dosing especially when precise measurements are required | To replace all the quantities delivered to the costumers. | 2025-04-10 11:28
GE Healthcare Finland Oy, , located at Kuortaneenkatu 2, FI-00510 Helsinki, Finland. | Portrait HSWXB | 100 | - | GE HealthCare has become aware of a software issue that prevents SpO probe x D off and Check Resp patch alarms initiated before or during reboot from being x D displayed after the reboot on certain Portrait Mobile Patient Monitor Hub x D systems These alarms alert clinicians of the loss of pulse oximetry or respiration x D rate monitoring A Hub reboot occurs in response to specific software or hardware x D issues It lasts a maximum of seconds and occurs rarely estimated one in ten x D thousand hours of use x D x D There have been no complaints or injuries reported as a result of this issue x D | A software update that will be installed by a GE HealthCare Field Engineer. The Field Engineer will remove and destroy any copies of the old software (v1.1.4). | 2025-04-10 12:00
Randox Laboratories Ltd | Human Assayed Mutli Sera Level 3 | 771 | - | Randox Laboratories can confirm the analyte NEFA in the above listed lots and packed into x D the above listed batches has increased over the shelf life of the product and requires the x D targets and ranges to be reassigned Please see below table for updated targets and ranges x D The value sheets have been updated and are available on www randox com please discard x D the previous version and download the updated version | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Please accept our apologies for any inconvenience caused. Thank you for your patience and understanding. If you have any questions or concerns please contact Randox Technical Services. | 2025-04-10 14:28
Randox Laboratories Ltd | Human Assayed Mutli Sera Level 3 | 771 | - | Randox Laboratories can confirm the analyte NEFA in the above listed lots and packed into x D the above listed batches has increased over the shelf life of the product and requires the x D targets and ranges to be reassigned Please see below table for updated targets and ranges x D The value sheets have been updated and are available on www randox com please discard x D the previous version and download the updated version | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Please accept our apologies for any inconvenience caused. Thank you for your patience and understanding. If you have any questions or concerns please contact Randox Technical Services. | 2025-04-10 14:28
Medtronic, Inc. | Ascenda™ Intrathecal Catheter | 468 | - | Medtronic has developed a revised design and manufacturing process for the Ascenda catheter The intent of the design changes is to reduce the risk of a specific possible pathway for tissue in growth inside the catheter interface with the catheter connector x D x D Design Changes x D The Ascenda design change focuses on improvement of the catheter connection seal quality functionality to reduce the probability of tissue growth into the Leak Path that could potentially cause an occlusion of the drug path within the connector The design updates to the Ascenda implement the same design of tube inner diameter ID and cone angle as that of the SC catheter to improve seal quality x D | Advice on actions to be taken by the distributor and the user: • Complete the enclosed Customer Acknowledgement Form acknowledging that you have received this information. • Share this notice with all those who need to be aware of this design update within or outside your organization or to any organization where the potentially affected product has been transferred or distributed and maintain a copy of this notice in your records. Update March 2025 Please review all inventory of Ascenda Catheter(s) and take the following actions: • Identify the Ascenda catheter product manufactured on or before 09-MAY-2024 that does not have the updated design. The FSN provides an example of the symbols used to identify the Manufacturing date on the outer packaging of the Model 8780, 8781, and 8784 Ascenda kit. • Return unused, affected product in your inventory to Medtronic. Your Medtronic representative can assist you in the return of affected product as necessary. • Complete the enclosed Customer Acknowledgement Form This form must be returned even if you do not have any affected product in your possession. • This notice should be passed on to all those who need to be aware within your organization or to any organization where the potentially affected devices have been transferred. Please maintain a copy of this notice in your records. | 2025-04-11 16:26
Glooko, Inc. | Glooko Web Application | - | - | Glooko imports data from diabetes devices and displays the data according to the device s settings If the time settings on a diabetes device are different from the actual local time for example due to incorrect time settings on the device or while traveling between time zones the data will be displayed using the device s time setting This means that your reference data displayed on Glooko may appear with a time discrepancy x D Users may not realize that there is a time difference when uploading their diabetes data to Glooko In most cases there are no negative health consequences but there is a slight possibility that insulin dosages may need adjustment | Actions Users Should Take To avoid this issue, users are advised to ensure that the time settings on their personal diabetes device match the local time, so the data is displayed correctly in the Glooko Web Application. Actions Taken by the Manufacturer Glooko is unable to modify a patient’s diabetes data or set the correct time without violating data integrity standards. Therefore, to prevent future time discrepancies, Glooko is updating its current Instructions for Use and adding a user warning about the issue and how to correct it. You can view the updated Instructions for Use here, and additional information and FAQs are available here. Contact If you have any additional questions regarding this matter, please contact support at https://support.glooko.com or via email at help@glooko.com. | 2025-04-13 10:12
Glooko, Inc. | Glooko Web Application | - | - | Glooko imports data from diabetes devices and displays the data according to the device s settings If the time settings on a diabetes device are different from the actual local time for example due to incorrect time settings on the device or while traveling between time zones the data will be displayed using the device s time setting This means that your reference data displayed on Glooko may appear with a time discrepancy x D Users may not realize that there is a time difference when uploading their diabetes data to Glooko In most cases there are no negative health consequences but there is a slight possibility that insulin dosages may need adjustment | Actions Users Should Take To avoid this issue, users are advised to ensure that the time settings on their personal diabetes device match the local time, so the data is displayed correctly in the Glooko Web Application. Actions Taken by the Manufacturer Glooko is unable to modify a patient’s diabetes data or set the correct time without violating data integrity standards. Therefore, to prevent future time discrepancies, Glooko is updating its current Instructions for Use and adding a user warning about the issue and how to correct it. You can view the updated Instructions for Use here, and additional information and FAQs are available here. Contact If you have any additional questions regarding this matter, please contact support at https://support.glooko.com or via email at help@glooko.com. | 2025-04-13 10:12
Randox Laboratories Ltd | Liquid Assayed Specific Protein Control Level 1 | 1782 | - | Randox Laboratories can confirm that the Rheumatoid Factor RF concentrations in Liquid x D Assayed Specific Protein Controls Level Lots LPC and LPC are low below the x D measurement range of nephelometric methods on Siemens instruments Randox have since x D updated the IFU s and removed the associated targets and ranges x D Randox Laboratories can also confirm that Anti Streptolysin O ASO target and range values x D have been reassigned in lot no s LPC LPC LPC LPC and LPC Please see x D table below for updated target values | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Download corrected IFU from Smartdocs.com and discard of previous copies. Please accept our apologies for any inconvenience caused. Thank you for your patience and understanding. If you have any questions or concerns, please contact Randox Technical Services. | 2025-04-13 15:20
Randox Laboratories Ltd | Liquid Assayed Specific Protein Control Level 2 | 1782 | - | Randox Laboratories can confirm that the Rheumatoid Factor RF concentrations in Liquid x D Assayed Specific Protein Controls Level Lots LPC and LPC are low below the x D measurement range of nephelometric methods on Siemens instruments Randox have since x D updated the IFU s and removed the associated targets and ranges x D Randox Laboratories can also confirm that Anti Streptolysin O ASO target and range values x D have been reassigned in lot no s LPC LPC LPC LPC and LPC Please see x D table below for updated target values | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Download corrected IFU from Smartdocs.com and discard of previous copies. Please accept our apologies for any inconvenience caused. Thank you for your patience and understanding. If you have any questions or concerns, please contact Randox Technical Services. | 2025-04-13 15:20
Randox Laboratories Ltd | Liquid Assayed Specific Protein Control Level 3 | 1782 | - | Randox Laboratories can confirm that the Rheumatoid Factor RF concentrations in Liquid x D Assayed Specific Protein Controls Level Lots LPC and LPC are low below the x D measurement range of nephelometric methods on Siemens instruments Randox have since x D updated the IFU s and removed the associated targets and ranges x D Randox Laboratories can also confirm that Anti Streptolysin O ASO target and range values x D have been reassigned in lot no s LPC LPC LPC LPC and LPC Please see x D table below for updated target values | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Download corrected IFU from Smartdocs.com and discard of previous copies. Please accept our apologies for any inconvenience caused. Thank you for your patience and understanding. If you have any questions or concerns, please contact Randox Technical Services. | 2025-04-13 15:20
Welch Allyn | Connex Vital Signs Monitor (CVSM), CP150 Electrocardiograph & Connex Spot Monitor Accessory Power Management Stand (APM) | 2300 | - | THE USE OF AN UNAPPROVED BATTERY MAY LEAD TO COMBUSTION FIRE OR DEVICE PERFORMANCE ISSUES | 1. DO NOT USE UNAPPROVED BATTERIES. Baxter asks customers to only use the approved Baxter / Welch Allyn BATT99 battery, consistent with product labeling due to their safety and performance testing. 2. INSPECT ALL DEVICES. If you have a device identified in the affected product table above, check the device for the correct battery type. See Figure 1 below for examples of labels from approved batteries labelled with the BATT99 product code, “Welch Allyn”, and the Skaneateles Falls manufacturing facility address. Figure 1: Examples of Label for Approved Batteries 3. STOP USING THE DEVICE if an unapproved battery is found. 4. IMMEDIATELY REPLACE THE UNAPPROVED BATTERY with an approved battery. Please contact your local Baxter sales representative to order the approved Baxter / Welch Allyn BATT99 batteries. | 2025-04-13 15:36
CareFusion 303, Inc. | Alaris™ Pump | 640 | - | BD has identified that Alaris System Maintenance ASM testing was either not performed or incomplete as part of the preventative maintenance activities by BD technicians for the specific models and serial numbers of BD Alaris System with Guardrails Suite MX listed in Appendix A x D x D | Corrective actions by BD BD is investigating this issue. Corrective actions have been implemented to prevent recurrence. Actions to be taken by BD BD will be sending a technician to your site to perform and verify the necessary tests and complete the relevant documentation. Actions to be taken by Clinical Users Continue to use your BD Alaris™ System with Guardrails™ Suite MX devices in line with the applicable Directions for Use. If you experience any issue with using the device, contact BD at EMEA_productComplaints@bd.com. Actions to be taken by Biomedical Departments 1. Circulate this Field Safety Notice to all those within your organisation that may use the BD Alaris™ System with Guardrails™ Suite MX devices listed in Appendix A. 2. If you have further distributed the devices, please identify those users and notify them at once of this Field Safety Notice. 3. Locate the serial numbers listed in Appendix A and conduct the necessary preparations to make them available for a site visit by the BD technician. 4. Complete the Customer Response Form on page 3 and return it to Abdulrahman.Hanbazazah@bd.com or Ahmed.Shebah@bd.com as soon as possible or no later than the 14th May, 2025. o A BD technician will be in contact to schedule a site visit. | 2025-04-14 12:51
GE Healthcare Finland Oy, , located at Kuortaneenkatu 2, FI-00510 Helsinki, Finland. | CHA CC | 4 | - | GE HealthCare has become aware that the CHA system discontinues infusion orders with a Continuous timing schedule once the ordered amount is given even if the prescribed time duration is not met This can lead to undermedication if the intention was for the medication administration to continue until further notice or for a set duration x D x D There have been no injuries reported to GE HealthCare as a result of this issue x D | UMDC letter to be sent to the affected customers. Manual supplement will be delivered to the affected customers with detailed instructions | 2025-04-15 09:46
GE Healthcare Finland Oy, , located at Kuortaneenkatu 2, FI-00510 Helsinki, Finland. | CHA A | 4 | - | GE HealthCare has become aware that the CHA system discontinues infusion orders with a Continuous timing schedule once the ordered amount is given even if the prescribed time duration is not met This can lead to undermedication if the intention was for the medication administration to continue until further notice or for a set duration x D x D There have been no injuries reported to GE HealthCare as a result of this issue x D | UMDC letter to be sent to the affected customers. Manual supplement will be delivered to the affected customers with detailed instructions | 2025-04-15 09:46
Beckman Coulter Ireland, Inc | Creatinine | 632 | - | Beckman Coulter has determined through internal testing that the above listed lots of Creatinine reagent REF OSR do not consistently meet the icteric bilirubin interference specification for serum plasma as stated in the Creatinine Instructions for Use IFU x D Interference less than or mol L up to mg dL or mol L bilirubin |  Beckman Coulter recommends sharing the content of this letter with your laboratory and/or Medical Director to evaluate the requirement for a retrospective review of serum and plasma creatinine results for highly icteric patient samples measured with the identified lots.  Discontinue use and dispose of any remaining stock of the affected lots listed above. Contact your local Beckman Coulter representative for reimbursement and replacement of affected stock | 2025-04-15 23:51
Remote Diagnostic Technologies Ltd. | TEMPUS Pro Patient Monitor | 4 | - | When a user connects the Video Laryngoscope VL to the Tempus Pro Monitor during boot up the VL screen is displayed instead of the expected patient details dialog screen If the user attempts to capture an image to clear the screen an error dialog screen is shown that asks the user to restart the device x D | Philips is updating the Tempus Pro User/Operator Manual and the Tempus Pro Video Laryngoscope Supplement Guide to include the following Warning: DO NOT connect the video laryngoscope USB probe to the Tempus Pro during device boot-up. Once ‘Patient Details’ dialog appears, enter patient data then plug in the video laryngoscope USB probe. Additionally, the video laryngoscope start-up steps will be updated to plug the laryngoscope into either of the two USB ports on the Tempus Pro after boot-up sequence is complete. Philips will notify you again when the updated Tempus Pro User/Operator Manual and Tempus Pro Video Laryngoscope Supplement Guide is available for download. | 2025-04-16 09:54
Finland | Portrait Mobile Monitoring Solution | 1 | - | GE HealthCare has become aware of a software issue that can result in the loss of patient monitoring on the Portrait Central Viewer Application after days of continuous run time x D x D There have been no complaints or injuries reported as a result of this issue x D x D | A software update that will be installed by a GE HealthCare Field Engineer. The Field Engineer will remove any Portrait Core Services v1.1 and earlier from the customer site and destroy. | 2025-04-16 11:13
Abbott Diagnostics Scarborough, Inc. | ID NOW COVID-19 | 997 | - | Please find the attached FSN letter | Please find the attached FSN letter | 2025-04-17 14:34
Alere Scarborough Inc. | ID NOW Influenza A & B 2 | 997 | - | Please find the attached FSN letter | Please find the attached FSN letter | 2025-04-17 14:34
Alere Scarborough Inc. | ID NOW RSV | 997 | - | Please find the attached FSN letter | Please find the attached FSN letter | 2025-04-17 14:34
Cepheid | Xpert BCR-ABL Ultra p190 | 12 | - | Cepheid has determined that the potentially affected product noted above was stored outside of x D its labeled storage conditions i e C and has experienced a temperature excursion before it x D was delivered to you | Stop using the products listed above and we recommend re-testing of samples that were tested with cartridges from the affected products to assure accuracy of test results. You should document (see Response Form), stop using, and dispose of any remaining product from this lot that is in your inventory. Cepheid will provide replacement products used or in your current inventory. To receive your replacement product, please provide information on the Response Form acknowledging receipt of this letter by either emailing the completed form to CFQ@cepheid.com or faxing it to +1 (408) 716-3143. Replacement product should ship to you within 3-5 business days following receipt of the Response Form. | 2025-04-17 16:08
ICU Medical | ParaPAC plus™ | 14 | - | When a paraPAC plusTM ventilator is switched to the operating mode of Ventilate paraPAC plusTM ventilators may intermittently provide continuous positive gas flow instead of the intended cycling like a human breath This non cycling and continuous positive gas flow when in the cycling mode is a malfunction not allowing the ventilator to properly function as designed | Actions to be taken by the User/Customer: There is no need to return or discontinue using your paraPAC plusTM, at this time. When using the device, all instructions, including warnings and cautions in the User Manual must be followed with heightened awareness. This is inclusive, but not limited to the following: • Constant monitoring of the patient • Blood oxygenation and expired carbon dioxide levels should be monitored independently using pulse oximetry and capnography. • All pre-use checks must be performed before each use. • Alternative means of ventilation such as bag mask ventilation, must be available in the event of ventilator failure or malfunction. If the paraPac plusTM ventilator experiences continuous flow, remove the ventilator from clinical use, set the device aside for repair and use another device or alternative means of ventilation Report the continuous flow experience by filing a complaint, per instructions below. | 2025-04-17 18:19
Remote Diagnostic Technologies Ltd. | TEMPUS Pro Patient Monitor | 4 | - | I The device may fail to startup normally freeze or reboot Error messages may be displayed resulting in the device becoming inoperable These issues could occur during startup or clinical use in the following ways x D o The Tempus Pro Monitor may freeze at the startup screen display an MPM Application software error message or may show an Attention Restart Required message x D o The Tempus Pro Monitor may freeze and automatically reboot either when the user starts the process of recording a lead ECG by pressing the Start Record button or after the lead has been recorded when the user presses another button in the lead interface such as Print or View interpretation x D o The Tempus Pro Monitor may automatically reboot during an intermittent connection with the Tempus Pro Smart Mount x D x D II lead ECG measurements of the Tempus Pro Monitor generated by the standard Louvain algorithm could produce inaccurate ECG Rhythm and Morphology statements This issue impacts SW versions before vx x D x D These issues were identified via customer complaints and Philips internal investigation Philips has not received any reports of patient harm related to these issues x D x D Intended use of affected product x D The Tempus Pro is a portable vital signs monitor intended to be used by clinicians and medically qualified personnel for the attended or unattended monitoring of single or multiple vital signs in clinical and pre hospital care applications x D | Philips has developed a software update to resolve these issues that is expected to be available in Q3 2025 via FCO86700012. Customers will be contacted by Philips to implement the software update | 2025-04-18 13:43
siemens healthineers | STRATUS CS ACUTE CARE CTNL TEST PACK | 20 | - | manufacture lot problem | kits replacment | 2025-04-20 14:21
Beckman Coulter Inc. | AU480 and AU680 Clinical Chemistry Analyzers | 72 | - | Coulter has identified that using MU sample probes with lot numbers from to on AU and AU clinical chemistry analyzers There is remote probability of false low test results due to sample dispensing issues |  Check the analyzers, which meet the condition below, as they may include a defective sample probe. o Analyzers whose Serial No. is from 2023070718 to 2025021482. o Analyzers whose sample probe has been replaced since July 1, 2023.  For analyzers that meet either of the above criteria, remove the sample probe from the instrument in accordance with IFU Section 6, Maintenance / As needed Maintenance/Replace a Sample or Reagent Probe. And then, check the groove on the head of the removed sample probe. (see Fig.-1)  If it has “a marking line”, please continue to use.  If it does not have “a marking line”, please replace the sample probe with a probe not included in the impacted sample probe LOT numbers.  Analyzers that do not fall into the above condition can be used as is.  Please check your stock of sample probes as follows: o Check the outer box label of your sample probes. o If you have MU993400 sample probes with LOT numbers from 178713114 to 179433670, please dispose of the impacted probe to ensure it is not used and contact your BEC representative for a replacement (see Fig.-2). Note: Also check the sample probes included in the maintenance kits for AU480 (PN: B66751) and AU680 (PN: B66752) (see Fig.- 3). Contact your BEC representative for a replacement if impacted probes are identified.  If all the results of the actions above show only defective sample probes on your analyzers and in stock, stop measuring the assay items which sample volume is 2.0 uL or less and contact your BEC representative for a replacement.  Beckman Coulter recommends that customers who have replaced probes on your analyzers per this letter should share the content of this letter with your laboratory and/or medical director to assess if a retrospective review of the patient results is needed. | 2025-04-20 23:38
Beckman Coulter Inc. | Access EPO | 13 | - | Beckman Coulter has determined that the Access Erythropoetin EPO reagent lot number listed in the table above exhibits an approximate concentration bias of with patient samples when compared to alternate EPO reagent lots |  Discontinue using and discard all remaining reagent packs of the Access EPO reagent lot listed above.  Beckman Coulter recommends sharing the content of this letter with your laboratory and/or medical director regarding the need to review previous patient test results.  Contact your local Beckman Coulter representative for replacement product requests. o For customers in the United States and Canada: Complete the attached replacement order form. o For customers outside of the United States and Canada: Contact your local Beckman Coulter representative | 2025-04-20 23:59
Welch Allyn, Inc. | Connex Integrated Wall System | 1934 | - | n Baxter received a customer complaint for the Flexiport Reusable Blood Pressure Cuff kit reporting the packaging indicates not made with natural rubber latex however there was a rubber band inside each package x D On following the complaint investigation a nonconformance with an expanded investigation was initiated The FlexiPort Reusable and Two Piece Reusable blood pressure cuff labeling states the product is not made with natural rubber latex The media specification SAP DIR for the IFU specifies using a band to collate the pages but does not specify that the band must not be made from natural rubber latex Inquiries were made to the IFU suppliers Grupo Baca and Curtis Grupo Baca supplies IFUs for Baxter s Tijuana Mexico and Baxter s Skaneateles NY facilities and this supplier provided a certificate of conformance that the rubber bands are not made with natural rubber latex A different supplier Curtis supplies the IFUs for Baxter Skaneateles and informed Baxter that the rubber bands supplied with their IFUs are made with natural rubber latex x D On a non conforming materials report was created by supplier quality to contain devices with the impacted IFU and on an additional non conforming materials report was initiated to block additional subcomponents kits built with the IFU for the blood pressure cuffs part number x D On a hold was placed in the Skaneateles plant for the products containing not made with natural rubber latex labeling and an IFU banded by natural rubber latex x D On Baxter decided to issue an Urgent Medical Device Recall to inform customers of the issue | ctions to be Taken by Customers 1. The list of impacted product codes and UDI-DI numbers is in Attachment A. Immediately locate the impacted products at your facility that are unopened. Follow the instructions in Attachment A to determine if your unopened products are affected and how to arrange for return as appropriate. 2. Complete the enclosed customer reply form and return it to Baxter by either faxing it or scanning and e-mailing it or sending it by post even if you don’t have any inventory. Returning the customer reply form promptly will confirm your receipt of this notification and prevent you from receiving repeat notices. 3. If you distributed this product to other facilities or departments within your institution, please forward a copy of this communication to them. 4. If you purchased this product from a distributor, please note that the Baxter customer reply form is not applicable. If a reply form is provided by your distributor or wholesaler, please return it to the supplier according to their instructions. Actions to be Taken by Dealers, Wholesalers, Distributors/Resellers, or Original Equipment Manufacturers (OEM) 1. If you are a dealer, wholesaler, distributor/reseller, or original equipment manufacturer (OEM) and you have an affected product, please do not distribute. Contact Baxter for additional instructions. 2. If you are a dealer, wholesaler, distributor/reseller, or original equipment manufacturer (OEM) that distributed any affected product to other facilities, please conduct a consumer-level recall of the affected product that you distributed to customers. | 2025-04-21 14:34
Maquet Cardiopulmonary GmbH | Wall Holder HKH 8820 | 1 | - | The European Harmonized Standard DIN EN entitled Medical Vehicles and Their Equipment x D Road Ambulances section Holding Securing and Restraint Systems mandates that products such x D as the CARDIOHELP i System and the accessories utilized for transportation such as the HKH Wall x D Holder undergo the performance of a simulated crash test x D MCP conducted a crash test according to this Harmonized Standard on the HKH Wall Holder in December x D During this test the CARDIOHELP i system detached from the HKH Wall Holder resulting in the x D test not being passed Therefore the HKH Wall Holder does not comply with standard DIN EN x D All HKH Wall Holders are affected by this issue | Inform all customers possessing the affected products promptly about this Field Action by sending the Field Safety Notice for Customers. Develop a new design for the HKH 8820 Wall Holder, compliant with standard DIN EN 1789:2020. Upon availability of replacement products with new design: Inform all customers and replace customer devices. | 2025-04-22 15:15
Beckman Coulter Inc. | APO B | 1 | - | The Apo B Lot VAS LOSR ML contains an incorrect target value of mg dL for the level calibrator ODR Lot The correct level target value is mg dL The value in g L unit is correct | If your laboratory is reporting Apo B results generated using the combination of reagent OSR6143, Lot 2631 and calibrator ODR3022, Lot 3000 in mg/dL units, a retrospective review of patient results greater than 77 mg/dL and less than 176 mg/dL is recommended. Beckman Coulter requests that each customer  Discard the value assignment sheet located within each kit of Apo B OSR6143 Lot 2631.  Download the updated VAS kit insert from the Beckman Coulter website by searching ‘LOSR6143_2631_ML.02. Beckman Coulter Diagnostics ¦ Beckman Coulter  Ensure that the level 4 calibrator ODR3022, Lot 3000 target value used for calibration of Apo B OSR6143 Lot 2631 is 116 mg/dL and aligns with the corrected VAS. | 2025-04-22 23:14
Dexcom, Inc | G7 iOS CGM App | 1 | - | A bug was identified in the G iOS CGM App SW version which can cause the app to enter a state where it stops updating the estimated glucose value EGV and continues to show a stale EGV x D x D The Trend graph may continue to progress in time or may disappear New alerts and alarms will not trigger in this state however the pre scheduled Signal Loss alert and associated re alerts will occur in enabled of users have it enabled | Actions to be taken: Dexcom has access to contact all remaining users of the affected App version through deployment of in-App notifications and conduct a mandatory app update facilitated by blocking use of the affected app version | 2025-04-23 14:19
Smith Medical | Infusion pump | 8 | - | Damaged or dislodged battery separators or foreign material in the battery compartment may cause an electrical short condition between the battery contacts | 1. Inform all affected CADD-Solis and CADD-Solis VIP users (or potential users) of this notice and provide theinstructions below. 2. Visually inspect the external condition of the battery pack or Wireless Communication Module and look forevidence of any physical damage. As stated in the battery pack Instructions for Use, if the battery packhousing is cracked or otherwise damaged, replace the battery pack. NEVER use a battery pack that appearsdamaged. A rechargeable battery pack must be replaced with either another CADD-Solis rechargeablebattery pack or with 4 AA batteries. 3. Visually inspect the condition of the battery compartment for evidence of any physical damage. ContactTechnical Support for repairs. 4. Always check the battery compartment for fluid or debris before inserting the batteries, and do not allowany fluid or debris to fall into the battery compartment 5. Complete and return the attached Customer Response Form to EMEA-FSN@icumed.com within ten days ofreceipt to acknowledge your understanding of this notification. 6. DISTRIBUTORS: If you have distributed potentially affected products to your customers, please immediatelyforward this notice to them and request that they complete the response form and return it to YOU. Thenthe DISTRIBUTOR must complete a SINGLE form with the required details and return to EMEA-FSN@icumed.com | 2025-04-23 16:57
smith medical | infusion pump | 8 | - | Under certain conditions a CADD Solis pump may trigger an erroneous false Upstream Occlusion USO Alarm The erroneous USO alarm may occur when there is a delay of more than one hour between the first prime or infusion of a new CADD Administration Set and the next prime or infusion of the same CADD Administration Set The USO Alarm is a high priority alarm that will interrupt an ongoing infusion or delay initiation of an infusion The pump cannot resume or start an infusion until the alarm is cleared x D A USO Alarm may potentially occur if all these conditions are met x D x D Using a CADD administration set i e not a medication cassette reservoir and x D x D Enabling the Upstream Occlusion Alarm and x D x D Not programming Keep Vein Open KVO or Continuous rate and x D x D Priming or infusing shortly after attaching the administration set and the next infusion does not occurfor approximately one hour or longer x D This issue will not occur when using a medication cassette reservoir if the USO Alarm is disabled if the KVO setting is programmed or if a Continuous rate setting is programmed | Smiths Medical is sending this notification to all affected CADD-Solis and CADD-Solis VIP customers. An update to the Infusion Pump Operator’s Manual will be made to include these conditions. | 2025-04-23 17:11
Codman & Shurtleff, Inc | CODMAN® Disposable Perforator | 1220 | - | Integra LifeSciences is voluntarily issuing this Field Safety Notice for the recall of CODMAN Disposable Perforator mm listed in Table x D During an investigation of complaints Integra LifeSciences identified that there is an inadequate weld proud weld on specific lots of mm Codman Disposable Perforators that can potentially cause the product to disassemble break separate during its use | Actions to be taken by Customers: 1. Please review and understand the information provided in this letter. 2. Determine if the product you have is subject to the recall: a. Identify the impacted reference and lot number. b. See Appendix 2. below for a sample of product label for where to locate the reference and lot number. The lot numbers are 7 digits long (only numbers) c. Open the excel file, use the find function Ctrl+F or use the dropdown arrow on the top of the column and see if your lot number(s) is (are) on the list. 3. If you do have affected product(s): a. Quarantine the units immediately. b. Check the box “I do have affected units.” on the acknowledgement form. c. Record on the form the total quantity of affected products and lot number(s) that you have. 4. If you do not have affected product(s), check the box, “I do not have affected units.” 5. Please return the completed reply form by email to emea-fsca@integralife.com, By filling in this form, you confirm that you have received this Safety Notice, and you intend to fully comply with this notification. We expect a response within 3 weeks. You also confirm that this notification has been forwarded to every person concerned in your organization. 6. At receipt of your form, and if it is noted that you have affected units available for return, Integra Customer Service will contact you and provide a Return Material Authorization (RMA) number and directions to return the affected product(s). A credit note will be processed upon receipt and verification of returned goods. Note: credit will only be given for the impacted lot(s) that are returned. 7. We recommend that you retain a copy of the form for your records. PLEASE NOTE THAT REGARDLESS OF WHETHER YOU HAVE THE AFFECTED PRODUCTS TO RETURN OR NOT – A COMPLETED ACKNOWLEDGEMENT IS REQUIRED | 2025-04-27 13:38
Beckman Coulter Inc. | Access Testosterone Reagent Kit | 443 | - | Beckman Coulter has determined that Access Testosterone reagent packs PN from the lot listed in the table above may have experienced cross contamination between reagent pack wells during manufacturing |  Discontinue using and discard all remaining Access Testosterone reagent packs from the lot listed above.  Beckman Coulter recommends performing a retrospective review on patient results generated with the Access Testosterone lot listed above and assessing the need for retesting.  Contact your local Beckman Coulter representative for replacement product requests. | 2025-04-27 23:04
PHYSIO-ASSIST | SIMEOX_H_EU | 5 | - | The user manual supplied with the device lists the contra indications known at the time of writing After x D years on the market this list must be updated along with the list of expected side effects based on x D data collected during post marketing surveillance | Corrective action An addendum on the contra-indications and expected side-effects related to the use of SIMEOX_H is provided with this safety information. This addendum should be read and taken into account by all patients and healthcare professionals using Simeox_H. For distributors, clinics or medical practices: please complete the attached reply form and return it to us. Please respect this safety information and all related measures. Distribution This Safety Notification must be distributed to all persons who need to be informed within your institution, to any institution where the devices have been transferred and to all patients using a SIMEOX_H at home. Please take into account this notification and the resulting actions during the period of use of the device to ensure the effectiveness of the corrective actions. In the event that you, as the customer, choose not to implement the required corrective action as described above, PHYSIO-ASSIST will not be responsible for any safety related issues or legal liability arising from your failure to respond to this Safety Notification. | 2025-04-28 08:41
Covidien llc | Hugo™ RAS Tower (120VAC), Hugo™ RAS Tower (240VAC) | 2 | - | Complaints were received reporting a loss of motion during use of the Hugo Robotic Assisted Surgery RAS System involving the use of the Monopolar Curved Shears Model MRASI x D The loss of motion has been found to occur intraoperatively after the user has commanded the shears to perform consecutive small actuations opening and closing of the jaw which can cause the system to halt the teleoperation of the arm In response to the system halt the triggering of an error notification code NOTCODE occurs when sensing excessive torque Halting of teleoperation is an intended response function of the system to prevent further use of the device until the issue is resolved or the portion of the system with issue is replaced x D An evaluation was performed on the Monopolar Curved Shears MCS to evaluate the over torque conditions after consecutive small actuations of the instrument The testing indicates that the issue is specific to the MCS as the instrument has specific design of the blades The torque and shear forces necessary by the shear controller during the small actuations opening and closing of the shears results in the potential to exceed torque thresholds and the NOTCODE x D | • Notify relevant personnel in which the HugoTM RAS system is used about this medical device field safety notice in alignment with internal hospital procedures. • If you experience this issue, follow the steps provided in the supplemental labeling and report any incidents related to this issue to your local Medtronic representative. • Complete the attached Customer Acknowledgement Form. • Please maintain a copy of this notice in your records. • Utilize the below workflow for troubleshooting disabled instruments until new labeling is provided by Medtronic. | 2025-04-28 15:41
Olympus | SOLTIVE™ Premium SuperPulsed Laser System | 9 | - | Software update to the Olympus SOLTIVE SuperPulsed Laser System SOLTIVE Laser models Pro TFL SLS and Premium TFL PLS The SOLTIVE Laser is intended for incision excision resection ablation coagulation hemostasis and vaporization of soft tissue with or without an endoscope in urology lithotripsy gastroenterological surgery and gynecological surgery x D The attached Addendum to the SOLTIVE Laser System Instructions for Use details the new treatment parameters of the Procedure Presets to be implemented through the upcoming software update This is an update for FSCA SA and SA x D | To update the software and to deliver IFU Addendum corresponding with this update. The full IFU can be located electronically at https://www.olympus-europa.com/medical/en/Contact-and-Support/search_page.html | 2025-05-01 10:06
GE Healthcare Finland Oy, , located at Kuortaneenkatu 2, FI-00510 Helsinki, Finland. | CHACC and CHAA | 7 | - | GE HealthCare has become aware of a potential cybersecurity vulnerability within one of the CHA components Exposure of this potential vulnerability is limited to the local network x D x D There have been no injuries exploits of this potential vulnerability or unauthorized access to patient data reported to GE HealthCare as a result of this issue x D x D | Customers having the affected CHA version (v4.2 – 5.8.2B) UMDC letter will be sent. The CHA system will be upgraded to the CHA patch 5.8.2C. The deployment can be performed remotely or in-person by a qualified GEHC employee/distributor. No removal actions needed as this is sw only product. | 2025-05-01 13:37
Siemens Healthcare Diagnostics Inc. | Atellica IM HCY CAL | 5 | - | Issue Description x D Siemens Healthineers has confirmed through the investigation of customer complaints that the Atellica IM Homocysteine Calibrator and ADVIA Centaur HCY Calibrator Lots listed in the table above do not meet the assay Instructions for Use IFU reference interval due to calibrator instability As the calibrator ages homocysteine results were observed to increase with bias up to at months post manufacture x D The described calibrator instability impacts homocysteine on the Atellica CI Analyzer Atellica IM Analyzer ADVIA Centaur XP System ADVIA Centaur XPT System and ADVIA Centaur CP System x D x D Kindly review the attached Urgent Field Safety Notice | Customer Actions • Perform the instructions provided below: o Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. o Discontinue use of and discard the kit lots listed in the table above (Products Section). o Review your inventory of these products to determine your laboratory’s replacement needs and to provide information to Siemens Healthineers for reporting to the authorities. o Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. o Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached Urgent Field Safety Notice | 2025-05-01 14:42
Siemens Healthcare Diagnostics Inc. | Advia Centaur Cal HCY CALIBRATOR | 5 | - | Issue Description x D Siemens Healthineers has confirmed through the investigation of customer complaints that the Atellica IM Homocysteine Calibrator and ADVIA Centaur HCY Calibrator Lots listed in the table above do not meet the assay Instructions for Use IFU reference interval due to calibrator instability As the calibrator ages homocysteine results were observed to increase with bias up to at months post manufacture x D The described calibrator instability impacts homocysteine on the Atellica CI Analyzer Atellica IM Analyzer ADVIA Centaur XP System ADVIA Centaur XPT System and ADVIA Centaur CP System x D x D Kindly review the attached Urgent Field Safety Notice | Customer Actions • Perform the instructions provided below: o Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. o Discontinue use of and discard the kit lots listed in the table above (Products Section). o Review your inventory of these products to determine your laboratory’s replacement needs and to provide information to Siemens Healthineers for reporting to the authorities. o Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. o Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached Urgent Field Safety Notice | 2025-05-01 14:42
Abbott Molecular Division Inc. | Alinity m STI AMP Kit | 1 | - | Please find the attached FSN letter | Please find the attached FSN letter | 2025-05-04 16:45
OrthoPediatrics Corp,. | Small Bone Clamp | 6 | - | OrthoPediatrics has identified Small Bone Clamps that have been manufactured with a retempering process leading to embrittlement causing the clamp to fracture There have been no adverse events reported to date and four complaints were reported concerning this issue | The affected product must be returned to Orthopediatrics. A replacement product will be provided. Please contact customer service to arrange return of the affected product. Distributors should retrieve affected products that may be with customers. Please complete and return the attached response acknowledgement form confirming your understanding of these instructions. | 2025-05-05 15:14
3M Health Care | 3M™ Ranger™ Blood/Fluid Warmer | 9488 | - | The design requirements did not clearly specify the location of the fluid x D temperature measurement and corresponding flow rate based on fluid inlet x D temperature | Ensure all your internal users are informed about this corrective action and Keep this notice on hand until the updated IFUs and operator’s manuals are available. | 2025-05-06 10:25
3M Health Care | 3M Ranger™ High Flow Blood/Fluid Disposable Set | 9488 | - | The design requirements did not clearly specify the location of the fluid x D temperature measurement and corresponding flow rate based on fluid inlet x D temperature | Ensure all your internal users are informed about this corrective action and Keep this notice on hand until the updated IFUs and operator’s manuals are available. | 2025-05-06 10:25
Meridian Bioscience Inc. | ImmunoCard STAT! Crypto/Giardia | 2 | - | A single lot of the ImmunoCard STAT Crypto Giardia Test Kit was inadvertently stored in Medline warehouses at temperatures outside of the labeled storage requirements | 1. Immediately check your stock for the affected item number and the affected lot numbers which can be located within the recall portal. 2. Please use the link and the information below to complete your response form. Please list the quantity of affected product you have in inventory on the form. Even if you do not have any affected product in inventory, please complete and submit the response form. 3. Please destroy any affected inventory. Medline will issue credit for the product based upon the quantity and lot numbers identified on the submitted response form The login for completing the response form is: Website link: https://recalls.medline.com Recall Reference #: R-25-063 Recall Code: Recall Code 4. If you are a distributor, or have resold or transferred this product to another company or individual, you are required by FDA regulations to notify them of this recall communication. Please request that customers return any affected product to you for collection and return to Medline Industries, LP. | 2025-05-07 12:47
Randox Laboratories Ltd | Urea | 5 | - | Randox is conducting a Device Recall for Urea reagent catalogue number UR batch The R b Enzyme Reagent is no longer meeting its reconstituted stability claims of weeks at to C or days at to C The calibration factor QC s and patient samples may decrease during repeat testing by up to within hrs of reagent reconstitution | Action to be taken: Discontinue use of and discard any of the above immediately. Provide Randox with photographic evidence of the destruction of the kits. Review your reagent inventory of these products and assess your laboratories needs for reimbursement for discarded inventory. Review results generated with the affected batches in line with the clinical profile of the patient. Discuss the contents of this notice with your Medical Director to determine the appropriate course of action, including for any previously generated results. If applicable follow your laboratories previously reported patient results protocol regarding the need for reviewing Complete and return the response form 12187-QA to technical.services@randox.com within five working days. | 2025-05-08 17:49
Alcon Laboratories, Inc. | TOTALPLUS | 2268 | - | There is potential for some trays within impacted lots to have damage in the Tyvek cover | We are asking that you locate and dispose of any affected lots of Alcon Constellation Pak® remaining in your inventory, To comply with this FIELD SAFETY CORRECTIVE ACTION and request the replacement of any unused product, please take the following steps: 1. Review your inventory to determine if you have any unused affected product within your facility. See table on page 1 for affected Constellation Pak® lots shipped to your location. 2. Segregate and dispose of any unused affected product from your inventory. 3. Post this notification letter near where affected products are stored to ensure awareness of the recall and continued shipment of corrected Constellation Pak®. 4. Respond to Alcon indicating your understanding of these instructions even if you have zero (0) units remaining in inventory by completing and returning the attached “Response Form” and returning to Alcon via email. 5. Please forward this notification to all departments within your organization who may be in possession of this affected product; and any other organization to which this product may have been transferred. | 2025-05-08 18:32
Siemens healthineers | Chek-Stix® UA Diptube Liquid QC, 15 mL | - | - | Siemens Healthineers received a Product Correction Market Withdrawal Notification from Kova International regarding foreign material identified in some quality control products The affected products lots are listed in the table below These products are for use on the CLINITEK Status and Clinitek Advantus instruments For more details about the issue and the actions to be taken please refer to the attached Product Correction Market Withdrawal Notification from Kova International | Review Kova’s Product Correction - Market Withdrawal Notification attached. • Review this letter with your Medical Director and/or Point of Care Supervisor to determine the appropriate course of action. • If you are a distributor, ensure your customers receive this customer notification letter. • Retain this letter with your laboratory records and forward this letter to those who may have received this product in your facility. • If you have any of the kit lots listed above, complete and return the Product Replacement Form attached to this letter. Please contact Siemens Healthineers Technical Solution Center or your local Siemens Healthineers representative for further questions. | 2025-05-11 14:28
Siemens Healthineers | Chek-Stix® UA Diptube Liquid QC, 15 mL | 429 | - | Siemens Healthineers received a Product Correction Market Withdrawal Notification from Kova International regarding foreign material identified in some quality control products The affected products lots are listed in the table below These products are for use on the CLINITEK Status and Clinitek Advantus instruments For more details about the issue and the actions to be taken please refer to the attached Product Correction Market Withdrawal Notification from Kova International | Review Kova’s Product Correction - Market Withdrawal Notification attached. • Review this letter with your Medical Director and/or Point of Care Supervisor to determine the appropriate course of action. • If you are a distributor, ensure your customers receive this customer notification letter. • Retain this letter with your laboratory records and forward this letter to those who may have received this product in your facility. • If you have any of the kit lots listed above, complete and return the Product Replacement Form attached to this letter. Please contact Siemens Healthineers Technical Solution Center or your local Siemens Healthineers representative for further questions. | 2025-05-11 14:36
Randox Laboratories Ltd | Liquid Urine Control – Level 2 | 101 | - | Randox can confirm there has been a transcription error in the Instructions for Use IFU for x D Liquid Urine Control Level lot UC and Assayed Urine Control Level lot UC x D For Liquid Urine Control Level lot UC the targets and ranges for the analyte Creatinine x D in units mmol l have been listed incorrectly in the Method Siemens Atellica Solution and x D Siemens Dimension RxL Max Xpand EXL sections on the IFU | Action to be taken: Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Download corrected IFU from Smartdocs.com and discard of previous copies. Please accept our apologies for any inconvenience caused. Thank you for your patience and understanding. If you have any questions or concerns please contact Randox Technical Services. | 2025-05-12 18:08
Randox Laboratories Ltd | Assayed Urine Control – Level 3 | 101 | - | Randox can confirm there has been a transcription error in the Instructions for Use IFU for x D Liquid Urine Control Level lot UC and Assayed Urine Control Level lot UC x D For Liquid Urine Control Level lot UC the targets and ranges for the analyte Creatinine x D in units mmol l have been listed incorrectly in the Method Siemens Atellica Solution and x D Siemens Dimension RxL Max Xpand EXL sections on the IFU | Action to be taken: Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Download corrected IFU from Smartdocs.com and discard of previous copies. Please accept our apologies for any inconvenience caused. Thank you for your patience and understanding. If you have any questions or concerns please contact Randox Technical Services. | 2025-05-12 18:08
Werfen | ACL TOP Family / ACL TOP Family 50 Series / ACL TOP Family 70 Series, including ACL TOP 970 CL | 108 | - | We have become aware of a potential carryover issue that may cause elevated quality control and sample results when HemosIL Free Protein S Part No is tested in the same run as HemosIL D Dimer HS Part No The reagent carryover may result in a falsely higher Free Protein S result increasing the risk of excluding a protein S deficiency However a misclassification is not likely to result in harm to the patient that requires medical or surgical intervention x D Regardless of your territory sells both reagents the following parameter upgrade is mandatory x D Mandatory Parameters P for the ACL TOP Family SW x D Mandatory Parameters P SW and P SW for the ACL TOP Family Series and x D Mandatory Parameters P SW or a higher version for the ACL TOP Family Series x D contain the final mitigations to address this carryover issue x D x D The following instrument models are included under each Family of Instruments x D x D The ACL TOP Family includes the ACL TOP ACL TOP CTS ACL TOP ACL TOP CTS ACL TOP LAS ACL TOP CTS ACL TOP CTS models x D x D The ACL TOP Family Series includes the ACL TOP ACL TOP CTS ACL TOP LAS ACL TOP CTS ACL x D TOP CTS models x D x D The ACL TOP Family Series includes the ACL TOP s LAS and models x D | Customers are requested to take the following actions: • Share this information with your laboratory staff and update your internal procedures, as needed. • Forward this notification to all affected locations in your facility. • Post this notification on all affected systems in your facility. • Contact your local Werfen Representative to install the Mandatory Test Parameters at your next service or applications visit, no later than December 30, 2026 • Retain a copy of this notification for your records. | 2025-05-14 09:28
Bard Access System | PowerPICC Catheter | 460 | - | In March BD issued a global product removal due to an increase of material fatigue leaks on Fr single lumen PowerPICC catheters both SOLO and non SOLO versions in specific geographies These leaks are primarily characterised by a transverse circumferential crack in the catheter tubing Figure x D Based upon customer feedback and continued investigations BD has further strengthened the specification associated with the MFI to assure product performance which has resulted in an increased scope of the FSN see attachment and is clarifying its instructions for customers x D x D x D x D Figure Example of transverse circumferential crack in the catheter tubing x D BD s investigation has identified certain factors that contribute to material fatigue leaks in the Fr single lumen PowerPICC catheter tubing specifically x D The resin used to extrude the catheter tubing exceeded our supplier s specification for a material property called melt flow index MFI BD s investigation has concluded that a higher MFI could make the PowerPICC catheter tubing more susceptible to leakage when placed under stress x D Insertion and securement techniques that contradict the PowerPICC IFU requirements such as the requirement to insert the catheter as close to the zero mark as possible x D o BD PICCs have a taper design that increases in diameter near the zero centimeter mark on the device x D o When using any BD PICC ensure practice is consistent with BD IFUs for insertion depth and securement including insertion to the zero mark Fully insert the PICC as close as possible to the zero centimeter mark position B in Figure This allows the kink resistant tapered region to be utilised and is associated with lower catheter leakage rates x D o Select securement devices that can be used consistent with the BD PICC IFU x D x D x D Figure Image of the PowerPICC Taper Region x D x D | Clinical User Actions Consider the patient's infusion needs, alternative access options and the risks and benefits of continued catheter usage. BD is not recommending to explant product in-situ from any recalled lots unless catheter damage is suspected, as outlined below: Actions if catheter damage is not suspected: 1. Carefully examine the visible portion of the catheter to assess for any sign of damage to the catheter shaft. 2. Monitor the patient closely for signs and symptoms of catheter damage, such as increased extremity circumference, infusate leakage, or reports of pain. 3. When using any BD PICC ensure practice is consistent with BD Instructions for Use for insertion depth and securement, including insertion to the zero mark. If the IFU for a securement device contradicts the IFU for the BD PICC, follow the IFU for BD’s PICCs. Actions if catheter damage is suspected: 1. Immediately stop any infusion if catheter damage is suspected. 2. Follow your institution's guidelines for catheters with suspected damage. 3. If the catheter is confirmed to have a fracture, the catheter should be removed and an alternative route for access should be obtained. Action TAKEN by BD • BD has implemented additional controls around MFI that provides increased level of product assurance. BD Actions • BD is removing additional unexpired lots of 4Fr Single Lumen PowerPICC, both SOLO and non-SOLO versions as described in Attachment 1. • BD is notifying users to ensure practice is consistent with BD Instructions for Use (IFU) for insertion depth and securement. • BD will provide product replacement or credit for all destroyed products. Customer Actions: • Cease use of any unused affected lot numbers of 4Fr Single Lumen PowerPICC. • Identify and quarantine all unused affected lot numbers of 4Fr Single Lumen PowerPICC. • Make a note of the lot numbers and immediately destroy all unused affected units. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 16th June 2025. • Circulate this notice to all those who need to be aware within your organisation or to any organisation where the potentially affected products have been transferred. • Please report all device-related incidents to the manufacturer, distributor or local representative, and the national Competent Authority if appropriate, as this provides important feedback. Distributor Actions: • Cease distribution of any unused affected lot numbers of 4Fr Single Lumen PowerPICC. • Identify, quarantine, making a note of the lot numbers then destroy all undistributed affected lot numbers of 4Fr Single Lumen PowerPICC. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 16th June. • Complete and return the Customer Response Form following completion of your reconciliation activities. • Please report all device-related incidents to the manufacturer, distributor or local representative, and the national Competent Authority if appropriate, as this provides important feedback. | 2025-05-14 17:44
Olympus | Duodenovideoscope(Evis Exera III,Evis Lucera Elite) | 22 | - | When manual cleaning was delayed beyond one hour and a presoak was performed compared to those TJF duodenoscopes where manual cleaning began within one hour after patient procedure The current TJF reprocessing manual Section Presoak the endoscope includes validated instructions to perform an extended detergent soak of the TJF Q V V V in cases where manual cleaning cannot begin within one hour and is delayed up to hours after patient procedure However as a result of this post market surveillance data and as a precautionary measure Olympus no longer includes an option for delaying manual cleaning beyond one hour after patient procedure The reprocessing manual for the TJF Q V V V is being updated to remove the presoak the endoscope step and will include instructions that require manual cleaning to begin within one hour after patient procedure for the TJF duodenoscope Olympus review of complaint data identified four reports of potential delayed cleaning None of these complaints reviewed included a report of infection or serious injury | 1. Ensure all reprocessing and endoscopy unit personnel are completely knowledgeable and thoroughly trained on the updated reprocessing instructions contained in this letter when using the TJF duodenoscope. 2. Acknowledge receipt of this letter. | 2025-05-15 01:21
Stryker Instruments | System 8 Reciprocating Saw | 60 | - | There is a potential for the Compression Spring inside the saw s drivetrain assembly to disengage during blade loading and unloading A disengaged spring disables the blade collar from springing back to its original home position potentially leading to blade retention failure when loading a blade | 1. Review your inventory to locate and quarantine any affected products at your facility. Reach out to your Sales Representative for assistance in locating your serial number(s). 2. Return the enclosed Business Reply Form (BRF), even if the affected product is no longer in inventory, to the email xxx@stryker.com. 3. Upon receipt of your BRF, Stryker In-house Service will contact you with the shipping information to return your device for repair. 4. Your device will be returned to you once repairs have been completed. 5. Maintain awareness of this communication internally until all required actions have been completed. | 2025-05-15 16:14
Stryker Instruments | SYSTEM 8 STERNUM SAW | 60 | - | There is a potential for the Compression Spring inside the saw s drivetrain assembly to disengage during blade loading and unloading A disengaged spring disables the blade collar from springing back to its original home position potentially leading to blade retention failure when loading a blade | 1. Review your inventory to locate and quarantine any affected products at your facility. Reach out to your Sales Representative for assistance in locating your serial number(s). 2. Return the enclosed Business Reply Form (BRF), even if the affected product is no longer in inventory, to the email xxx@stryker.com. 3. Upon receipt of your BRF, Stryker In-house Service will contact you with the shipping information to return your device for repair. 4. Your device will be returned to you once repairs have been completed. 5. Maintain awareness of this communication internally until all required actions have been completed. | 2025-05-15 16:14
Stryker Instruments | SYSTEM 9 RECIPROCATING SAW | 60 | - | There is a potential for the Compression Spring inside the saw s drivetrain assembly to disengage during blade loading and unloading A disengaged spring disables the blade collar from springing back to its original home position potentially leading to blade retention failure when loading a blade | 1. Review your inventory to locate and quarantine any affected products at your facility. Reach out to your Sales Representative for assistance in locating your serial number(s). 2. Return the enclosed Business Reply Form (BRF), even if the affected product is no longer in inventory, to the email xxx@stryker.com. 3. Upon receipt of your BRF, Stryker In-house Service will contact you with the shipping information to return your device for repair. 4. Your device will be returned to you once repairs have been completed. 5. Maintain awareness of this communication internally until all required actions have been completed. | 2025-05-15 16:14
Stryker Instruments | SYSTEM 9 - STERNUM SAW | 60 | - | There is a potential for the Compression Spring inside the saw s drivetrain assembly to disengage during blade loading and unloading A disengaged spring disables the blade collar from springing back to its original home position potentially leading to blade retention failure when loading a blade | 1. Review your inventory to locate and quarantine any affected products at your facility. Reach out to your Sales Representative for assistance in locating your serial number(s). 2. Return the enclosed Business Reply Form (BRF), even if the affected product is no longer in inventory, to the email xxx@stryker.com. 3. Upon receipt of your BRF, Stryker In-house Service will contact you with the shipping information to return your device for repair. 4. Your device will be returned to you once repairs have been completed. 5. Maintain awareness of this communication internally until all required actions have been completed. | 2025-05-15 16:14
Ortho Clinical Diagnostics | ORTHO VISION Analyzer | 132 | - | The purpose of this notification is to provide information regarding a software anomaly identified on the ORTHO VISION and ORTHO VISION Max analyzers where the analyzer is unable to perform inventory of loaded samples and reagents if the loading station door is opened immediately after the allotted second countdown in the second window before the door re locks | • When accessing the door to the loading station visually confirm that the timer has not reached the end of its 20 second window. If so, let the door lock reengage before accessing the load station through the user interface. • If the loading door was opened at the '0:00' time, as timer expires, and no inventory is detected on the GUI, ensure a new loading event is initiated. To initiate a new loading event, open the loading station door, wait a few seconds and subsequently close it to facilitate an accurate inventory event. • Complete and return the enclosed Confirmation of Receipt form no later than April XX,2025. • Please forward this notification if the affected product was distributed outside of your facility | 2025-05-15 19:43
Ortho Clinical Diagnostics | ORTHO VISION Max Analyzer | 132 | - | The purpose of this notification is to provide information regarding a software anomaly identified on the ORTHO VISION and ORTHO VISION Max analyzers where the analyzer is unable to perform inventory of loaded samples and reagents if the loading station door is opened immediately after the allotted second countdown in the second window before the door re locks | • When accessing the door to the loading station visually confirm that the timer has not reached the end of its 20 second window. If so, let the door lock reengage before accessing the load station through the user interface. • If the loading door was opened at the '0:00' time, as timer expires, and no inventory is detected on the GUI, ensure a new loading event is initiated. To initiate a new loading event, open the loading station door, wait a few seconds and subsequently close it to facilitate an accurate inventory event. • Complete and return the enclosed Confirmation of Receipt form no later than April XX,2025. • Please forward this notification if the affected product was distributed outside of your facility | 2025-05-15 19:43
Philips Medical Systems Nederland B.V. | Zenition 50 | 3 | - | A wireless foot switch pedal may get stuck in the active position when the user releases the pedal resulting in the emission of unintended radiation because of x D Build up of dense or sticky fluids on the wireless foot switch e g blood or contrast fluid if not properly cleaned x D Use of protective covers that are either the wrong size or incorrectly placed on the wireless foot switch x D Dislodgment of a screw holding the pick up bar see Fig of the wireless foot switch subsequently becoming lodged within the wireless foot switch housing x D An additional screw inadvertently left inside the wireless foot switch housing during manufacturing | Philips will contact customers to schedule a visit to inspect the wireless foot switch and hand over an IFU addendum (reference FCO71800111). | 2025-05-16 14:41
Philips Medical Systems Nederland B.V. | Zenition 70 | 3 | - | A wireless foot switch pedal may get stuck in the active position when the user releases the pedal resulting in the emission of unintended radiation because of x D Build up of dense or sticky fluids on the wireless foot switch e g blood or contrast fluid if not properly cleaned x D Use of protective covers that are either the wrong size or incorrectly placed on the wireless foot switch x D Dislodgment of a screw holding the pick up bar see Fig of the wireless foot switch subsequently becoming lodged within the wireless foot switch housing x D An additional screw inadvertently left inside the wireless foot switch housing during manufacturing | Philips will contact customers to schedule a visit to inspect the wireless foot switch and hand over an IFU addendum (reference FCO71800111). | 2025-05-16 14:41
Philips Medical Systems Nederland B.V. | Veradius Unity | 3 | - | A wireless foot switch pedal may get stuck in the active position when the user releases the pedal resulting in the emission of unintended radiation because of x D Build up of dense or sticky fluids on the wireless foot switch e g blood or contrast fluid if not properly cleaned x D Use of protective covers that are either the wrong size or incorrectly placed on the wireless foot switch x D Dislodgment of a screw holding the pick up bar see Fig of the wireless foot switch subsequently becoming lodged within the wireless foot switch housing x D An additional screw inadvertently left inside the wireless foot switch housing during manufacturing | Philips will contact customers to schedule a visit to inspect the wireless foot switch and hand over an IFU addendum (reference FCO71800111). | 2025-05-16 14:41
Olympus | SOLTIVE™ Premium SuperPulsed Laser System | 29 | - | A fiber break at the strain relief or along the length of the fiber is the cause of or contributes to these incidents Olympus is continuing to investigate the issue and will take appropriate actions regarding the investigation findings | Ensure all surgeons and personnel are completely knowledgeable and thoroughly trained on the content of this letter, the instructions provided for detecting fiber damage, Appendix 1, and the SOLTIVE Laser System and Fibers Instructions for Use. Include a copy of this letter with the Instructions for Use. You may continue to use the device as per this letter and the Instructions for Use. | 2025-05-18 17:37
Ethicon Endo-Surgery, LLC | Endopath Echelon™ Vascular Reloads | 73 | - | Ethicon Endo Surgery LLC has received an increase in reports regarding inadvertent instrument lockout during surgical procedures involving ENDOPATH ECHELON Vascular White Reload for Advanced Placement Tip mm Row product code VASECR This reload is used exclusively with the ECHELON Flex Powered Vascular Stapler product code PVE A The instrument s lockout feature is designed to prevent a used or improperly installed reload from being fired or an instrument from being fired without a reload x D If a user encounters this issue the device will momentarily activate but will not cut or staple tissue and additional steps will be required to open it and remove it from tissue This communication includes reinforcement of mitigating instructions for addressing lockouts which can be found in the current Instructions for Use IFU for the ECHELON Flex Powered Vascular Stapler product code PVE A | 1. Share this notification with all users of ENDOPATH ECHELON™ Vascular White Reload for Advanced Placement Tip (35 mm, 4 Row), product code VASECR35. 2. Confirm that personnel using ENDOPATH ECHELON™ Vascular White Reload for Advanced Placement Tip (35 mm, 4 Row), product code VASECR35, understand the Instructions for Use (IFU) for the ECHELON™ Flex Powered Vascular Stapler (product code PVE35A) and this notification. 3. Post a copy of this communication. 4. If any subject product has been forwarded to another facility, contact that facility to share this information. Please share a copy of this notification when communicating. 5. Complete the Business Reply Form (BRF) Attachment 2 confirming receipt of this notice and e-mail it to HAlzamel@ITS.JNJ.com within three (3) business days. Please return the BRF even if you do not have product subject to this Field Safety Corrective Action. | 2025-05-19 09:43
Fritz Stephan GmbH | SOPHIE | 4 | - | Fritz Stephan GmbH has received reports about the failure of the Sophie touch display that can lead to situations in which the device can no longer be operated x D If the touch input fails in such a way that the touch functionality is no longer available the device can still be operated using a rotary push button x D However the investigation showed that also touch failures may occur in which the touch controller sends irregular signals that could be interpreted as touch inputs by the device software In this case it is no longer possible to operate the device using the rotary push button as the incorrect touch signals always overwrite the inputs via the rotary push button | The false positive touch input signals are caused by to close tolerances in the touch calibration data in the controller firmware. Fritz Stephan GmbH has already released a new touch controller firmware (V2025-02-06) for Sophie 2020 that will avoid these false positive touch input signals. | 2025-05-19 11:39
Becton Dickinson Infusion Therapy Systems, Inc | BD Insyte™ Autoguard™ BC Pro | 37600 | - | Based on customer feedback BD has identified that during insertion of the catheter the needle is slow to retract or fails to retract The slow or no retraction may require repeated attempts at retraction of the needle by pressing the button There is also the potential for a needle to remain exposed despite shielding attempts As a result we are removing specific lots listed | Clinical User Actions • Cease use and destroy any unused affected devices. • For devices that have deployed and are currently in use or if an impacted device was previously used, no additional follow-up activities are required for patients that have completed therapy. BD Actions BD is continuing to investigate the root cause and will take action to prevent reoccurrence of this issue. BD will provide replacement for all affected product. Customer Actions: • Cease use of any unused affected Insyte™ Autoguard™ Shielded IV Catheter. • Identify and quarantine all unused affected Insyte™ Autoguard™ Shielded IV Catheter. • Make a note of the lot numbers and destroy all unused affected units. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 19th June 2025. • Circulate this notice to all those who need to be aware within your organization or to any organization where the potentially affected products have been transferred. • If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: • Cease distribution. • Identify, quarantine, making a note of the lot numbers then destroy all undistributed affected Insyte™ Autoguard™ Shielded IV Catheter. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 19th June 2025. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. | 2025-05-20 11:04
Becton Dickinson and Company | BD BACTEC™ MGIT™ 960 PZA Kit | 357 | - | In mid BD initiated a field action for the BD BACTEC MGIT PZA Kits Later that year following an internal investigation and based on positive expanded performance test results BD implemented a modification to the raw material to address the issue BD resumed production while conducting further stability testing x D However BD has received additional complaints regarding the affected product of intermittent false resistance results for pyrazinamide PZA during susceptibility testing of Mycobacterium tuberculosis isolates Concurrently BD s additional stability testing suggests that the implemented corrective action may not fully resolve the issue and further investigation is necessary x D | Clinical risk PZA is a widely used component in the treatment of tuberculosis, and its exclusion based on false resistance results can result in a less optimal treatment regimen. This could include an extended length of treatment and increased risk of medication side effects, such as hepatotoxicity, peripheral neuropathy, and hypersensitivity reactions. To date, there has been four (4) adverse events worldwide related to this issue. Clinical User Actions • Customers should discard the affected product immediately. • There are no recommendations for retesting or reviewing previous patient test results; however, this does not preclude any clinician or institution from performing additional retesting or review if they deem necessary and have resources available to them. Laboratories should be aware that the availability and validation status of PZA Drug Susceptibility Testing (DST) methods vary by country and laboratory setting. We recommend consulting your national TB program, public health authority, or reference laboratory to determine the most appropriate testing strategy based on local testing infrastructure and regulatory guidance. Action taken by BD • BD has paused manufacturing of the affected product and will issue credit as applicable. Replacement product is currently unavailable. • BD will continue to investigate this issue. • BD will communicate when supply is available or when appropriate next steps are decided. Customer Actions: • Cease use of any unused affected lots of BD BACTEC™ MGIT™ 960 PZA Kit. • Identify and quarantine all unused affected lots of BD BACTEC™ MGIT™ 960 PZA Kit. • Make a note of the lot numbers and destroy all unused affected units. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 30th June 2025. • Circulate this notice to all those who need to be aware within your organization or to any organization where the potentially affected products have been transferred. • If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: • Cease distribution. • Identify, quarantine, making a note of the lot numbers then destroy all unused affected lots of BD BACTEC™ MGIT™ 960 PZA Kit. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 30th June 2025. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. | 2025-05-20 11:28
DeVilbiss Healthcare LLC | iGo2 | 155 | - | The iGo POC car charger cord is a non critical accessory to the iGo POC It allows the user x D to operate their iGo POC from a DC power supply such as the charging port in a vehicle x D Under normal operating conditions the car charger cord will power the unit while concurrently x D charging the device s internal battery x D After receiving complaints related to overheating of the plug portion of the car charger cord an x D investigation determined the root cause was an assembly error specifically the fuse and the x D retention spring were installed in the reverse positions It was confirmed that during use the x D adapter connection may become hot to the touch and in some cases cause deformation of the x D plastic housing The manufacturing assembly error has now been corrected x D Drive Devilbiss Healthcare is voluntarily implementing this recall of affected car charger cords x D Please note that the car charger cord defect has no impact on the functioning of the iGo x D Portable Oxygen Concentrator device or any other accessories | Iden fy any inventory in your possession using the following iden fiers: a. model DV6X-619 Rev E car charger cords of lot numbers 2023-01-01 through 2024-12-31 (see Figure 1 below). Not all Rev E cords are recalled. Refer to step 1c below to confirm if the Rev E unit is recalled. b. If the cord that was originally included with the iGo®2 POC unit is s ll with the POC, the affected cords can also be iden fied by the serial number on the iGo®2 POC unit (see Figure 2 below). The range of applicable serial numbers is provided in Table 1 on page 5. c. If there is any doubt, or if you do not maintain the car charger cord with the original POC unit as part of your rental distribu on process (i.e., used cords are stored and comingled in a box), inspec on and confirma on of the condi on of the car charger cord adapter can be easily performed by unscrewing the metal cap. (See Figure 3 for spring and fuse posi on.) i. If it reveals the internal spring next to the cap, this configura on is being recalled. ii. If it reveals the fuse next to the cap, this configura on is not recalled. Note: For every product confirmed to be unaffected by this recall, we recommend that you mark the label as acceptable in some permanent manner to ensure future inspec ons are unnecessary a er rental redistribu on and return. 2. Destroy all affected car charger cords by cu ng the cord with wire-cu ers or unscrewing the cap and then discarding the car charger parts. 3. To obtain replacement car charger cords, click on the Verifica on Code hyperlink, located on the first page of this no ce directly below the address block, or go to www.recallrtr.com/iGo2DME and input the Verifica on Code. Using that code on the website, you will be able to: a. Please confirm receipt of this recall and confirm the quan ty of affected car charger cords in your possession (see Table 1 on page 5 for the full range of affected lots and serial numbers) b. Order replacement products; and c. Fill out a web form to acknowledge receipt of and compliance with this recall. If you experience any issues with the website or wish to communicate via telephone, Drive DeVilbiss Healthcare has established a dedicated Customer Service phone number, (888-770-7017), which will be available Monday through Friday, from 7:30 a.m. to 5:00 p.m. Central Time. You can also contact us via email regarding this recall at iGo2DME@real meresults.net 4. Please contact all customer/end users who purchased or rented an affected unit and provide them no ce of and instruc ons for compliance with this recall. A dra no fica on to end-users is a ached for your convenience. The recall website will also be enabled to assist end-users with registering for the recall and receiving a replacement of affected car charger cords. If you would like assistance no fying your customers, please contact Dan Hinkebein from Real me Results at dan.hinkebein@real meresults.net to provide your customer list and sign any BAA so that no fica on of end-users may be performed on your behalf. 5. As these items are typically rented to end users, we recommend that you con nue to check units as they return from rental use. 6. Un l such me as you obtain replacement car charger cords, we recommend you consider the following: a. If you do not have inventory with unaffected car chargers, distribute the iGo®2 POC system without the car charger as this is a non-cri cal accessory. b. Advise end users that if they must use the car charger before it has been inspected and/or replaced, they should ensure the ba ery is fully charged and/or flow se ng is at or below 3 lpm. | 2025-05-20 17:35
DeVilbiss Healthcare LLC | iGo2 | 155 | - | The iGo POC car charger cord is a non critical accessory to the iGo POC It allows the user x D to operate their iGo POC from a DC power supply such as the charging port in a vehicle x D Under normal operating conditions the car charger cord will power the unit while concurrently x D charging the device s internal battery x D After receiving complaints related to overheating of the plug portion of the car charger cord an x D investigation determined the root cause was an assembly error specifically the fuse and the x D retention spring were installed in the reverse positions It was confirmed that during use the x D adapter connection may become hot to the touch and in some cases cause deformation of the x D plastic housing The manufacturing assembly error has now been corrected x D Drive Devilbiss Healthcare is voluntarily implementing this recall of affected car charger cords x D Please note that the car charger cord defect has no impact on the functioning of the iGo x D Portable Oxygen Concentrator device or any other accessories | Iden fy any inventory in your possession using the following iden fiers: a. model DV6X-619 Rev E car charger cords of lot numbers 2023-01-01 through 2024-12-31 (see Figure 1 below). Not all Rev E cords are recalled. Refer to step 1c below to confirm if the Rev E unit is recalled. b. If the cord that was originally included with the iGo®2 POC unit is s ll with the POC, the affected cords can also be iden fied by the serial number on the iGo®2 POC unit (see Figure 2 below). The range of applicable serial numbers is provided in Table 1 on page 5. c. If there is any doubt, or if you do not maintain the car charger cord with the original POC unit as part of your rental distribu on process (i.e., used cords are stored and comingled in a box), inspec on and confirma on of the condi on of the car charger cord adapter can be easily performed by unscrewing the metal cap. (See Figure 3 for spring and fuse posi on.) i. If it reveals the internal spring next to the cap, this configura on is being recalled. ii. If it reveals the fuse next to the cap, this configura on is not recalled. Note: For every product confirmed to be unaffected by this recall, we recommend that you mark the label as acceptable in some permanent manner to ensure future inspec ons are unnecessary a er rental redistribu on and return. 2. Destroy all affected car charger cords by cu ng the cord with wire-cu ers or unscrewing the cap and then discarding the car charger parts. 3. To obtain replacement car charger cords, click on the Verifica on Code hyperlink, located on the first page of this no ce directly below the address block, or go to www.recallrtr.com/iGo2DME and input the Verifica on Code. Using that code on the website, you will be able to: a. Please confirm receipt of this recall and confirm the quan ty of affected car charger cords in your possession (see Table 1 on page 5 for the full range of affected lots and serial numbers) b. Order replacement products; and c. Fill out a web form to acknowledge receipt of and compliance with this recall. If you experience any issues with the website or wish to communicate via telephone, Drive DeVilbiss Healthcare has established a dedicated Customer Service phone number, (888-770-7017), which will be available Monday through Friday, from 7:30 a.m. to 5:00 p.m. Central Time. You can also contact us via email regarding this recall at iGo2DME@real meresults.net 4. Please contact all customer/end users who purchased or rented an affected unit and provide them no ce of and instruc ons for compliance with this recall. A dra no fica on to end-users is a ached for your convenience. The recall website will also be enabled to assist end-users with registering for the recall and receiving a replacement of affected car charger cords. If you would like assistance no fying your customers, please contact Dan Hinkebein from Real me Results at dan.hinkebein@real meresults.net to provide your customer list and sign any BAA so that no fica on of end-users may be performed on your behalf. 5. As these items are typically rented to end users, we recommend that you con nue to check units as they return from rental use. 6. Un l such me as you obtain replacement car charger cords, we recommend you consider the following: a. If you do not have inventory with unaffected car chargers, distribute the iGo®2 POC system without the car charger as this is a non-cri cal accessory. b. Advise end users that if they must use the car charger before it has been inspected and/or replaced, they should ensure the ba ery is fully charged and/or flow se ng is at or below 3 lpm. | 2025-05-20 17:35
DeVilbiss Healthcare LLC | iGo2 | 155 | - | The iGo POC car charger cord is a non critical accessory to the iGo POC It allows the user x D to operate their iGo POC from a DC power supply such as the charging port in a vehicle x D Under normal operating conditions the car charger cord will power the unit while concurrently x D charging the device s internal battery x D After receiving complaints related to overheating of the plug portion of the car charger cord an x D investigation determined the root cause was an assembly error specifically the fuse and the x D retention spring were installed in the reverse positions It was confirmed that during use the x D adapter connection may become hot to the touch and in some cases cause deformation of the x D plastic housing The manufacturing assembly error has now been corrected x D Drive Devilbiss Healthcare is voluntarily implementing this recall of affected car charger cords x D Please note that the car charger cord defect has no impact on the functioning of the iGo x D Portable Oxygen Concentrator device or any other accessories | Iden fy any inventory in your possession using the following iden fiers: a. model DV6X-619 Rev E car charger cords of lot numbers 2023-01-01 through 2024-12-31 (see Figure 1 below). Not all Rev E cords are recalled. Refer to step 1c below to confirm if the Rev E unit is recalled. b. If the cord that was originally included with the iGo®2 POC unit is s ll with the POC, the affected cords can also be iden fied by the serial number on the iGo®2 POC unit (see Figure 2 below). The range of applicable serial numbers is provided in Table 1 on page 5. c. If there is any doubt, or if you do not maintain the car charger cord with the original POC unit as part of your rental distribu on process (i.e., used cords are stored and comingled in a box), inspec on and confirma on of the condi on of the car charger cord adapter can be easily performed by unscrewing the metal cap. (See Figure 3 for spring and fuse posi on.) i. If it reveals the internal spring next to the cap, this configura on is being recalled. ii. If it reveals the fuse next to the cap, this configura on is not recalled. Note: For every product confirmed to be unaffected by this recall, we recommend that you mark the label as acceptable in some permanent manner to ensure future inspec ons are unnecessary a er rental redistribu on and return. 2. Destroy all affected car charger cords by cu ng the cord with wire-cu ers or unscrewing the cap and then discarding the car charger parts. 3. To obtain replacement car charger cords, click on the Verifica on Code hyperlink, located on the first page of this no ce directly below the address block, or go to www.recallrtr.com/iGo2DME and input the Verifica on Code. Using that code on the website, you will be able to: a. Please confirm receipt of this recall and confirm the quan ty of affected car charger cords in your possession (see Table 1 on page 5 for the full range of affected lots and serial numbers) b. Order replacement products; and c. Fill out a web form to acknowledge receipt of and compliance with this recall. If you experience any issues with the website or wish to communicate via telephone, Drive DeVilbiss Healthcare has established a dedicated Customer Service phone number, (888-770-7017), which will be available Monday through Friday, from 7:30 a.m. to 5:00 p.m. Central Time. You can also contact us via email regarding this recall at iGo2DME@real meresults.net 4. Please contact all customer/end users who purchased or rented an affected unit and provide them no ce of and instruc ons for compliance with this recall. A dra no fica on to end-users is a ached for your convenience. The recall website will also be enabled to assist end-users with registering for the recall and receiving a replacement of affected car charger cords. If you would like assistance no fying your customers, please contact Dan Hinkebein from Real me Results at dan.hinkebein@real meresults.net to provide your customer list and sign any BAA so that no fica on of end-users may be performed on your behalf. 5. As these items are typically rented to end users, we recommend that you con nue to check units as they return from rental use. 6. Un l such me as you obtain replacement car charger cords, we recommend you consider the following: a. If you do not have inventory with unaffected car chargers, distribute the iGo®2 POC system without the car charger as this is a non-cri cal accessory. b. Advise end users that if they must use the car charger before it has been inspected and/or replaced, they should ensure the ba ery is fully charged and/or flow se ng is at or below 3 lpm. | 2025-05-20 17:35
Randox Laboratories Ltd | Liquid Assayed Chemistry Control Premium Plus Level 2 | 2 | - | Randox can confirm there has been a transcription error in the Instructions for Use IFU for x D Liquid Assayed Chemistry Control Premium Plus Level LAN Lot UN The targets x D and ranges for the analytes Caeruloplasmin and Paracetamol have been listed incorrectly on x D the IFU Please refer to Table for the correct targets and ranges | Action to be taken: Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Download corrected IFU from Smartdocs.com and discard of previous copies. Please accept our apologies for any inconvenience caused. Thank you for your patience and understanding. If you have any questions or concerns, please contact Randox Technical Services. | 2025-05-22 10:37
Randox Laboratories Ltd | Liquid Assayed Chemistry Control Premium Plus Level 3 | 2 | - | Randox can confirm there has been a transcription error in the Instructions for Use IFU for x D Liquid Assayed Chemistry Control Premium Plus Level LAN Lot UN The targets x D and ranges for the analytes Caeruloplasmin and Paracetamol have been listed incorrectly on x D the IFU Please refer to Table for the correct targets and ranges | Action to be taken: Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Download corrected IFU from Smartdocs.com and discard of previous copies. Please accept our apologies for any inconvenience caused. Thank you for your patience and understanding. If you have any questions or concerns, please contact Randox Technical Services. | 2025-05-22 10:37
SCHILLER Medical SAS | MAGLIFE RT-1 | 11 | - | PLEASE FIND THE ATTACHED FORM | PLEASE FIND THE ATTACHED FORM | 2025-05-22 14:28
Philips | Trilogy Evo | 337 | - | Pressure delivery may incrementally exceed the specified limits cmH O for pressure control accuracy when using the following circuit configurations with preset circuit settings configured on the device User Interface x D F P Adult RT single limb inspiratory heated mm ID circuit configured with the Adult mm circuit setting x D F P Adult MR reusable single limb Evatherm circuit configured with the Adult Pediatric mm circuit setting x D All other critical parameters including Tidal Volume measurements and alarms continue to perform as expected and are not impacted by the circuit setting x D This situation may happen when the clinician selects a preset circuit configuration based on the diameter presented in the labelling provided by the circuit manufacturer Philips Respironics testing indicates that any potential deviation will not deviate by more than cmH O from the published accuracy specification for pressure control All specifications are published in the device Instructions for Use | Philips Respironics is currently working on a solution to resolve this issue. A separate notification will be provided when a solution is available. | 2025-05-22 19:01
Philips | Trilogy Evo O2 | 337 | - | Pressure delivery may incrementally exceed the specified limits cmH O for pressure control accuracy when using the following circuit configurations with preset circuit settings configured on the device User Interface x D F P Adult RT single limb inspiratory heated mm ID circuit configured with the Adult mm circuit setting x D F P Adult MR reusable single limb Evatherm circuit configured with the Adult Pediatric mm circuit setting x D All other critical parameters including Tidal Volume measurements and alarms continue to perform as expected and are not impacted by the circuit setting x D This situation may happen when the clinician selects a preset circuit configuration based on the diameter presented in the labelling provided by the circuit manufacturer Philips Respironics testing indicates that any potential deviation will not deviate by more than cmH O from the published accuracy specification for pressure control All specifications are published in the device Instructions for Use | Philips Respironics is currently working on a solution to resolve this issue. A separate notification will be provided when a solution is available. | 2025-05-22 19:01
Philips | Trilogy EV300 | 337 | - | Pressure delivery may incrementally exceed the specified limits cmH O for pressure control accuracy when using the following circuit configurations with preset circuit settings configured on the device User Interface x D F P Adult RT single limb inspiratory heated mm ID circuit configured with the Adult mm circuit setting x D F P Adult MR reusable single limb Evatherm circuit configured with the Adult Pediatric mm circuit setting x D All other critical parameters including Tidal Volume measurements and alarms continue to perform as expected and are not impacted by the circuit setting x D This situation may happen when the clinician selects a preset circuit configuration based on the diameter presented in the labelling provided by the circuit manufacturer Philips Respironics testing indicates that any potential deviation will not deviate by more than cmH O from the published accuracy specification for pressure control All specifications are published in the device Instructions for Use | Philips Respironics is currently working on a solution to resolve this issue. A separate notification will be provided when a solution is available. | 2025-05-22 19:01
Philips Medical Systems Nederland B.V. | Panorama HFO | 4 | - | Philips has been informed of an event in which the structural integrity of the Panorama T HFO system components failed due to unintended excessive pressure buildup of helium gas during a magnet quench x D During a quench a large amount of helium evaporates and is vented outside the building through a venting system If an unknown blockage is present in the venting system and the pressure exceeds design limits the structural integrity of the system may be compromised The magnet may experience an unsolicited quench during normal use or will quench when initiated in an emergency situation by the operator pressing the Magnet EMERGENCY STOP button x D Philips has received one complaint of system and property damage after quenching the magnet on Panorama T systems as of November There was no report of injury or serious harm | Philips is providing this Urgent Field Safety Notice Letter which contains recommendations for continued use of the systems referenced in Section 4. The advice provided in Section 4 is applicable for systems in operational status and is intended to address the issue identified in this letter. It is not intended to extend the service life of the product or designed to address any of the risks related to continued use of the system after its service life (meaning the timespan during which the system is expected to perform and remain safe when used in accordance with its labelling). | 2025-05-23 09:07
TECHNOLOGIE MEDICALE | TUBE 10 FILTERS PLASTIQUW RV2/RV3. | - | - | Untheorized Medical Device and Counterfeit | • Recall affected devices • Submit a Field Safety Corrective Action through the NCMDR website. https://ncmdr.sfda.gov.sa/ within 3 working days. Publish date: 8/3/2023 Close date: 9/15/2023 | 2025-05-25 08:06
CareDx Pty Ltd. | AlloSeq cfDNA | 4 | - | The root cause has been identified to be linked to the raw material used in the AlloSeq cfDNA PCR Mix components of this lot The issue has been identified to be isolated to this lot based on ingoing components All ingoing components have since been replaced hence the issue is isolated to this lot | Remove AlloSeq cfDNA lot 24-300 from the market to prevent further incidents, and discarding of the faulty raw material to prevent reuse. | 2025-05-25 14:33
Covidien llc | Newport HT Ventilator with Accessories, Plus model; Newport HT70™ Plus Ventilator with oxygen monitoring, SpO2 monitoring, on­airway flow sensor; | 46 | - | Capacitor Loss of Ventilation x D Current post market surveillance data indicates an ongoing signal related to the C capacitor located on the main control PCBA of HT Ventilators C is a bypass capacitor to a power management chip U mounted on the control board PCBA If the C capacitor fails it will disable the output of the power management chip U causing the ventilator to cease operation and stop ventilation An audible alarm piezo will be initiated and the safety valve will open Investigation indicates that although the C capacitor is operating beyond its nominal specifications it is durable and failure as a result of over voltage is unlikely and not considered the primary cause of failure per the manufacturer s testing and published white paper Statistical analysis of the field failures has shown that the highest failure rates are confined to PCBAs manufactured with C capacitors from only two lot codes and x D Capacitor Shutdown Alarm x D Complaint trending review identified three customer complaints that reported failures of the Shutdown Buzzer Alarm Test as the alarm hardly sounded i e the alarm was of insufficient duration or volume per specification x D Ventilators with Control Printed Circuit Board Assemblies PCBAs manufactured since Nov containing C capacitors are at higher risk x D | • Review attachments A and B to determine which of your ventilators or service parts may be affected. • If an affected ventilator or service part is currently in clinical use, remove the device from service and replace it with an alternate means of ventilation. • After completing the transition to the alternate ventilator(s), quarantine all affected ventilators and service parts identified in attachment A and B and follow your facility’s processes and procedures for removal and disposal of the impacted devices. Complete the Customer Acknowledgement Form and return to Medtronic. • Contact your local Medtronic representative for any additional questions or concerns. • Distribute this notice to all relevant personnel within your organization and any organizations where these ventilators or service parts may have been transferred or distributed. • Provide the enclosed patient letter to any of your patients using this ventilator in the home care setting to inform them of this recall. • If you are aware of any incidents related to these issues, please contact your local Medtronic Representative. | 2025-05-26 11:12
Draeger Medical, Inc. | Filter SafeStar | 204600 | - | If a therapeutic decision is made solely on the basis of a slowly rising CO curve as described above Fig unnecessary medication might lead to a serious deterioration in the state of health of certain patient groups e g with pre existing cardiac conditions The risk assessement revealed an unacceptable risk x D The FSCA is intended to mitigate the risk | All involved users will be informed within the frame of a Field Safety Corrective Action (FSCA) by means of a Field Safety Notice (FSN) about the potential risk. | 2025-05-27 15:02
Draeger Medical, Inc. | Filter/HME TwinStar | 204600 | - | If a therapeutic decision is made solely on the basis of a slowly rising CO curve as described above Fig unnecessary medication might lead to a serious deterioration in the state of health of certain patient groups e g with pre existing cardiac conditions The risk assessement revealed an unacceptable risk x D The FSCA is intended to mitigate the risk | All involved users will be informed within the frame of a Field Safety Corrective Action (FSCA) by means of a Field Safety Notice (FSN) about the potential risk. | 2025-05-27 15:02
Medtron AG | All consumables from MEDTRON AG | 72103 | - | Currently the consumables from MEDTRON AG do not provide evidence that they fulfill the requirements for ethylene oxide residuals according to DIN EN ISO Biological evaluation of medical devices Part Ethylene oxide sterilization residuals for use in special patient populations in particular children infants neonates and premature infants x D There might be a health risk for patients kg due to increased ethylene oxide residuals x D Thus the use of the consumables is restricted to patients kg x D | Please make sure that all users of the above-mentioned products in your facility and other persons to be informed take note of this "Field Safety Notice". Please do not use the products anymore for patients < 10 kg. If the products have already been used no further measures are required. The products can still be used for patients ≥ 10 kg and no further measures are required. | 2025-05-27 15:40
Dexcom Inc. | Dexcom One+ iOS CGM App | 50 | - | A bug was identified within app version that can cause the app to enter a state where it stops x D updating the estimated glucose values EGVs which could result in alerts not triggering x D Additional indicators of this bug may include the trend graph progressing with no new EGVs x D added to the graph plotted points disappearing from the graph or the trend graph disappearing x D completely You are using an older version of the Dexcom One iOS CGM app which may be x D prone to this issue and thereby not function as intended | Upgrading from app version 1.2 is mandatory and you will not be able to use app version 1.2 after [Month, Day, Year]. Update your app now to continue using the app: 1. Tap Update App to go to the app store 2. Install the latest Dexcom One+ app version 3. Open the Dexcom One+ app If you need assistance, contact Technical Support. Local Tech Support contact information is available at Dexcom.com | 2025-05-27 16:49
Integra Lifesciences | Extended Tip Applicator, 8CM, Box of 5 | 3 | - | The decision to conduct a voluntary recall of the product was based on an internal Integra investigation x D that identified products could potentially contain out of specification levels of endotoxin Additionally x D the full investigation demonstrated an incomplete bioburden assessment and an incomplete sterilization x D location transfer documentation to ensure the effectiveness of the sterilization process x D Changes have been made to the manufacturing process and a verification of the effectiveness is being x D conducted to remediate the out of specification endotoxin Comprehensive bioburden monitoring and x D further validation of the applicator sterilization process will also be performed prior to the release of this x D product back to the market | Actions to be taken by Distributors: 1. Please review and understand the information provided in this letter. 2. Determine if the product you have is subject to the recall: a. Identify the impacted reference and lot number. b. See Appendix 2 below for a sample of product label for where to locate the reference and lot number. 3. If you do have affected product(s) in your warehouse: a. Quarantine them immediately. b. Check the box “I do have affected unit(s)” in the enclosed reply form. c. Record on the form the total quantity of affected unit(s) and lot number(s) that you have. 4. If you do not have affected product(s) in your warehouse, check the box, “I do not have affected unit(s)”. 5. Please check your customer traceability records for shipments of affected products. 6. If you have shipped impacted products to your customers, please complete below: a. Create a customer reply form with your contact details. b. Forward a copy of the Field Safety Notice to any of your customers that have purchased the affected products and lot numbers. c. Collect completed response forms and affected product(s) from your customers and indicate the total quantities and lot(s) in the distributor reply form (Appendix 1). 7. Please return the completed Reply form by email to emea-fsca@integralife.com, By filling in this form, you confirm that you have received this Safety Notice and you intend to fully comply with this notification. We expect a response within 3 weeks. You also confirm that this notification has been forwarded to every person concerned in your organization. 8. At receipt of the reply form, and if it is noted that you or your customers have affected products available for return, Customer Service will contact you and provide an RMA number and directions to return the product(s). The credit will be processed upon receipt and verification of returned goods. 9. We recommend that you retain a copy of the form for your records. PLEASE NOTE THAT REGARDLESS OF WHETHER YOU HAVE THE AFFECTED PRODUCT TO RETURN OR NOT – A COMPLETED ACKNOWLEDGEMENT IS REQUIRED The receipt of this form ensures that Integra has achieved a level of effectiveness in communicating this information. National Competent Authorities may perform audits of field actions of this nature to verify that our customers have been notified and understand the nature of the field action being taken. If required by the national medical device regulation of your distribution area, please ensure this Field Safety Corrective Action is notified to the national competent authorities. Thank you for your cooperation with this Field Safety Corrective Action and for returning the attached Reply Form. Please feel free to contact our Post Market Surveillance Department at emea-fsca@integralife.com for any additional questions. Your cooperation is appreciated, and we thank you for your continued collaboration. | 2025-05-28 11:53
Integra Lifesciences | Extended Tip Applicator, 15 CM, Box of 5 | 3 | - | The decision to conduct a voluntary recall of the product was based on an internal Integra investigation x D that identified products could potentially contain out of specification levels of endotoxin Additionally x D the full investigation demonstrated an incomplete bioburden assessment and an incomplete sterilization x D location transfer documentation to ensure the effectiveness of the sterilization process x D Changes have been made to the manufacturing process and a verification of the effectiveness is being x D conducted to remediate the out of specification endotoxin Comprehensive bioburden monitoring and x D further validation of the applicator sterilization process will also be performed prior to the release of this x D product back to the market | Actions to be taken by Distributors: 1. Please review and understand the information provided in this letter. 2. Determine if the product you have is subject to the recall: a. Identify the impacted reference and lot number. b. See Appendix 2 below for a sample of product label for where to locate the reference and lot number. 3. If you do have affected product(s) in your warehouse: a. Quarantine them immediately. b. Check the box “I do have affected unit(s)” in the enclosed reply form. c. Record on the form the total quantity of affected unit(s) and lot number(s) that you have. 4. If you do not have affected product(s) in your warehouse, check the box, “I do not have affected unit(s)”. 5. Please check your customer traceability records for shipments of affected products. 6. If you have shipped impacted products to your customers, please complete below: a. Create a customer reply form with your contact details. b. Forward a copy of the Field Safety Notice to any of your customers that have purchased the affected products and lot numbers. c. Collect completed response forms and affected product(s) from your customers and indicate the total quantities and lot(s) in the distributor reply form (Appendix 1). 7. Please return the completed Reply form by email to emea-fsca@integralife.com, By filling in this form, you confirm that you have received this Safety Notice and you intend to fully comply with this notification. We expect a response within 3 weeks. You also confirm that this notification has been forwarded to every person concerned in your organization. 8. At receipt of the reply form, and if it is noted that you or your customers have affected products available for return, Customer Service will contact you and provide an RMA number and directions to return the product(s). The credit will be processed upon receipt and verification of returned goods. 9. We recommend that you retain a copy of the form for your records. PLEASE NOTE THAT REGARDLESS OF WHETHER YOU HAVE THE AFFECTED PRODUCT TO RETURN OR NOT – A COMPLETED ACKNOWLEDGEMENT IS REQUIRED The receipt of this form ensures that Integra has achieved a level of effectiveness in communicating this information. National Competent Authorities may perform audits of field actions of this nature to verify that our customers have been notified and understand the nature of the field action being taken. If required by the national medical device regulation of your distribution area, please ensure this Field Safety Corrective Action is notified to the national competent authorities. Thank you for your cooperation with this Field Safety Corrective Action and for returning the attached Reply Form. Please feel free to contact our Post Market Surveillance Department at emea-fsca@integralife.com for any additional questions. Your cooperation is appreciated, and we thank you for your continued collaboration. | 2025-05-28 11:53
Ortho Clinical Diagnostics Inc. | VITROS 4600 Chemistry System | 131 | - | The VITROS System software is designed to apply a Reagent Expired RE code to results x D if the Use Expired Reagents option has been enabled and an expired reagent is used If x D Use Expired Reagents has not been enabled the VITROS System software is expected to x D report a No Result and apply an Insufficient Inventory II code when attempting to x D process samples using expired reagents including diluents x D QuidelOrtho received a customer complaint stating that their VITROS System had applied x D an RE code to results despite the Use Expired Reagents option not enabled on their VITROS x D System x D We investigated the complaint and determined that a software anomaly exists in the x D VITROS System software affecting all available versions Despite the Use Expired x D Reagents option not enabled if a VITROS Diluent Pack Lot is manually loaded on VITROS x D and XT Systems with the shelf expiration date left blank the VITROS x D System will flag the pack as expired but report results applying an RE code to the results x D Please note that the anomaly will allow the VITROS System to report results using VITROS x D Diluent Packs which have not reached the shelf expiration date AND expired VITROS Diluent x D Packs as printed on the product labeling if the VITROS Diluent Pack was manually loaded x D with the shelf expiration date left blank | • To prevent this issue from occurring, the shelf expiration date must be entered when manually loading VITROS Diluent Packs on VITROS 4600/5600/XT 7600 Systems. • View the Reagent Management Supply 3 screen to check for any onboard VITROS Diluent Packs marked as expired. o For all VITROS Diluent Packs marked as expired, navigate to View By Reagent → View Dil/ANC and confirm the shelf expiration date matches the shelf expiration date on the product labeling. o If the shelf expiration dates do not match, discard the VITROS Diluent Pack and load a new VITROS Diluent Pack. QuidelOrtho will credit your discarded inventory. (Credit will only be issued for within-expiry products.) • Complete and return the enclosed Confirmation of Receipt form no later than Month DD, 2025. Upon receipt of your completed Confirmation of Receipt form, QuidelOrtho will credit or replace your discarded inventory. • Save this notification with your User Documentation or post this notification by each VITROS 4600/5600/XT 7600 System in your laboratory until the issue has been resolved. • Please forward this notification if the affected product was distributed outside of your facility. • If your laboratory has experienced the issue with this product and you have not already done so, please report the occurrence to your local Global Services Organization. | 2025-05-28 17:29
Ortho Clinical Diagnostics Inc. | VITROS 5600 Integrated System | 131 | - | The VITROS System software is designed to apply a Reagent Expired RE code to results x D if the Use Expired Reagents option has been enabled and an expired reagent is used If x D Use Expired Reagents has not been enabled the VITROS System software is expected to x D report a No Result and apply an Insufficient Inventory II code when attempting to x D process samples using expired reagents including diluents x D QuidelOrtho received a customer complaint stating that their VITROS System had applied x D an RE code to results despite the Use Expired Reagents option not enabled on their VITROS x D System x D We investigated the complaint and determined that a software anomaly exists in the x D VITROS System software affecting all available versions Despite the Use Expired x D Reagents option not enabled if a VITROS Diluent Pack Lot is manually loaded on VITROS x D and XT Systems with the shelf expiration date left blank the VITROS x D System will flag the pack as expired but report results applying an RE code to the results x D Please note that the anomaly will allow the VITROS System to report results using VITROS x D Diluent Packs which have not reached the shelf expiration date AND expired VITROS Diluent x D Packs as printed on the product labeling if the VITROS Diluent Pack was manually loaded x D with the shelf expiration date left blank | • To prevent this issue from occurring, the shelf expiration date must be entered when manually loading VITROS Diluent Packs on VITROS 4600/5600/XT 7600 Systems. • View the Reagent Management Supply 3 screen to check for any onboard VITROS Diluent Packs marked as expired. o For all VITROS Diluent Packs marked as expired, navigate to View By Reagent → View Dil/ANC and confirm the shelf expiration date matches the shelf expiration date on the product labeling. o If the shelf expiration dates do not match, discard the VITROS Diluent Pack and load a new VITROS Diluent Pack. QuidelOrtho will credit your discarded inventory. (Credit will only be issued for within-expiry products.) • Complete and return the enclosed Confirmation of Receipt form no later than Month DD, 2025. Upon receipt of your completed Confirmation of Receipt form, QuidelOrtho will credit or replace your discarded inventory. • Save this notification with your User Documentation or post this notification by each VITROS 4600/5600/XT 7600 System in your laboratory until the issue has been resolved. • Please forward this notification if the affected product was distributed outside of your facility. • If your laboratory has experienced the issue with this product and you have not already done so, please report the occurrence to your local Global Services Organization. | 2025-05-28 17:29
Ortho Clinical Diagnostics Inc. | VITROS XT 7600 Integrated System | 131 | - | The VITROS System software is designed to apply a Reagent Expired RE code to results x D if the Use Expired Reagents option has been enabled and an expired reagent is used If x D Use Expired Reagents has not been enabled the VITROS System software is expected to x D report a No Result and apply an Insufficient Inventory II code when attempting to x D process samples using expired reagents including diluents x D QuidelOrtho received a customer complaint stating that their VITROS System had applied x D an RE code to results despite the Use Expired Reagents option not enabled on their VITROS x D System x D We investigated the complaint and determined that a software anomaly exists in the x D VITROS System software affecting all available versions Despite the Use Expired x D Reagents option not enabled if a VITROS Diluent Pack Lot is manually loaded on VITROS x D and XT Systems with the shelf expiration date left blank the VITROS x D System will flag the pack as expired but report results applying an RE code to the results x D Please note that the anomaly will allow the VITROS System to report results using VITROS x D Diluent Packs which have not reached the shelf expiration date AND expired VITROS Diluent x D Packs as printed on the product labeling if the VITROS Diluent Pack was manually loaded x D with the shelf expiration date left blank | • To prevent this issue from occurring, the shelf expiration date must be entered when manually loading VITROS Diluent Packs on VITROS 4600/5600/XT 7600 Systems. • View the Reagent Management Supply 3 screen to check for any onboard VITROS Diluent Packs marked as expired. o For all VITROS Diluent Packs marked as expired, navigate to View By Reagent → View Dil/ANC and confirm the shelf expiration date matches the shelf expiration date on the product labeling. o If the shelf expiration dates do not match, discard the VITROS Diluent Pack and load a new VITROS Diluent Pack. QuidelOrtho will credit your discarded inventory. (Credit will only be issued for within-expiry products.) • Complete and return the enclosed Confirmation of Receipt form no later than Month DD, 2025. Upon receipt of your completed Confirmation of Receipt form, QuidelOrtho will credit or replace your discarded inventory. • Save this notification with your User Documentation or post this notification by each VITROS 4600/5600/XT 7600 System in your laboratory until the issue has been resolved. • Please forward this notification if the affected product was distributed outside of your facility. • If your laboratory has experienced the issue with this product and you have not already done so, please report the occurrence to your local Global Services Organization. | 2025-05-28 17:29
Boston Scientific Corporation | ACURATE neo2™ and ACURATE Prime™ Aortic Valve Systems | 697 | - | Boston Scientific is discontinuing global commercialization of the ACURATE neo and ACURATE Prime Aortic Valve Systems and is voluntarily implementing a removal of any unused inventory of these devices While data continue to support the performance of the ACURATE valve system when the product s optimized instructions for use are followed this decision was made due to the increased clinical and regulatory requirements to maintain regulatory approvals in global markets The investments and additional resources needed to satisfy these requirements are prohibitive for the company | Discontinuing global commercialization of the ACURATE neo2™ and ACURATE Prime™ Aortic Valve Systems and is voluntarily implementing a removal of any unused inventory of these devices. | 2025-05-30 14:07
Philips Medical Systems Nederland B.V. | AZURION 3 M12 | 85 | - | Philips has identified a software issue in the internal communication process between the system software and the X ray generator firmware This issue may be triggered under either of the following situations x D Situation Pedal Tap When the foot switch is pressed and released rapidly without generating X rays due to a short duration of contact the system may enter a state where X ray generation is inhibited When this occurs the system does not show any message to the user indicating that X ray is not possible A warm restart may resolve the situation in some cases A cold restart will always resolve the situation x D Situation Phase Fault Azurion systems are designed to automatically recover from phase faults a phase fault is an abnormal condition in which one or more phase voltages drop to near zero When a phase fault occurs the X ray generator software initiates an automatic recovery process to restore fully system functionality during which imaging X ray functionality is not available for approximately seconds When the identified software issue occurs during this recovery communication with the system software is lost resulting in the system not recognizing that the generator has recovered and is ready for operation In this situation the system shows the user message Generator is busy starting up no X ray possible If the phase fault is detected by the generator during acquisition the system also shows the message Run Aborted Tube Problem To restore system functionality a cold restart of the system is required | Philips will address the identified issue by implementing a software update in all affected systems. For systems featuring an Interventional Workspot and/or EchoNavigator, Philips will upgrade the Interventional Workspot and/or EchoNavigator software version to ensure compatibility with the updated Philips Azurion system software. Philips expects to have the updated software released by June 2025. Your local Philips representative will contact you to schedule a visit to install the software update once available. | 2025-05-30 14:39
Philips Medical Systems Nederland B.V. | AZURION 3 M15 | 85 | - | Philips has identified a software issue in the internal communication process between the system software and the X ray generator firmware This issue may be triggered under either of the following situations x D Situation Pedal Tap When the foot switch is pressed and released rapidly without generating X rays due to a short duration of contact the system may enter a state where X ray generation is inhibited When this occurs the system does not show any message to the user indicating that X ray is not possible A warm restart may resolve the situation in some cases A cold restart will always resolve the situation x D Situation Phase Fault Azurion systems are designed to automatically recover from phase faults a phase fault is an abnormal condition in which one or more phase voltages drop to near zero When a phase fault occurs the X ray generator software initiates an automatic recovery process to restore fully system functionality during which imaging X ray functionality is not available for approximately seconds When the identified software issue occurs during this recovery communication with the system software is lost resulting in the system not recognizing that the generator has recovered and is ready for operation In this situation the system shows the user message Generator is busy starting up no X ray possible If the phase fault is detected by the generator during acquisition the system also shows the message Run Aborted Tube Problem To restore system functionality a cold restart of the system is required | Philips will address the identified issue by implementing a software update in all affected systems. For systems featuring an Interventional Workspot and/or EchoNavigator, Philips will upgrade the Interventional Workspot and/or EchoNavigator software version to ensure compatibility with the updated Philips Azurion system software. Philips expects to have the updated software released by June 2025. Your local Philips representative will contact you to schedule a visit to install the software update once available. | 2025-05-30 14:39
Philips | Azurion 5 M12 | 85 | - | Philips has identified a software issue in the internal communication process between the system software and the X ray generator firmware This issue may be triggered under either of the following situations x D Situation Pedal Tap When the foot switch is pressed and released rapidly without generating X rays due to a short duration of contact the system may enter a state where X ray generation is inhibited When this occurs the system does not show any message to the user indicating that X ray is not possible A warm restart may resolve the situation in some cases A cold restart will always resolve the situation x D Situation Phase Fault Azurion systems are designed to automatically recover from phase faults a phase fault is an abnormal condition in which one or more phase voltages drop to near zero When a phase fault occurs the X ray generator software initiates an automatic recovery process to restore fully system functionality during which imaging X ray functionality is not available for approximately seconds When the identified software issue occurs during this recovery communication with the system software is lost resulting in the system not recognizing that the generator has recovered and is ready for operation In this situation the system shows the user message Generator is busy starting up no X ray possible If the phase fault is detected by the generator during acquisition the system also shows the message Run Aborted Tube Problem To restore system functionality a cold restart of the system is required | Philips will address the identified issue by implementing a software update in all affected systems. For systems featuring an Interventional Workspot and/or EchoNavigator, Philips will upgrade the Interventional Workspot and/or EchoNavigator software version to ensure compatibility with the updated Philips Azurion system software. Philips expects to have the updated software released by June 2025. Your local Philips representative will contact you to schedule a visit to install the software update once available. | 2025-05-30 14:39
Philips | Azurion 5 M20 | 85 | - | Philips has identified a software issue in the internal communication process between the system software and the X ray generator firmware This issue may be triggered under either of the following situations x D Situation Pedal Tap When the foot switch is pressed and released rapidly without generating X rays due to a short duration of contact the system may enter a state where X ray generation is inhibited When this occurs the system does not show any message to the user indicating that X ray is not possible A warm restart may resolve the situation in some cases A cold restart will always resolve the situation x D Situation Phase Fault Azurion systems are designed to automatically recover from phase faults a phase fault is an abnormal condition in which one or more phase voltages drop to near zero When a phase fault occurs the X ray generator software initiates an automatic recovery process to restore fully system functionality during which imaging X ray functionality is not available for approximately seconds When the identified software issue occurs during this recovery communication with the system software is lost resulting in the system not recognizing that the generator has recovered and is ready for operation In this situation the system shows the user message Generator is busy starting up no X ray possible If the phase fault is detected by the generator during acquisition the system also shows the message Run Aborted Tube Problem To restore system functionality a cold restart of the system is required | Philips will address the identified issue by implementing a software update in all affected systems. For systems featuring an Interventional Workspot and/or EchoNavigator, Philips will upgrade the Interventional Workspot and/or EchoNavigator software version to ensure compatibility with the updated Philips Azurion system software. Philips expects to have the updated software released by June 2025. Your local Philips representative will contact you to schedule a visit to install the software update once available. | 2025-05-30 14:39
Philips | Azurion 7 B12/12 | 85 | - | Philips has identified a software issue in the internal communication process between the system software and the X ray generator firmware This issue may be triggered under either of the following situations x D Situation Pedal Tap When the foot switch is pressed and released rapidly without generating X rays due to a short duration of contact the system may enter a state where X ray generation is inhibited When this occurs the system does not show any message to the user indicating that X ray is not possible A warm restart may resolve the situation in some cases A cold restart will always resolve the situation x D Situation Phase Fault Azurion systems are designed to automatically recover from phase faults a phase fault is an abnormal condition in which one or more phase voltages drop to near zero When a phase fault occurs the X ray generator software initiates an automatic recovery process to restore fully system functionality during which imaging X ray functionality is not available for approximately seconds When the identified software issue occurs during this recovery communication with the system software is lost resulting in the system not recognizing that the generator has recovered and is ready for operation In this situation the system shows the user message Generator is busy starting up no X ray possible If the phase fault is detected by the generator during acquisition the system also shows the message Run Aborted Tube Problem To restore system functionality a cold restart of the system is required | Philips will address the identified issue by implementing a software update in all affected systems. For systems featuring an Interventional Workspot and/or EchoNavigator, Philips will upgrade the Interventional Workspot and/or EchoNavigator software version to ensure compatibility with the updated Philips Azurion system software. Philips expects to have the updated software released by June 2025. Your local Philips representative will contact you to schedule a visit to install the software update once available. | 2025-05-30 14:39
Philips | Azurion 7 B20/15 | 85 | - | Philips has identified a software issue in the internal communication process between the system software and the X ray generator firmware This issue may be triggered under either of the following situations x D Situation Pedal Tap When the foot switch is pressed and released rapidly without generating X rays due to a short duration of contact the system may enter a state where X ray generation is inhibited When this occurs the system does not show any message to the user indicating that X ray is not possible A warm restart may resolve the situation in some cases A cold restart will always resolve the situation x D Situation Phase Fault Azurion systems are designed to automatically recover from phase faults a phase fault is an abnormal condition in which one or more phase voltages drop to near zero When a phase fault occurs the X ray generator software initiates an automatic recovery process to restore fully system functionality during which imaging X ray functionality is not available for approximately seconds When the identified software issue occurs during this recovery communication with the system software is lost resulting in the system not recognizing that the generator has recovered and is ready for operation In this situation the system shows the user message Generator is busy starting up no X ray possible If the phase fault is detected by the generator during acquisition the system also shows the message Run Aborted Tube Problem To restore system functionality a cold restart of the system is required | Philips will address the identified issue by implementing a software update in all affected systems. For systems featuring an Interventional Workspot and/or EchoNavigator, Philips will upgrade the Interventional Workspot and/or EchoNavigator software version to ensure compatibility with the updated Philips Azurion system software. Philips expects to have the updated software released by June 2025. Your local Philips representative will contact you to schedule a visit to install the software update once available. | 2025-05-30 14:39
Philips Medical Systems Nederland B.V. | AZURION 7 M12 | 85 | - | Philips has identified a software issue in the internal communication process between the system software and the X ray generator firmware This issue may be triggered under either of the following situations x D Situation Pedal Tap When the foot switch is pressed and released rapidly without generating X rays due to a short duration of contact the system may enter a state where X ray generation is inhibited When this occurs the system does not show any message to the user indicating that X ray is not possible A warm restart may resolve the situation in some cases A cold restart will always resolve the situation x D Situation Phase Fault Azurion systems are designed to automatically recover from phase faults a phase fault is an abnormal condition in which one or more phase voltages drop to near zero When a phase fault occurs the X ray generator software initiates an automatic recovery process to restore fully system functionality during which imaging X ray functionality is not available for approximately seconds When the identified software issue occurs during this recovery communication with the system software is lost resulting in the system not recognizing that the generator has recovered and is ready for operation In this situation the system shows the user message Generator is busy starting up no X ray possible If the phase fault is detected by the generator during acquisition the system also shows the message Run Aborted Tube Problem To restore system functionality a cold restart of the system is required | Philips will address the identified issue by implementing a software update in all affected systems. For systems featuring an Interventional Workspot and/or EchoNavigator, Philips will upgrade the Interventional Workspot and/or EchoNavigator software version to ensure compatibility with the updated Philips Azurion system software. Philips expects to have the updated software released by June 2025. Your local Philips representative will contact you to schedule a visit to install the software update once available. | 2025-05-30 14:39
Philips Medical Systems Nederland B.V. | AZURION 7 M20 | 85 | - | Philips has identified a software issue in the internal communication process between the system software and the X ray generator firmware This issue may be triggered under either of the following situations x D Situation Pedal Tap When the foot switch is pressed and released rapidly without generating X rays due to a short duration of contact the system may enter a state where X ray generation is inhibited When this occurs the system does not show any message to the user indicating that X ray is not possible A warm restart may resolve the situation in some cases A cold restart will always resolve the situation x D Situation Phase Fault Azurion systems are designed to automatically recover from phase faults a phase fault is an abnormal condition in which one or more phase voltages drop to near zero When a phase fault occurs the X ray generator software initiates an automatic recovery process to restore fully system functionality during which imaging X ray functionality is not available for approximately seconds When the identified software issue occurs during this recovery communication with the system software is lost resulting in the system not recognizing that the generator has recovered and is ready for operation In this situation the system shows the user message Generator is busy starting up no X ray possible If the phase fault is detected by the generator during acquisition the system also shows the message Run Aborted Tube Problem To restore system functionality a cold restart of the system is required | Philips will address the identified issue by implementing a software update in all affected systems. For systems featuring an Interventional Workspot and/or EchoNavigator, Philips will upgrade the Interventional Workspot and/or EchoNavigator software version to ensure compatibility with the updated Philips Azurion system software. Philips expects to have the updated software released by June 2025. Your local Philips representative will contact you to schedule a visit to install the software update once available. | 2025-05-30 14:39
Abbott Ireland Diagnostics Division | Alinity i Total PSA Reagent Kit | 219 | - | Please find attached FSN letter | Please find attached FSN letter | 2025-06-01 15:49
Becton, Dickinson and Company BD Biosciences | BD FACS 7-Color Setup Beads | 39 | - | Based on customer feedback BD has identified errors in the spectral overlap factors SOF table provided for specific lots of BD FACS Color Setup Beads as listed in the table above This issue may affect the accuracy of flow cytometry data due to incorrect flow cytometry setup compensation values for multicolor reagents for which proper fluorescence compensation is required prior to analyzing samples | Clinical User Actions 1. Immediately inspect your inventory for the specific catalog and lot numbers listed. It is recommended the clinical laboratory staff review and segregate the affected lots. 2. Remove and discard the affected label insert with incorrect spectral overlap factor values included in the kit box prior to use. 3. The wrong SOF (Spectral Over Factors) values can be corrected by entering the right SOF values provided in Appendix 1. 4. Ensure the right SOF values are entered in the flow cytometry instrument prior to sample acquisition. • Appendix 1 shows the corrected spectral overlap factors for the affected product lot numbers. • Follow the instructions in Appendix 2 for manual entry or to scan the barcodes provided in Appendix 3 to obtain the correct Spectral Over Factors (SOF) values. 5. Customers who have not used the affected product may proceed to enter the SOF values shown in Appendix 1. 6. If the affected products have been fully or partially used and subjects' samples have been processed and acquired on an instrument with the incorrect SOF values, review the results. If the subjects’ sample is available, after entering the correct SOF values in the flow cytometry instrument, reacquire the sample(s). 7. If the subject sample(s) is not available anymore, this may require the subject to return to collect an additional sample for testing purposes. BD Actions: BD is investigating root cause and will implement appropriate corrective and preventative measures to prevent recurrence Customer Actions: • Review the information in Table 1 to determine if BD FACS™ 7-Color Setup Beads in your possession are impacted. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 3rd July 2025. • Circulate this notice to all those who need to be aware within your organisation or to any organisation where the potentially affected product has been transferred. • If you experience any issues please report as a complaint as per your normal process. Distributor Actions: • Review the information in Table 1 and determine if the BD FACS™ 7-Color Setup Beads in your possession are impacted. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 3rd July 2025. o There is no requirement to return your customer response forms to BD, you should maintain these on file at your facility. Return only your final consolidated response form. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues please report as a complaint as per your normal process. | 2025-06-02 13:54
Philips Electronics Nederland B.V. | Philips IntelliSite Pathology Solution | 5 | - | A problem has been detected in the Philips IntelliSite Pathology Solution PIPS x and the Philips IntelliSite Pathology Solution PIPS x for IVD and Non IVD systems x D The Image Management System IMS provides information on free space of available servers If there is insufficient free space e g less than a notification is shown on the screen See Fig below x D For PIPS x the notification is on the information bar displayed at the bottom of the screen You can see x D The connection status for available servers x D An indication for server storage space only visible when storage space on server is low x D For PIPS x the notification is on the information bar displayed at the bottom of the screen x D The message The system is low on storage space is shown when the available free space on the server is less than default setting x D The message The system is running out of storage space Please clear space to continue scanning is shown when the available free space on the server is less than default setting x D The scanner will show the Error message Slide skipped Could not communicate with IMS x D When there is not sufficient free space on the primary storage no new image data from the scanners can be added As described in the Instructions for Use See appendix B regular data cleanup should be performed to create free space on the server to enable the import of new Whole Slide Images WSIs x D The free space calculation depends on the configuration of primary storage and availability of additional storage configured during installation of the IMS software x D a For IMS configured with primary storage and no additional storage x D The free space calculation and free space notifications are correct x D b For IMS configured with primary storage and additional storage x D The total free space percentage of primary storage AND additional storage is calculated x D This may result in the following incorrect situations x D A warning being issued while there is adequate space in the primary storage x D No warning while the primary storage is nearly full but the additional storage has sufficient space | Philips will be contacting you to schedule a remote installation of new software to address the free-space notification issue. FCO76000060 will be released in May 2025. FCO76000061 is already available for remote installation. | 2025-06-03 17:04
Stago | STA OWREN® KOLLER | 135 | - | performance defect with some vials of STA OWREN KOLLER ref lot | The necessary measures were taken, the hospitals were informed, and the reagents were destroyed | 2025-06-04 10:20
Covidien llc | Hugo™ RAS Surgeon Console | 2 | - | Since the release of FA in March there continued to be complaints coming from the field for Surgeon Console power supply failures Investigation uncovered that a bad date code of MOSFETs on the power supply modules were contributing to this second round of power supply failures | • Immediately notify all personnel in all care environments in which the Hugo TM RAS system is used about this Medical Device Correction notice. • The continued use of Hugo™ RAS System is considered appropriate based on an internal review taking into account the benefit provided to patients compared to any potential risk that may be posed. This assessment may be augmented in individual surgeries by determining any circumstances that materially change the benefit or risk. • Complete the attached Customer Confirmation Form and return it as directed to confirm your receipt and understanding of this information. • If you are aware of any incidents related to this issue, please reach out to your Medtronic contact to provide information regarding those events. | 2025-06-04 14:45
Philips Medical Systems Nederland B.V. | AZURION 7 M20 | 1 | - | Philips has identified a potential safety issue in a limited number of Azurion systems manufactured between October and July x D Philips has identified that the bolts connecting the gearbox to the mounting flange of the C arm roll motor may have not been adequately torqued The C arm roll motor is located on the side of the C arm and drives the roll motion of the C arm Figure x D Over time if the bolts are not adequately torqued they may get loose and cause minor shift movements of the gearbox relative to the mounting flange These movements may cause unexpected noise during the C arm roll motion The misalignment between the gearbox and the mounting flange may be detected by the C arm positioning software in which case the message Motorized movement is not available will be shown to the user and motorized roll and propellor movements of the C arm will be disabled Motorized longitudinal stand movement stand rotation and detector movements will remain available A warm or a cold restart may temporarily resolve the issue x D x D If the bolts get loose over time the repeated rotation movement may lead to fatigue and eventually breakage of the bolts In this case the C arm may rotate freely until the tube reaches its lowest position as shown in Figure Motorized roll and propellor motion of the C arm will be unavailable | Philips will replace the C-arm roll motor in all affected systems free of charge. Your local Philips representative will contact you to schedule a visit of a Field Service Engineer to perform this activity. Philips expects the solution to be released in July 2025. | 2025-06-04 17:15
DiaSorin Italia S.p.A. | LIAISON Biotrin Parvovirus B19 IgM | 11 | - | Invalid calibrations or Positive Control out of range Low due to a drop in integral reactivity signal RLU If after obtaining a valid calibration and a valid run by testing calibrators and Kit Controls right after integral shaking as per IFU the kit is stored on board of the instrument there is a potential risk for obtaining False Negative results on daily routines | A Field Safety Corrective Action is recommended to inform customers of the need to: - stop to use the involved lot; - re-evaluate the negative results obtained with this lot of product, considering that, as reported in the IFU, a negative result for IgM antibodies to parvovirus B19 generally indicates that the patient has not been infected, but does not exclude the possibility of acute parvovirus B19 infection. If clinical exposure to parvovirus B19 is suspected despite a negative finding, a retest with a different lot should be evaluated, especially on pregnant women, immunocompromised patients or individuals with underlying hemolytic disorders. The root cause is under investigation and it is probably linked to a decreased reactivity of the component Conjugate. | 2025-06-11 18:14
Terumo | LIFEPEARL™ Drug Elutable Microspheres | 74 | - | the actual x D average diameter of the LIFEPEARL microspheres is smaller than expected in the identified lots x D and they are not within specification | close patient monitoring if implanted and replacement if not implanted | 2025-06-12 14:47
Intersurgical Limited | GUEDEL AIRWAY | 96340 | - | The device is potentially contaminated with small plastic burrs inside the Guedel or packaging from the manufacturing process | Identify, Quarantine and return the device | 2025-06-15 12:04
Randox Laboratories Ltd | RX Daytona + with ISE/ RX Daytona + without ISE / RX Imola / RX Modena /RX Monaco / RX Misano | 125 | - | Randox Laboratories have updated the following analyser manuals to remove references to compliance to x D EU Directive EC and replace this with In vitro diagnostic devices are designed and developed in x D compliance to the regulations and country requirements listed in the regulatory assessment x D Randox Laboratories have also added the RX carryover avoidance information to the RX Imola RX Daytona x D and RX Modena operator manuals | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete updated RX analyser operator manuals in PDF format and individual page updates are available on request. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2025-06-16 12:47
Dexcom, Inc | Dexcom G7 Receiver | 1056 | - | The root cause for the receiver speaker failures observed in the field is a material issue specifically the material of construction Poron foam without PET backing used in speakers obtained from one supplier utilized in Receiver manufacturing East Electronics The investigation found that the material of construction from the other applicable supplier was not deficient and those devices are unimpacted by this issue | Customers will receive notification via email or via physical mail. This notification will advise them to navigate to an online resource to determine if their receiver may be impacted be this issue. If their receiver is identified as impacted, they will then be advised that they may contact technical support to obtain a no-charge replacement. The notification will also remind users to check their systems regularly in accordance with their instructions for use and the prompt received at each charging cycle. Dexcom will also be placing a banner on our website to direct users to this information and will be coordinating with our Distributors to replace any remaining on-hand inventory. | 2025-06-16 17:37
Dexcom, Inc | Dexcom ONE+ Receiver | 1056 | - | The root cause for the receiver speaker failures observed in the field is a material issue specifically the material of construction Poron foam without PET backing used in speakers obtained from one supplier utilized in Receiver manufacturing East Electronics The investigation found that the material of construction from the other applicable supplier was not deficient and those devices are unimpacted by this issue | Customers will receive notification via email or via physical mail. This notification will advise them to navigate to an online resource to determine if their receiver may be impacted be this issue. If their receiver is identified as impacted, they will then be advised that they may contact technical support to obtain a no-charge replacement. The notification will also remind users to check their systems regularly in accordance with their instructions for use and the prompt received at each charging cycle. Dexcom will also be placing a banner on our website to direct users to this information and will be coordinating with our Distributors to replace any remaining on-hand inventory. | 2025-06-16 17:37
Ortho Clinical Diagnostics Inc. | Performance Verifier I | 10 | - | QuidelOrtho received complaints indicating that quality control QC results were reporting below the Range of Means listed on the Performance Verifier Assay Sheets when using VITROS Performance Verifiers I and II to monitor the performance of VITROS CHE Slides x D We investigated the issue and determined that VITROS Performance Verifiers I and II when stored at C will maintain reconstituted stability up to hours but may not maintain reconstituted stability up to days for CHE as specified in the Performance Verifier Assay Sheets x D Our investigation determined that up to hours post reconstitution QC results were within x Within lab Standard Deviation WLSD Beyond hours we determined that VITROS Performance Verifiers I and II may drift below x WLSD limits x D QuidelOrtho advises that customers may continue to use VITROS Performance Verifiers I and II to process QC of VITROS CHE Slides monitoring QC results to detect a drift in performance below x WLSD limits If a drift below x WLSD limits is detected customers are advised to perform QC again using freshly reconstituted vials of VITROS Performance Verifiers I and II x D Summary Cont x D When performing QC to verify a new calibration of VITROS CHE Slides freshly reconstituted vials of VITROS Performance Verifiers should be used as aged VITROS Performance Verifiers may not detect a sub optimal calibration x D x D QuidelOrtho has not seen evidence that suggests other assays supported by VITROS Performance Verifiers are affected by this issue VITROS Performance Verifiers may be used to perform Quality Control according to the stability guidelines listed in the Assay Sheet for all other supported VITROS assays | • Monitor VITROS CHE Slides QC results to detect a drift in VITROS Performance Verifiers I and II below 2x Within-lab SD limits. If QC results are below 2x WLSD, perform QC again using freshly reconstituted vials of VITROS Performance Verifiers I and II. • Complete the enclosed Confirmation of Receipt form no later than Month DD, 2025. • Save this notification with your User Documentation or post this notification near your laboratory’s fluid reconstitution location until the issue has been resolved. • Please forward this notification if the affected products were distributed outside of your facility. • If your laboratory has experienced the issue described in this notification and you have not already done so, please report the occurrence to your local Global Services Organization. | 2025-06-17 13:24
Ortho Clinical Diagnostics Inc. | Performance Verifier II | 10 | - | QuidelOrtho received complaints indicating that quality control QC results were reporting below the Range of Means listed on the Performance Verifier Assay Sheets when using VITROS Performance Verifiers I and II to monitor the performance of VITROS CHE Slides x D We investigated the issue and determined that VITROS Performance Verifiers I and II when stored at C will maintain reconstituted stability up to hours but may not maintain reconstituted stability up to days for CHE as specified in the Performance Verifier Assay Sheets x D Our investigation determined that up to hours post reconstitution QC results were within x Within lab Standard Deviation WLSD Beyond hours we determined that VITROS Performance Verifiers I and II may drift below x WLSD limits x D QuidelOrtho advises that customers may continue to use VITROS Performance Verifiers I and II to process QC of VITROS CHE Slides monitoring QC results to detect a drift in performance below x WLSD limits If a drift below x WLSD limits is detected customers are advised to perform QC again using freshly reconstituted vials of VITROS Performance Verifiers I and II x D Summary Cont x D When performing QC to verify a new calibration of VITROS CHE Slides freshly reconstituted vials of VITROS Performance Verifiers should be used as aged VITROS Performance Verifiers may not detect a sub optimal calibration x D x D QuidelOrtho has not seen evidence that suggests other assays supported by VITROS Performance Verifiers are affected by this issue VITROS Performance Verifiers may be used to perform Quality Control according to the stability guidelines listed in the Assay Sheet for all other supported VITROS assays | • Monitor VITROS CHE Slides QC results to detect a drift in VITROS Performance Verifiers I and II below 2x Within-lab SD limits. If QC results are below 2x WLSD, perform QC again using freshly reconstituted vials of VITROS Performance Verifiers I and II. • Complete the enclosed Confirmation of Receipt form no later than Month DD, 2025. • Save this notification with your User Documentation or post this notification near your laboratory’s fluid reconstitution location until the issue has been resolved. • Please forward this notification if the affected products were distributed outside of your facility. • If your laboratory has experienced the issue described in this notification and you have not already done so, please report the occurrence to your local Global Services Organization. | 2025-06-17 13:24
Ortho Clinical Diagnostics Inc. | VITROS Chemistry Products OP Reagent | 2 | - | QuidelOrtho advises that customers who are unable to generate QC results within acceptable range should discontinue using render unusable and discard your inventory of lots from GEN of VITROS OP Reagent QuidelOrtho will provide credit or replace your discarded inventory with lots of VITROS OP Reagent that are unaffected by this issue x D Please note that our investigation has also confirmed that the OP HI protocol is not affected by this issue | • If your laboratory is unable to generate QC results within acceptable range using lots from GEN 15 of VITROS OP Reagent, discontinue using, render unusable, and discard your inventory of VITROS OP Reagent, GEN 15. • Complete and return the enclosed Confirmation of Receipt form no later than Month DD, 2025. Upon receipt of your completed Confirmation of Receipt form, QuidelOrtho will provide credit or replace your discarded inventory with unaffected lots of VITROS OP Reagent. • Save this notification with your User Documentation or post this notification by each VITROS 5,1/4600/5600/XT 7600 System in your laboratory until the issue has been resolved. • Please forward this notification if the affected product was distributed outside of your facility. • If your laboratory has experienced the issue with this product and you have not already done so, please report the occurrence to your local Global Services Organization. | 2025-06-17 13:29
Maquet Cardiopulmonary GmbH | SPRINTER CART | 123 | - | According to IEC AMD AMD clause Instability from horizontal and x D vertical forces if an equipment overbalances when being pushed leaned on etc with of its weight the x D medical electrical equipment shall be marked with a clearly legible warning e g safety sign according to ISO x D P Figure x D Starting April MCP implemented a safety sign on newly produced INFUSION POLEs but did not x D correct INFUSION POLEs manufactured before this date Therefore all INFUSION POLEs manufactured x D before April are affected by this issue | Verify presence of ‘Pushing prohibited’ label on each INFUSION POLE of SPRINTER CART/ SPRNTER CART XL during next service visit • If not present, affix ‘Pushing prohibited’ label to INFUSION POLE of SPRINTER CART/ SPRNTER CART XL during next service visit | 2025-06-17 15:29
Randox Laboratories Ltd | Evidence + Analyser | 24 | - | The Evidence Operator Manual has been updated to include a section on Cyber Security x D informing customers of security and cybersecurity recommendations x D Additionally Randox Laboratories Ltd has identified a minor issue in software version x D that affects the sample code search function Specifically the function may not return results x D for samples that include special characters or symbols in the sample code and may be x D accompanied by an error message | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to support@randoxtoxicology.com within five working days. | 2025-06-18 09:42
Canon Medical Systems Corporation | Vantage Elan-Vantage Orian | 8 | - | We have recently received a report of a case in which a patient experienced auditory disturbances while using Canon MRI device In this case a patient underwent an MRI scan without sufficient hearing protection and subsequently reported symptoms of hearing loss due to exposure to high levels of noise x D x D | The device is equipped with a noise reduction mechanism, called Pianissimo that lowers the noise level; however, depending on the imaging method, the noise level may reach a level where hearing protection is essential. To ensure the safety of patients in using MRI System, please carefully follow the guidelines outlined in Chapter 8 of the safety manual. | 2025-06-18 19:25
Interojo INC | Bella Elite Green Olive | 39 | - | The product in question has not been traded or sold as it is a sample | The product in question has not been traded or sold, as it is a sample. | 2025-06-18 20:28
Roche Diagnostics GmbH | aPTT Screen cobas t 600T aPTT Lupus cobas t 600T aPTT cobas t 600T | 109 | - | The carryover occurs when the coupling nut is not tightened sufficiently In this case aPTT Screen aPTT Lupus and x D aPTT can be affected by carryover | Customers are requested to mount the heated probe according to the maintenance workflow or user assistance. Special attention is required during the mounting of the heated probe (Also see attachment 1). ● In addition, to ensure correct results for samples with aPTT Screen clotting times >50 sec, customers are requested to keep the already implemented reflex test and to make sure to have the latest carryover file installed (carryover evasion files are automatically installed with SW 2.3.0 or higher). See attached instructions (attachment 1), aPTT Screen IFU and user assistance. | 2025-06-19 16:29
Abiomed Inc. | Automated Impella Controller (AIC) | 58 | - | Abiomed has identified an AIC issue that may prevent the detection of an Impella pump when connected to an AIC The pump detection issue may occur with any of the Abiomed Impella pumps and may occur during console to console transfer or at initiation of therapy case start There is no visual alarm displayed on the AIC screen to indicate the detection issue in these situations x D | 1-Product is NOT being removed from the field and does not need to be returned. 2-As indicated in IFUs, have a back-up Automated Impella Controller (AIC) available in the unlikely event of a device failure. 3- Corrective actions will be defined and managed through the service team (CAPA-014370). 4-Keep the copy of the FSN together with the IFU | 2025-06-24 16:06
Beckman Coulter Inc. | UniCel DxI 800 Access Immunoassay System | 21 | - | Beckman Coulter has determined that UniCel DxI systems may experience an increase in pick and place gantry events Specific events you may encounter in the Event Log include x D Pick and Place hit an unexpected vessel while attempting to place vessel x D X Home sensor in unexpected position x D Vessel missing or not detected during pick and place move Wash Pick and Place x D Pick and place detected an unexpected vessel Wash Pick and Place x D Vessel not detected for pickup Wash Pick and Place x D Vessel missing or not detected during pick and place move Incubator Pick and Place x D Pick and place detected an unexpected vessel Incubator Pick and Place x D Vessel not detected for pickup Incubator Pick and Place x D Vessel missing or not detected during pick and place move Sample Pick and Place x D Pick and place detected an unexpected vessel Sample Pick and Place Vessel not detected for pickup Sample Pick and Place | If your laboratory encounters any of the events described in this letter: 􀁸 Perform a system initialization to resume normal operation. o If repeated initialization attempts fail, contact Customer Technical Support. 􀁸 Repeat testing for any cancelled tests. 􀁸 No action is required if your laboratory has not observed the events listed above. | 2025-06-24 16:11
Intuitive surgical Inc. | Universal Surgical Manipulators | 33 | - | This notice provides updated information regarding the Universal Surgical Manipulators USMs affected by this Field Action x D In early Intuitive notified customers of a potential issue with certain USMs and implemented corrective actions only two complaints have been reported with no adverse events x D A recent investigation while indicating low likelihood suggests that additional USMs may be affected in the field x D As a precaution Intuitive will inspect the USMs listed in Appendix B there is no increased risk and the device remains safe for use x D This letter is to inform you that a population of the da Vinci X and Xi system instrument arms USMs have Instrument Carriages that may separate from the Insertion Axis Linear Rail during shipping handling or repositioning of the arms when the system is powered down The Instrument Carriage would remain attached to the drive motor which prevents free fall motion or complete detachment from the USM However the Instrument Carriage may feel like it is not tightly connected to the rail This condition may also lead to increased friction along the insertion axis x D Therefore Intuitive would like to schedule a visit to perform inspection of the affected USM and conduct replacement of the USM if necessary x D Instrument Carriages that are intact will function as intended and are expected to remain firmly attached throughout a surgical procedure x D | All affected systems remain safe for use if the following instructions are followed: see atatched FSN. An Intuitive Representative will schedule a site visit to inspect the insertion linear rail and replace the USM if necessary. | 2025-06-24 18:26
Medtronic Inc. | CareLink SmartSync | 87 | - | The affected device may be susceptible to erroneous electrical reset pop up messages The erroneous message indicates an electrical reset has occurred However no reset has actually occurred and there are no parameter changes or loss of diagnostics | Until the updated versions are available, if an electrical reset (pop-up) message presents during a SmartSync interrogation, it may be erroneous. If it is unclear if an electrical reset has occurred or if the message is erroneous, please contact your local Medtronic Representative and provide a Save to Media file and the SmartSync application ID. When the updated versions are approved and available, your local Medtronic representative can assist you with the installation. | 2025-06-25 09:07
Medtronic Inc. | CareLink SmartSync | 87 | - | The affected device may be susceptible to Abort button behavior inconsistent with user expectations and legacy programmers | Until the updated versions are available, the abort button will terminate an in-progress test (ex. T shock) but will not terminate detection, high-voltage charging, and/or therapy deliveries following an induction. When the updated versions are available, your local Medtronic representative can assist you with the installation. | 2025-06-25 09:18
GE Medical Systems, LLC | SIGNA Premier | 74 | - | GE HealthCare has become aware that for certain MR systems see affected products list below the planned maintenance steps to check the functionality of the caster locks on the MR patient table may not have been performed The functional check is intended to proactively detect potential failure of table caster locks which could result in unintended table movement and unstable patient support when the table is not docked to the system | Perform Table PM inspection and tests as prescribed in Service Methods PMs. | 2025-06-25 14:55
GE Medical Systems, LLC | SIGNA Artist | 74 | - | GE HealthCare has become aware that for certain MR systems see affected products list below the planned maintenance steps to check the functionality of the caster locks on the MR patient table may not have been performed The functional check is intended to proactively detect potential failure of table caster locks which could result in unintended table movement and unstable patient support when the table is not docked to the system | Perform Table PM inspection and tests as prescribed in Service Methods PMs. | 2025-06-25 14:55
Randox Laboratories Ltd | Ransel Glutathione Peroxidase | 2 | - | Randox Laboratories can confirm that Ransel Glutathione Peroxidase reagent catalogue number RS is now x D discontinued along with its associated reagents Ransel Diluting Agent catalogue number RS and x D Haemoglobin Reagent catalogue number HG The catalogue number RS will remain available to x D purchase however it will no longer be CE marked from the th of July From this date the intended use of x D RS will move to research and veterinary use only and will not be submitted for IVDR The Instructions For x D Use IFU have been updated to reflect this and are available on request | Action to be taken: Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Please accept our apologies for any inconvenience caused. Thank you for your patience and understanding. If you have any questions or concerns, please contact Randox Technical Services. | 2025-06-26 14:19
.Siemens Healthcare Diagnostics Inc | 3-in-1 TESTpoint ABN1 Control | 109 | - | Issue Description x D Siemens Healthineers has confirmed through the investigation of customer complaints that the products listed in the above table have incorrect platelet PLT value assignments for the ADVIA i Hematology Systems x D Erroneously decreased patient platelet results may be reported if calibrator SETpoint lot SP or SP are used and controls are recovering within the expected assigned value ranges Siemens investigation revealed a deviation of up to in patient results across the assay range when these calibrators are used x D Since calibrators and controls may be used in any combination the table below provides the possible scenarios and impact to PLT results x D Siemens has confirmed that the behavior is limited to PLT results generated with the indicated SETpoint and TESTpoint lots x D x D Kindly review the attached Urgent Field Safety Notice | Customer Actions Please perform the instructions provided below: • Customers that have calibrated using SETPOINT lot SP251033 or SP251043 should perform a calibration using a subsequent lot. Customers that have not calibrated using SETPOINT lot SP251033 or SP251043 do not need to perform additional calibration. • Discontinue use of and discard the calibrator SETPOINT lots SP251033 and SP251043. • If you experience this issue, you may request credit for SETpoint products in Field Correction Effectiveness Check Form. Please review your inventory of these products and assess your laboratory’s needs. • Controls with PLT results that are within the expected assigned value ranges may continue to be used. • Customers who observe PLT results that are consistently out of the expected assigned value ranges for the TESTpoint lots TP251035, TP252035, TP253035, TP251045, TP252045, TP253045 should discontinue use of and discard them and you may request credit for TESTpoint products in Field Correction Effectiveness Check Form. • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Please retain this letter with your laboratory records and forward this letter to those who may have received this product. • Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. Kindly review the attached Urgent Field Safety Notice | 2025-06-26 15:05
.Siemens Healthcare Diagnostics Inc | 3-in-1 TESTpoint ABN2 Control | 109 | - | Issue Description x D Siemens Healthineers has confirmed through the investigation of customer complaints that the products listed in the above table have incorrect platelet PLT value assignments for the ADVIA i Hematology Systems x D Erroneously decreased patient platelet results may be reported if calibrator SETpoint lot SP or SP are used and controls are recovering within the expected assigned value ranges Siemens investigation revealed a deviation of up to in patient results across the assay range when these calibrators are used x D Since calibrators and controls may be used in any combination the table below provides the possible scenarios and impact to PLT results x D Siemens has confirmed that the behavior is limited to PLT results generated with the indicated SETpoint and TESTpoint lots x D x D Kindly review the attached Urgent Field Safety Notice | Customer Actions Please perform the instructions provided below: • Customers that have calibrated using SETPOINT lot SP251033 or SP251043 should perform a calibration using a subsequent lot. Customers that have not calibrated using SETPOINT lot SP251033 or SP251043 do not need to perform additional calibration. • Discontinue use of and discard the calibrator SETPOINT lots SP251033 and SP251043. • If you experience this issue, you may request credit for SETpoint products in Field Correction Effectiveness Check Form. Please review your inventory of these products and assess your laboratory’s needs. • Controls with PLT results that are within the expected assigned value ranges may continue to be used. • Customers who observe PLT results that are consistently out of the expected assigned value ranges for the TESTpoint lots TP251035, TP252035, TP253035, TP251045, TP252045, TP253045 should discontinue use of and discard them and you may request credit for TESTpoint products in Field Correction Effectiveness Check Form. • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Please retain this letter with your laboratory records and forward this letter to those who may have received this product. • Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. Kindly review the attached Urgent Field Safety Notice | 2025-06-26 15:05
.Siemens Healthcare Diagnostics Inc | 3-in-1 TESTpoint NORM Control | 109 | - | Issue Description x D Siemens Healthineers has confirmed through the investigation of customer complaints that the products listed in the above table have incorrect platelet PLT value assignments for the ADVIA i Hematology Systems x D Erroneously decreased patient platelet results may be reported if calibrator SETpoint lot SP or SP are used and controls are recovering within the expected assigned value ranges Siemens investigation revealed a deviation of up to in patient results across the assay range when these calibrators are used x D Since calibrators and controls may be used in any combination the table below provides the possible scenarios and impact to PLT results x D Siemens has confirmed that the behavior is limited to PLT results generated with the indicated SETpoint and TESTpoint lots x D x D Kindly review the attached Urgent Field Safety Notice | Customer Actions Please perform the instructions provided below: • Customers that have calibrated using SETPOINT lot SP251033 or SP251043 should perform a calibration using a subsequent lot. Customers that have not calibrated using SETPOINT lot SP251033 or SP251043 do not need to perform additional calibration. • Discontinue use of and discard the calibrator SETPOINT lots SP251033 and SP251043. • If you experience this issue, you may request credit for SETpoint products in Field Correction Effectiveness Check Form. Please review your inventory of these products and assess your laboratory’s needs. • Controls with PLT results that are within the expected assigned value ranges may continue to be used. • Customers who observe PLT results that are consistently out of the expected assigned value ranges for the TESTpoint lots TP251035, TP252035, TP253035, TP251045, TP252045, TP253045 should discontinue use of and discard them and you may request credit for TESTpoint products in Field Correction Effectiveness Check Form. • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Please retain this letter with your laboratory records and forward this letter to those who may have received this product. • Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. Kindly review the attached Urgent Field Safety Notice | 2025-06-26 15:05
.Siemens Healthcare Diagnostics Inc | SETpoint CAL | 109 | - | Issue Description x D Siemens Healthineers has confirmed through the investigation of customer complaints that the products listed in the above table have incorrect platelet PLT value assignments for the ADVIA i Hematology Systems x D Erroneously decreased patient platelet results may be reported if calibrator SETpoint lot SP or SP are used and controls are recovering within the expected assigned value ranges Siemens investigation revealed a deviation of up to in patient results across the assay range when these calibrators are used x D Since calibrators and controls may be used in any combination the table below provides the possible scenarios and impact to PLT results x D Siemens has confirmed that the behavior is limited to PLT results generated with the indicated SETpoint and TESTpoint lots x D x D Kindly review the attached Urgent Field Safety Notice | Customer Actions Please perform the instructions provided below: • Customers that have calibrated using SETPOINT lot SP251033 or SP251043 should perform a calibration using a subsequent lot. Customers that have not calibrated using SETPOINT lot SP251033 or SP251043 do not need to perform additional calibration. • Discontinue use of and discard the calibrator SETPOINT lots SP251033 and SP251043. • If you experience this issue, you may request credit for SETpoint products in Field Correction Effectiveness Check Form. Please review your inventory of these products and assess your laboratory’s needs. • Controls with PLT results that are within the expected assigned value ranges may continue to be used. • Customers who observe PLT results that are consistently out of the expected assigned value ranges for the TESTpoint lots TP251035, TP252035, TP253035, TP251045, TP252045, TP253045 should discontinue use of and discard them and you may request credit for TESTpoint products in Field Correction Effectiveness Check Form. • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Please retain this letter with your laboratory records and forward this letter to those who may have received this product. • Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. Kindly review the attached Urgent Field Safety Notice | 2025-06-26 15:05
HAMILTON MEDICAL AG | HAMILTON‐C6 | 10 | - | Damage to the capacitor could over time lead to a loss of its function due to x D oxidation or similar degradation processes The worst case scenario is a local short x D circuit of the aforementioned capacitor with different outcomes depending on the x D device state x D HAMILTON C is going to be started The affected HAMILTON C device could then x D not be started x D HAMILTON C is already started and in use for ventilation with a patient If the x D short circuit would occur during use of the HAMILTON C ventilator the x D ventilation would be interrupted but the HAMILTON C device would nevertheless x D alarm about ambient state or panel connection loss The ventilator must then be x D replaced by an alternative device | Hamilton Medical AG has decided to pro-actively repair affected HAMILTON-C6 devices by exchanging the ESM Board through the responsible local organizations/partners. | 2025-06-26 16:48
Invacare Portugal, Lda. - Sociedade Industrial e Comercial de Ortopedia | Birdie EVO | - | - | Following the awareness of a serious adverse event related to Invacare Birdie lifter and carabineer attachment Invacare wants to remind of the importance to ensure the correct assembly position of the carabineer in the Birdie lifter arm and as well the correct attachment of the hanger bar into the carabineer x D Invacare Birdie lifter instructions for use explains in detail the correct use and handling of the attachment of carabineer and hanger bar Section Mounting the Hanger bar x D Furthermore it is important to verify that carabineer remains always in the correct position of assembly x D To ensure that carabineer can t be able to completely rotate without resistance within the plastic parts where it is assembled If that occurs it means the plastic parts present wear and need to be replaced x D x D If the carabiner is not attached properly to the spreader bar meaning according to the Birdie lifter Instructions For Use it might led to the fall of the patient serious injuries death x D In total Invacare became aware of cases worldwide cases in Europe and cases in UK | This information is relevant for all users of the Invacare Birdie lifter, meaning all manufactured Birdie lifters regardless of the Serial number. Emails containing the FSN and Addendum will be sent by registered email to ensure that the recipient can verify receipt. If no confirmation is received within a month, FSNs will be sent by registered post, so we are able to get confirmation about receipt. 1. Provider should inform all affected users about the content of this Field Safety Notice. 2. The Addendum should be kept together with the current Birdie lifter Instructions for use. 3. This updated Maintenance checklist must always be the basis to perform the regular Maintenance of the device. | 2025-06-29 11:55
Beckman Coulter Inc. | UIBC | 1321 | - | Beckman Coulter has determined through internal testing that UIBC reagent REF OSR fails to meet the hemolytic interference specification as stated in the Instructions for Use IFU listed below x D IFU BAOSR X x D No significant interference up to mg dL Hemolysate where the criteria for no significant interference is within of the initial value x D IFU BLOSR X x D Interference less than up to g L haemoglobin x D x D The magnitude of impact varies depending on the level of hemolysis present in the sample Samples with a hemoglobin concentration above mg dL or g L demonstrate clinically significant hemolytic interference of approximately | Beckman Coulter recommends • Sharing the content of this letter with your laboratory and/or Medical Director to evaluate the requirement for a retrospective review of UIBC results for hemolysed samples. • Discontinuance or disposal of this product is not necessary. • Per the IFU, avoid hemolysed samples when using the UIBC assay. If the LIH influence check setting for hemolysis has been enabled on AU/DxC AU instruments for UIBC, update the setting as follows: AU480 / AU680 / AU5800: • Select Menu > Parameters > Specific Test Parameters > General • Select UIBC test name and select Edit • Update the LIH influence check to ++ DxC 700 AU: • Select Menu > Configuration Parameters > Specific Test Parameters > Test Volume and Methods > General • Select UIBC test name and select Edit • Update the LIH influence check to ++ DxC 500 AU / DxC 500i: • Select Menu > System Configuration > Test Menu • Select the UIBC test name and select Chemistry Details • Select Customer Configurable Options • Update the LIH influence check to ++ • Select System Configuration > Activate Draft | 2025-06-29 12:01
Philips Medical Systems Nederland B.V. | Achieva 3.0T | 60 | - | Philips has identified an issue where component failures in the Gradient Coil of the affected MR systems listed in Section may act as a heat source with a potential to produce smoke and or fire x D x D If a component failure occurs the user may observe the following x D Repeated scan abort error messages due to gradient amplifier fault detection described in Section x D Unusual noise from the system in the examination or technical room x D Burning smell x D Smoke and or fire | Philips is committed to addressing the issue and plans to start delivery of a software risk control measure starting in Q3 2025. This feature will prevent smoke/fire by prohibiting customers from continued scanning if gradient amplifier errors reach a threshold, thus triggering a system interlock. If this system interlock occurs, which prevents further scanning, an FSE will come to your site to investigate the issue and if it is identified that there is a problem with the gradient coil, it will be replaced | 2025-06-30 15:43
Philips Medical Systems Nederland B.V. | Ingenia 1.5T | 60 | - | Philips has identified an issue where component failures in the Gradient Coil of the affected MR systems listed in Section may act as a heat source with a potential to produce smoke and or fire x D x D If a component failure occurs the user may observe the following x D Repeated scan abort error messages due to gradient amplifier fault detection described in Section x D Unusual noise from the system in the examination or technical room x D Burning smell x D Smoke and or fire | Philips is committed to addressing the issue and plans to start delivery of a software risk control measure starting in Q3 2025. This feature will prevent smoke/fire by prohibiting customers from continued scanning if gradient amplifier errors reach a threshold, thus triggering a system interlock. If this system interlock occurs, which prevents further scanning, an FSE will come to your site to investigate the issue and if it is identified that there is a problem with the gradient coil, it will be replaced | 2025-06-30 15:43
Philips Medical Systems Nederland B.V. | Ingenia 1.5T S | 60 | - | Philips has identified an issue where component failures in the Gradient Coil of the affected MR systems listed in Section may act as a heat source with a potential to produce smoke and or fire x D x D If a component failure occurs the user may observe the following x D Repeated scan abort error messages due to gradient amplifier fault detection described in Section x D Unusual noise from the system in the examination or technical room x D Burning smell x D Smoke and or fire | Philips is committed to addressing the issue and plans to start delivery of a software risk control measure starting in Q3 2025. This feature will prevent smoke/fire by prohibiting customers from continued scanning if gradient amplifier errors reach a threshold, thus triggering a system interlock. If this system interlock occurs, which prevents further scanning, an FSE will come to your site to investigate the issue and if it is identified that there is a problem with the gradient coil, it will be replaced | 2025-06-30 15:43
Philips Medical Systems Nederland B.V. | Ingenia 3.0T | 60 | - | Philips has identified an issue where component failures in the Gradient Coil of the affected MR systems listed in Section may act as a heat source with a potential to produce smoke and or fire x D x D If a component failure occurs the user may observe the following x D Repeated scan abort error messages due to gradient amplifier fault detection described in Section x D Unusual noise from the system in the examination or technical room x D Burning smell x D Smoke and or fire | Philips is committed to addressing the issue and plans to start delivery of a software risk control measure starting in Q3 2025. This feature will prevent smoke/fire by prohibiting customers from continued scanning if gradient amplifier errors reach a threshold, thus triggering a system interlock. If this system interlock occurs, which prevents further scanning, an FSE will come to your site to investigate the issue and if it is identified that there is a problem with the gradient coil, it will be replaced | 2025-06-30 15:43
Philips Medical Systems Nederland B.V. | INGENIA AMBITION S | 60 | - | Philips has identified an issue where component failures in the Gradient Coil of the affected MR systems listed in Section may act as a heat source with a potential to produce smoke and or fire x D x D If a component failure occurs the user may observe the following x D Repeated scan abort error messages due to gradient amplifier fault detection described in Section x D Unusual noise from the system in the examination or technical room x D Burning smell x D Smoke and or fire | Philips is committed to addressing the issue and plans to start delivery of a software risk control measure starting in Q3 2025. This feature will prevent smoke/fire by prohibiting customers from continued scanning if gradient amplifier errors reach a threshold, thus triggering a system interlock. If this system interlock occurs, which prevents further scanning, an FSE will come to your site to investigate the issue and if it is identified that there is a problem with the gradient coil, it will be replaced | 2025-06-30 15:43
Philips Medical Systems Nederland B.V. | INGENIA AMBITION X | 60 | - | Philips has identified an issue where component failures in the Gradient Coil of the affected MR systems listed in Section may act as a heat source with a potential to produce smoke and or fire x D x D If a component failure occurs the user may observe the following x D Repeated scan abort error messages due to gradient amplifier fault detection described in Section x D Unusual noise from the system in the examination or technical room x D Burning smell x D Smoke and or fire | Philips is committed to addressing the issue and plans to start delivery of a software risk control measure starting in Q3 2025. This feature will prevent smoke/fire by prohibiting customers from continued scanning if gradient amplifier errors reach a threshold, thus triggering a system interlock. If this system interlock occurs, which prevents further scanning, an FSE will come to your site to investigate the issue and if it is identified that there is a problem with the gradient coil, it will be replaced | 2025-06-30 15:43
Philips Medical Systems Nederland B.V. | INGENIA ELITION X | 60 | - | Philips has identified an issue where component failures in the Gradient Coil of the affected MR systems listed in Section may act as a heat source with a potential to produce smoke and or fire x D x D If a component failure occurs the user may observe the following x D Repeated scan abort error messages due to gradient amplifier fault detection described in Section x D Unusual noise from the system in the examination or technical room x D Burning smell x D Smoke and or fire | Philips is committed to addressing the issue and plans to start delivery of a software risk control measure starting in Q3 2025. This feature will prevent smoke/fire by prohibiting customers from continued scanning if gradient amplifier errors reach a threshold, thus triggering a system interlock. If this system interlock occurs, which prevents further scanning, an FSE will come to your site to investigate the issue and if it is identified that there is a problem with the gradient coil, it will be replaced | 2025-06-30 15:43
Philips Medical Systems Nederland B.V. | MR 5300 | 60 | - | Philips has identified an issue where component failures in the Gradient Coil of the affected MR systems listed in Section may act as a heat source with a potential to produce smoke and or fire x D x D If a component failure occurs the user may observe the following x D Repeated scan abort error messages due to gradient amplifier fault detection described in Section x D Unusual noise from the system in the examination or technical room x D Burning smell x D Smoke and or fire | Philips is committed to addressing the issue and plans to start delivery of a software risk control measure starting in Q3 2025. This feature will prevent smoke/fire by prohibiting customers from continued scanning if gradient amplifier errors reach a threshold, thus triggering a system interlock. If this system interlock occurs, which prevents further scanning, an FSE will come to your site to investigate the issue and if it is identified that there is a problem with the gradient coil, it will be replaced | 2025-06-30 15:43
Stryker | HeartSine ® Samaritan ® PAD | 155 | - | It was determined during extensive quality testing that a manufacturing process x D issue related to a circuit board component may impair the device s ability to x D function or cause failure This failure could occur at any point when the device is holding a charge in preparation to deliver therapy while delivering a shock or x D after shock delivery The device becomes inoperable after the failure occurs x D | Stryker is notifying all customers who have received affected HeartSine Samaritan PAD devices to perform the actions outlined above. Once your response is received and a replacement is available, Stryker will be in contact as soon as possible to arrange the next steps for a replacement device | 2025-07-01 13:40
Changzhou Shuangma Medical Devices Co Ltd | Fine micro Cannula | - | المنتج غير حاصل على اذن التسويق | - | The medical device has been recalled from the market. | 2025-07-01 13:52
Philips Medical Systems Nederland B.V. | Allura Xper FD10 | 37 | - | Philips has identified a software issue in the internal communication process between the system software and the X ray generator firmware | Philips will address the identified issue by implementing a software update in all affected systems. | 2025-07-01 17:18
Philips Medical Systems Nederland B.V. | Allura Xper FD10/10 | 37 | - | Philips has identified a software issue in the internal communication process between the system software and the X ray generator firmware | Philips will address the identified issue by implementing a software update in all affected systems. | 2025-07-01 17:18
Philips Medical Systems Nederland B.V. | Allura Xper FD20 | 37 | - | Philips has identified a software issue in the internal communication process between the system software and the X ray generator firmware | Philips will address the identified issue by implementing a software update in all affected systems. | 2025-07-01 17:18
Philips Medical Systems Nederland B.V. | Allura Xper FD20/10 | 37 | - | Philips has identified a software issue in the internal communication process between the system software and the X ray generator firmware | Philips will address the identified issue by implementing a software update in all affected systems. | 2025-07-01 17:18
Philips Medical Systems Nederland B.V. | Allura Xper FD10 OR Table | 37 | - | Philips has identified a software issue in the internal communication process between the system software and the X ray generator firmware | Philips will address the identified issue by implementing a software update in all affected systems. | 2025-07-01 17:18
Philips Medical Systems Nederland B.V. | Allura Xper FD20 OR Table | 37 | - | Philips has identified a software issue in the internal communication process between the system software and the X ray generator firmware | Philips will address the identified issue by implementing a software update in all affected systems. | 2025-07-01 17:18
Philips Medical Systems Nederland B.V. | Allura Xper FD20/15 | 37 | - | Philips has identified a software issue in the internal communication process between the system software and the X ray generator firmware | Philips will address the identified issue by implementing a software update in all affected systems. | 2025-07-01 17:18
Randox Laboratories Ltd. | URINALYSIS CONTROL - LEVEL 2 (URNAL CONTROL 2) | 30 | - | Randox Laboratories can confirm that the urinalysis control Level lot UC and UC may show negative x D results for Leukocytes within the claimed storage and stability period for some Methods If you are experiencing x D this issue please contact technical services randox com with data to demonstrate the issue and replacements x D will be arranged No action is required if you are reporting positive results for Leucocytes | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Please accept our apologies for any inconvenience caused. Thank you for your patience and understanding. If you have any questions or concerns, please contact Randox Technical Services. | 2025-07-03 12:24
Medtronic Minimed | InPen Application | 2 | - | During in house testing of the InPen app several issues with the Insulin Notification Service INS were identified specifically affecting Daytime notifications Key findings include x D If the Nighttime schedule starts after midnight e g AM Nighttime notifications are disabled and Daytime notifications are enabled AM Daytime notifications fail to trigger x D When the start time for Day or Night is set to AM and the user swipe kills the app or navigates to the settings screen the app resets AM to default Breakfast or Nighttime start times x D Adjusting the Breakfast start time automatically updates the Daytime start time to match the new Breakfast start time x D These issues collectively prevent Daytime notifications from being triggered as expected x D | Users of InPen app software versions 7.0.0, 7.0.1, 7.1.0, 7.1.1, 7.2.0 and 7.5.0 should update to version 7.5.1 or higher as soon as possible to resolve this issue if they have not already done so. | 2025-07-06 09:39
Liko AB | Uno 102 EE | 2 | - | The Q link allows for an improper attachment false latching of the quick release hook used on sling bars and x D other accessories | 1. Until your Q-link 13 is replaced with the Q-link mobile, customers may continue to use the current Q-link 13 component with added precautions to ensure proper attachment prior to use with patients. 2. Please post this letter in areas where affected mobile lifts are stored and used. 3. A follow-up communication will be provided once sufficient Q-link Mobile components are available for distribution, and will include instructions to request replacements. 4. Complete the enclosed customer reply form and return it to Baxter by scanning and e-mailing it to Bandar_Alosaimi@Baxter.com, even if you don’t have any inventory. Returning the customer reply form promptly will confirm your receipt of this notification and prevent you from receiving repeat notices. 5. If you distributed this product to other facilities or departments within your institution, please forward a copy of this communication to them. 6. If you purchased this product from a distributor, please note that the Baxter customer reply form is not applicable. If a reply form is provided by your distributor or wholesaler, please return it to the supplier according to their instructions. | 2025-07-06 09:56
Liko AB | Viking | 2 | - | The Q link allows for an improper attachment false latching of the quick release hook used on sling bars and x D other accessories | 1. Until your Q-link 13 is replaced with the Q-link mobile, customers may continue to use the current Q-link 13 component with added precautions to ensure proper attachment prior to use with patients. 2. Please post this letter in areas where affected mobile lifts are stored and used. 3. A follow-up communication will be provided once sufficient Q-link Mobile components are available for distribution, and will include instructions to request replacements. 4. Complete the enclosed customer reply form and return it to Baxter by scanning and e-mailing it to Bandar_Alosaimi@Baxter.com, even if you don’t have any inventory. Returning the customer reply form promptly will confirm your receipt of this notification and prevent you from receiving repeat notices. 5. If you distributed this product to other facilities or departments within your institution, please forward a copy of this communication to them. 6. If you purchased this product from a distributor, please note that the Baxter customer reply form is not applicable. If a reply form is provided by your distributor or wholesaler, please return it to the supplier according to their instructions. | 2025-07-06 09:56
Randox Laboratories Ltd. | LIQ URN CONTROL 3 | 1182 | - | Randox Laboratories can confirm there has been a transcription error in the Instructions for Use IFU for x D Liquid Urine Control Level lot UC The analyte Inorganic Phosphate has been incorrectly listed as x D method Pyrogallol Red and should be Phosphomolybdate UV in the Beckman Coulter AU Series section x D The Targets and Ranges have not been affected by this error x D | Randox Laboratories can confirm there has been a transcription error in the Instructions for Use (IFU) for Liquid Urine Control Level 3, lot 1390UC. The analyte Inorganic Phosphate has been incorrectly listed as method Pyrogallol Red and should be Phosphomolybdate UV in the Beckman Coulter AU Series section. The Targets and Ranges have not been affected by this error. | 2025-07-06 12:19
INOVA Diagnostics, Inc. | Quanta Flash Centromere Controls | 4 | - | Inova Diagnostics Inc a Werfen Company is initiating this field action This notification only affects the QUANTA Flash Centromere Controls lot x D The QUANTA Flash Centromere Controls Instructions for Use IFU recommends that the controls be run once each day that the assay will be used however users should also consider national local regulatory requirements x D o Out of control QC results will be flagged by the BIO FLASH instrument if the results do not meet the criteria for acceptance x D o If out of control QC results are flagged sample results cannot be reported until the issue is resolved x D A valid QC is required to process patient samples While this issue does not impact patient results the reporting of results may be delayed until an invalid QC result is resolved | • If your facility currently has the QUANTA Flash® Centromere Controls lot 233974 in stock, discontinue use and discard any remaining inventory. o Follow local regulatory requirements for proper disposal of products. o Transition to an alternative lot for testing. • If QC was previously run and met established acceptance criteria, the results are considered valid and no further action is required. • Please contact your local Werfen representative if a replacement product is needed. | 2025-07-07 14:52
Maquet Cardiopulmonary GmbH | Heater Unit HU 35, 230V | 121 | - | Based on the internal complaints investigation and extensive material compatibility testing under aggrevated x D conditions simulating more robust real time aging and material stress it was exposed that the current HU x D water tank fittings made of polyamide have a limited resistance to Chloramine T used for disinfecting the water x D circuit of the unit x D Failure to correctly apply the disinfection procedures of the water circuit e g wrong dosing and or insufficient x D rinsings or frequent preventive use of the highly effective disinfection protocol with chloramine T solution x D for hours outside its application requirement instead of the recommended weekly routine disinfection with x D chloramine T for min increases the material stress and potential degradation of HU tank fittings | Replacement of the HU 35 tank fittings and the tank screw cap with the new qualified resistant material on the next preventive maintenance | 2025-07-08 17:18
Saudi Mais Co., for Medical Products | Blood Transfusion Set | 887 | - | improper assembly | recalling the product | 2025-07-09 13:21
Covidien llc | Nellcor Bedside SpO2 Patient Monitoring System | 900 | - | Through escalation of complaints related to deaths it was reported that the PM N did not alarm when used in a home care setting Through investigation Medtronic employees have found that the monitors were not set into Homecare modes Homecare Mode simplifies device operation and prevents users from making changes to patient mode and alarm setting Remaining in standard mode allows for user to adjust the alarm limits and patient mode x D Additional investigation was completed looking into the two devices that are similar in design and functionality to PM N Libra and Birdie Birdie is applicable in Japan only Through investigation it was found there were additional complaints related to No Alarm not hearing the alarm go off in the hospital and homecare settings x D The devices that were returned and investigated in relation to the No alarm complaints could not be confirmed as all devices passed functional testing Additionally as documented in the complaint investigation files device logs proved that all appropriate alarms were triggered as expected The combination of knowing the devices triggered alarms and the physical testing of the devices proving that alarms annunciate as designed the conclusion is that the devices are performing as intended x D | Advice on actions to be taken by the distributor and the user: Actions prescribing clinicians should take • Be sure you read and understand the information contained in this Notice. • Communicate the information from this notice to your home use patients and their caregivers. Suggested language to use with your patients and their caregivers: For the safety of the patient, be sure the monitor’s alarm volume is set to a level that can be heard throughout your home at all times of day or night. Respond to all alarms because whenever an alarm occurs the patient's health may be at risk. Notify your clinician if alarms occur. Whenever you reconnect the sensor to the patient or the monitor check that the oxygen level and heart rate are acceptable. Read and follow the Home Use Guide provided to you. This guide provides basic information for safe use of your monitoring system. • Consult the Operator’s and Service manuals to set alarm configurations, volume and confirm that the Covidien Nellcor™ Bedside SpO2 Patient Monitoring System is appropriately set for the patient being monitored and for the patient’s care environment. • Confirm that any model PM100N device in use outside of a hospital or hospital-type environment is set to Homecare Mode. Contact your Medtronic Representative to obtain a four-digit password required to set a monitor to Homecare Mode. Actions facility risk managers should take • Pass on and post this notice for all those who need to be aware within your organization and to any organization where the product has been transferred or distributed. • Please complete and return the enclosed Customer Acknowledgement Form even if you do not have inventory. | 2025-07-10 15:10
Boston Scientific | Vercise Genus | 73 | - | In August Boston Scientific reminded DBS customers to follow the steps outlined in the labeling IFU x D Instructions for Use to implant Vercise DBS IPGs within a subcutaneous pocket Boston Scientific is now x D providing an update to the previous communication An IFU revision is being implemented to include an x D additional warning statement x D Warning Submuscular implant in the pectoral region i e subpectoral may expose the Vercise Genus x D rechargeable IPG to frequent muscle tension forces which may cause component damage feedthrough x D wire s resulting in device malfunction that could lead to adverse events including explant of the IPG | IFU update | 2025-07-11 18:05
Boston Scientific Corporation | SureTek Burr Hole Cover Kit | 73 | - | Boston Scientific is initiating a Product Correction IFU update to address difficulties while locking x D SureTekTM Burr Hole Covers BHC to secure Deep Brain Stimulation DBS Leads part of the Vercise x D Deep Brain Stimulation System In some cases resistance has been encountered while closing the x D Retaining Clip locking mechanism to securely hold the lead in place | IFU update | 2025-07-11 18:10
.Siemens Healthcare Diagnostics Inc | Atellica® CH D_TBil | 26880 | - | Issue Description x D Siemens Healthineers has confirmed through complaint investigation that the lot calibration claim of days per the Instructions for Use may not be met when using the Atellica CH D TBil assay due to insufficient reagent P light protection Siemens investigation indicated that light exposure of the reagent could potentially lead to falsely elevated or depressed results upon switching reagent wells or packs and a potential for differing performance between reagent wells of the same pack The direction of result bias depends on the pack used for lot calibration and the relative light exposure of subsequent packs to which the lot calibration is applied All in date Atellica CH D TBil lots used on the Atellica CH and Atellica CI analyzers are impacted by this issue x D Note Once the reagent pack is loaded onto the analyzer the light shielded compartment prevents any further light exposure A design change will be performed to change to opaque black packs to address light sensitivity x D x D Kindly review the attached Urgent Field Safety Notice | Customer Actions • Perform the instructions provided below: 1. All D_TBil reagents must be stored in their outer kit box and out of direct light. ▪ If stored in a walk-in cold room, it is important to keep lighting off when not in use. ▪ Consider storing kit boxes on lower shelving and away from direct lighting. ▪ If possible, store kit boxes in a refrigerator with an opaque door or within a secondary opaque container. If stored in a refrigerator with a glass door, measures to block light from coming through the glass door should be implemented. 2. Update the following QC settings of the Atellica software (See appendix for instruction): a) Enable QC on Pack Change By Assay Type for CH in order for QC to be processed when switching between wells. Note: This setting enables QC on Pack Change for all CH assays. All QC Levels for CH assays for which QC on Pack Change is not required should be deselected in the QC Master List. b) Enable Patient QC Flagging in order for impacted patient results to be identified after a QC failure. Note: All impacted patient results will have a “QC Fail” flag. Ensure that a result with a flag of “QC Fail” is held for review so that it can be rerun after passing QC. 3. Upon QC failure, perform a Pack Calibration on the impacted well. Note: Ensure that Pack Calibration is performed on the intended well that failed QC. Alternatively, a user may perform a Pack Calibration for each reagent well proactively. Refer to Atellica Solution and Atellica CI Analyzer Online Help for instructions. 4. Run QC. 5. Once acceptable QC has been obtained, repeat all D_TBil testing for all patient results flagged with “QC Fail”. • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. • Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached Urgent Field Safety Notice | 2025-07-14 08:42
TOP GLOVE SDN. BHD. | OptiTect Latex Gloves | - | - | The medical device does not have Medical Devices Marketing Authorization MDMA which violates the medical device law Article | • Stop using the affected products. • Return the affected products to the distributor. Manufacturer action: Removal. | 2025-07-15 08:26
Getinge | Corin Operating Tables | 7 | - | Under certain conditions such as the simultaneous use of the operating table with other medical devices that may cause interference including high frequency HF surgical equipment we have identified a potential issue that could prevent the device from performing as intended If the issue occurs the Corin operating table XY does not respond to any command It is impossible to actuate any movement of the Corin operating table neither with the universal remote control nor the override panel Please refer to the attached FSN for more details | A software update that will mitigate the described issue has been developed. Getinge will initiate an immediate field action of all affected device units. | 2025-07-15 18:11
Ortho Clinical Diagnostics Inc. | Ca Slides | 111 | - | The purpose of this notification is to inform you that QuidelOrtho has confirmed an issue involving VITROS Chemistry Products Ca Slides Customers may experience an increased incidence of calibration failures when attempting to calibrate VITROS Ca Slides on VITROS XT Chemistry Systems and VITROS XT Integrated Systems | • Upon receipt, load ADD DRV 6338 on your VITROS XT 3400 and/or VITROS XT 7600 System(s). • Complete the enclosed Confirmation of Receipt form no later than Month DD, 2025. • Save this notification with your User Documentation or post this notification by each VITROS XT 3400/VITROS XT 7600 System in your laboratory until the issue has been resolved. • Please forward this notification if the affected product was distributed outside of your facility. • If your laboratory has experienced the issue with this product and you have not already done so, please report the occurrence to your local Global Services Organization. | 2025-07-16 16:32
TOP GLOVE SDN. BHD. | OPTITECT LATEX GLOVES | 50 | - | The medical device does not have Medical Devices Marketing Authorization MDMA | Withdraw the affected device from the market. | 2025-07-17 14:30
Draeger Medical, Inc. | VentStar Helix | 10 | - | Loosening the tubing connection can lead to termination of ventilation in the event of use Contamination of the tubing system can lead to health problems | 1. Dispatch of an important safety notice to all customers The safety notice is sent to the recipients named in the location list. If recipients do not respond in the first communication, remind them two more times. Then send a closing letter and archive it in TrackWise. 2. Delivery of replacement if requested by the customer | 2025-07-21 14:15
Dexcom, Inc. | Dexcom G7 CGM App | 2524 | - | Dexcom identified an issue in the Dexcom G iOS and Android Apps versions and earlier and Dexcom ONE iOS and Android Apps versions and earlier where when the transmitter of an associated wearable experiences a hardware or firmware failure upon the next communication cycle with the app the transmitter sends a transmitterFailed error message to the app However the app does not provide the expected Sensor Failed alert to inform the user of the wearable failure The app instead ends the session and displays the Start Sensor screen This error is due to the software development kit SDK alert type not being mapped to any of the G display alerts so the expected Alert model is not triggered when the G wearable s transmitter enters into a hardware firmware derived failure state x D x D Note The watchOS app Apple is also affected by this issue It is installed along with the iOS phone app and the versions used of each are the same for each user The watchOS app cannot be installed independently from the iOS phone app | Manufacturer action: Dexcom has access to contact all remaining users of the affected App versions through deployment of in-App notifications. Dexcom intends to utilize this in-App messaging and is targeting to notify all remaining users of the affected G7 CGM App versions and Dexcom ONE+ App version of this issue at a minimum weekly beginning on 23 July 2025. This message will also notify users that the affected app versions will be blocked 1 month after receipt. Users must either update upon receiving this message or may choose to defer the update until closer to the blocking date. An in-app obsolescence message is targeted for 20 August 2025 and will be utilized to block the affected app Android and iOS CGM Apps. Dexcom has the ability to track user app versions to confirm that the affected users have upgraded their App or discontinued use of the App. The blocking completed via the in-app obsolescence message will ensure that all users are no longer able to use the affected version of the App. Mandatory Customer Action for (Dexcom G7 app version 2.8 or earlier, Dexcom ONE+ app version 1.4 or earlier): Upgrading is mandatory and you will not be able to use app version 1.4 and 2.8 or earlier after August 20, 2025. Update your app now to continue using the app: 1. Tap Update App to go to the app store 2. Install the latest Dexcom ONE+ and G7 app version 3. Open the Dexcom ONE+ and G7 app If you need assistance, contact Technical Support. US: 1-844-478-1600 Global: Find local Technical Support contact information at Dexcom.com Adverse reactions or quality problems experienced with the use of these products should be reported to the National Competent Authority in your area. | 2025-07-21 17:30
Dexcom, Inc. | Dexcom ONE+ | 2524 | - | Dexcom identified an issue in the Dexcom G iOS and Android Apps versions and earlier and Dexcom ONE iOS and Android Apps versions and earlier where when the transmitter of an associated wearable experiences a hardware or firmware failure upon the next communication cycle with the app the transmitter sends a transmitterFailed error message to the app However the app does not provide the expected Sensor Failed alert to inform the user of the wearable failure The app instead ends the session and displays the Start Sensor screen This error is due to the software development kit SDK alert type not being mapped to any of the G display alerts so the expected Alert model is not triggered when the G wearable s transmitter enters into a hardware firmware derived failure state x D x D Note The watchOS app Apple is also affected by this issue It is installed along with the iOS phone app and the versions used of each are the same for each user The watchOS app cannot be installed independently from the iOS phone app | Manufacturer action: Dexcom has access to contact all remaining users of the affected App versions through deployment of in-App notifications. Dexcom intends to utilize this in-App messaging and is targeting to notify all remaining users of the affected G7 CGM App versions and Dexcom ONE+ App version of this issue at a minimum weekly beginning on 23 July 2025. This message will also notify users that the affected app versions will be blocked 1 month after receipt. Users must either update upon receiving this message or may choose to defer the update until closer to the blocking date. An in-app obsolescence message is targeted for 20 August 2025 and will be utilized to block the affected app Android and iOS CGM Apps. Dexcom has the ability to track user app versions to confirm that the affected users have upgraded their App or discontinued use of the App. The blocking completed via the in-app obsolescence message will ensure that all users are no longer able to use the affected version of the App. Mandatory Customer Action for (Dexcom G7 app version 2.8 or earlier, Dexcom ONE+ app version 1.4 or earlier): Upgrading is mandatory and you will not be able to use app version 1.4 and 2.8 or earlier after August 20, 2025. Update your app now to continue using the app: 1. Tap Update App to go to the app store 2. Install the latest Dexcom ONE+ and G7 app version 3. Open the Dexcom ONE+ and G7 app If you need assistance, contact Technical Support. US: 1-844-478-1600 Global: Find local Technical Support contact information at Dexcom.com Adverse reactions or quality problems experienced with the use of these products should be reported to the National Competent Authority in your area. | 2025-07-21 17:30
Spacelabs Healthcare Inc. | Ultraview DM3 Monitor and Spacelabs DM4 | 75 | - | This product labelling does not adequately provide warning about the risks associated with the use of damaged batteries or batteries past the intended use life x D | -The Operations Manual and Service Manual for the DM3 and DM4 have been updated to convey the risks associated with continuing to use the lithium batteries past the specified use life and to clarify that batteries older than two years, damaged, or that are no longer holding a charge must be replaced -Users must download current versions of the Operations Manual and Service Manual and dispose of any previous copies. The updated product labeling is available via the Spacelabs user manual site at https://manuals.spacelabshealthcare.com/. Hospital's Biomedical engineering must remove and replace any batteries that have been in use longer than two years, are damaged, or are no longer holding a charge. Contact Spacelabs Service to arrange for battery replacement. Use only manufacturer-supplied batteries. Hospital's Biomedical engineering must arrange preventive maintenance to replace the batteries upon reaching the conclusion of their two-year use life. | 2025-07-21 18:46
Hologic Inc. | Affirm Prone Biopsy System | 9 | - | Please find the attached report | Please find the attached report | 2025-07-22 13:19
SHIMADZU | RADspeed PRO | 41 | - | Please find the attached report | Please find the attached report | 2025-07-22 15:26
SHIMADZU | RADspeed PRO | 176 | - | Please find the attached letter | Please find the attached letter | 2025-07-22 16:24
Spacelabs Healthcare Inc | Spacelabs Reusable Cloth ABP Cuffs &UltraCheck Reusable Nylon ABP Cuffs | 36 | - | The plasticizer in the hose may degrade and leach over time causing it to become sticky and then hard and brittle | 1-Be aware of potential device degradation and avoid excessive use, cleaning, or exposure to cleaning agents. If the tubing becomes sticky or inflexible, please discard it. FSN (acknowledgement form) should be signed by the person in charge of the hospital. 2-All affected products will now be labeled with a shelf-life or manufactured with tubing that does not require a shelf life.. | 2025-07-22 20:18
- | Massage table Elemental , Massage table Elemental, Safari Massage table (Jaguar) | - | - | Counterfeit Devices which violates the medical device law Article | Withdraw the affected product from the market. | 2025-07-24 08:53
B.Braun Medical SAS | Celsite® Babyport SET PUR 4,5F IV - Implantable access port system, Venous | 117 | - | During the course of Post Market surveillance activities B Braun Medical identifi ed that the tear away introducer supplied in the kits of the above listed batches of Celsite Babyport access port may potentially present a defect at the level of the distal tip | Recall IMMEDIATELY the aff ected products from the customers who received the aff ected batches. | 2025-07-24 16:44
Eppendorf | DNA LoBind® Tubes 1.5mL, 250 Tubes | 28 | - | Tubes of the affected lots have a residual plastic spike from the injection mold process The spike can cut through a glove used as personal protective equipment PPE in the laboratory posing a risk for contamination or cause small cuts in the skin of the user | 1. Please do not use the tubes from the affected lot to avoid any risk during use. 2. Make pictures of the packaging with visible product label and LOT number. 3. Contact your supplier of the respective products to request free replacement or refund. 4. Dispose the affected products. Please adhere to the designated disposal procedures based on your laboratory classification. Remedies for customers | 2025-07-27 13:13
BD Switzerland Sàrl | BD Alaris neXus GP | 500 | - | BD has identified through customer feedback that the specific BD Alaris neXus GP pumps listed in attachment are not able to connect to Wi Fi which may be required for data transfer with the Alaris Communication Engine v x ACE and later versions If the Alaris Gateway Workstation AGW is present and the pump is docked the data transfer performs as intended | Clinical Actions • Ensure your pump, if listed in Attachment 1, is made available for the software upgrade as requested by your service organisation. o The following link may also be used to determine if a pump is impacted: https://bdx.my.site.com/CC360/s/impactedproducts?language=en_US&rn=MMS-25-5323%20GLOBAL Actions for Biomedical Engineers / Service Organisation • Locate the specific BD Alaris™ neXus GP Infusion pumps listed in Attachment 1 for the software upgrade. o The following link may also be used to determine if a pump is impacted: https://bdx.my.site.com/CC360/s/impactedproducts?language=en_US&rn=MMS-25-5323%20GLOBAL • The software upgrade may be performed by BD technicians or by qualified Biomedical Engineers / Service Organisations already trained to service the BD Alaris™ neXus GP Infusion pumps. • Qualified Biomedical Engineers and Service Organisations should contact BD (local BD representative or local BD office for service support) to receive the necessary materials and instructions if they wish to conduct the necessary software upgrade Actions TAKEN by BD: BD has updated the software of the pumps to ensure future shipments are unaffected by this issue. Actions TO BE taken by BD BD is investigating the root cause of this issue to avoid recurrence. Customer Actions: • Review the information in Attachment 1 to determine if the BD Alaris neXus GP pump(s) in your possession are impacted. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 28th August 2025. • Circulate this notification to all those who need to be aware within your organisation or to any organisation where the potentially affected product has been transferred. • If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: • Review the information in Attachment 1 and determine if the BD Alaris neXus GP pump(s) in your possession are impacted. • Identify the facilities where you have distributed affected product and notify them immediately of this notification. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 28th August 2025. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. | 2025-07-27 14:09
Baxter Healthcare SA | PrisMax | 142 | - | Vantive is issuing a Correction for the PrisMax Systems and TherMax Blood Warmer Units listed below due to the following three issues x D x D PrisMax System alarm T Air Detected in Prime x D PrisMax System alarm T Blood Leak Detected may be raised when no blood leak is present or the operator may have difficulty normalizing the Blood Leak Detector BLD leading to PrisMax System alarms T BLD Normalize Failed T Normalization Failed or T BLD Self Test Failure x D Thermax System may be unable to detect the presence of a bag on the Thermax Blood Warmer Unit leading to difficulty setting up therapy or PrisMax System alarm T Thermax Disposable Not Inserted x D | Vantive is currently investigating these issues and will be correcting the impacted devices. | 2025-07-27 14:20
Olympus | HF units ESG-410 | 10 | - | Olympus received forty three complaints between July to March from users where the ESG Electrosurgical Generator displayed error message E System Error followed by either automatic single reboots of the device or continuous reboot loops Users reported that these reboots occurred both during and prior to clinical use An investigation by Olympus identified that this issue is caused by a previous change to an electronic component This component change impacted an internal performance check parameter which could then result in the error message being displayed and an automatic system reboot In some cases the system will be in a continuous reboot None of the forty three complaints resulted in an adverse event x D An Olympus representative will reach out to you to arrange for a device update which resolves the cause of this issue x D | 1. Examine your inventory for affected serial numbers (refer to Table 1 for affected serial number information). 2. Customers may continue to use the device whilst they are waiting for the update, ensuring all users follow the product Instructions for Use (IFU). This includes the instruction to prepare a spare electrosurgical generator, or an alternative procedure, to avoid interruption due to an unexpected electrosurgical generator failure during the procedure. If any problems are encountered with the device, including error messages and automatic reboots, please contact Olympus. 3. Olympus representative will reach out to you to arrange a mutually convenient time to update the affected device at no charge. | 2025-07-27 14:23
EPPENDORF | DNA LoBind® Tubes 1.5mL, 250 Tubes | 28 | - | Tubes of the affected lots have a residual plastic spike from the injection mold process The spike can cut through a glove used as personal protective equipment PPE in the laboratory posing a risk for contamination or cause small cuts in the skin of the user | 1. Please do not use the tubes from the affected lot to avoid any risk during use. 2. Make pictures of the packaging with visible product label and LOT number. 3. Contact your supplier of the respective products to request free replacement or refund. 4. Dispose the affected products. Please adhere to the designated disposal procedures based on your laboratory classification. | 2025-07-27 15:18
Biomerieux | BIOFIRE® FILMARRAY® Pneumonia Panel and BIOFIRE® FILMARRAY® Pneumonia Panel plus | 17 | - | The update to the Pouch Module software modifies the melt range of the Coronavirus assay to exclude most instances of a non specific amplicon generated by cross reactivity with human genomic DNA hgDNA and reports a Not Detected Coronavirus result | update the kit SW | 2025-07-28 15:38
GE Healthcare | Enterprise Archive | 1 | - | GE HealthCare has identified that imaging studies or images within a study may be associated with an incorrect patient following migration of historical studies to the Enterprise Archive EA in True PACS If this occurs it could potentially result in misdiagnosis or a delay in care x D x D There have been no injuries reported to GE HealthCare as a result of this issue x D | GE HealthCare Service to confirm if site impacted then use steps to identify impacted studies and fix utilizing new FMI manual. | 2025-07-28 16:07
Boston Scientific Corporation | RELIANCE ePTFE | 1598 | - | An increased potential to exhibit a pattern of gradual rise in LVSI which is associated with calcification of the shocking coil s | Advisory- Customer notification | 2025-07-30 13:41
Immunotech SAS ,a Beckman Coulter Company | Access hsTnI | 784 | - | In remote circumstances if a sample containing high cardiac troponin cTnI is tested using an Access assay other than Access hsTnI cTnI carryover may occur if the next test performed using the same probe on that instrument is Access hsTnI on a different sample x D Following additional feedback Beckman Coulter is providing representative data for the magnitude of carryover observed in internal studies | • Follow your established laboratory protocols for analyzing and retesting discrepant samples if an observed Access hsTnI test result does not align with the patient’s clinical presentation. • Beckman Coulter recommends reviewing the content of this letter with your laboratory and/or medical director to determine appropriate next steps. | 2025-07-30 23:41
Philips Medical Systems Nederland B.V. | Allura Xper FD10/10 | 12 | - | The Phase UPS is an optional component that can be installed in Allura and Azurion systems to support a controlled shutdown of the PCs the computers that control the system in the event of a mains power failure The Phase UPS provides temporary power to the PCs allowing the system to save the data of the last acquired run before the PCs shut down x D Philips has identified that some internal components of the Phase UPS may fail When this occurs there will be a complete loss of power to the system causing the system to shut down and or not start up The Phase UPS system may fail without warning and prior to any mains power failure | Philips will bypass the 1 Phase UPS in the affected systems (disconnecting all connections of the 1 Phase UPS but not removing it physically). This will render the 1 Phase UPS inactive, preventing the occurrence of the above-described issue. Not having the 1 Phase UPS connected will not affect system usage during normal operation. Once the 1 Phase UPS is bypassed, controlled shutdown and data backup in the event of a hospital mains power interruption will no longer be available, which may cause the last acquired run to be lost when the power failure occurs during the acquisition or shortly after the run was acquired. Data loss from the last run on Azurion systems is not expected, as these systems are equipped with solid state drives (SSD). Philips will contact all affected customers to schedule the execution of this activity (reference: FCO72200594). As of the date of this Urgent Field Safety Notice, Philips expects to start bypassing the 1 Phase UPS by Q4 2025. | 2025-07-31 13:31
Philips Medical Systems Nederland B.V. | Allura Xper FD20 | 12 | - | The Phase UPS is an optional component that can be installed in Allura and Azurion systems to support a controlled shutdown of the PCs the computers that control the system in the event of a mains power failure The Phase UPS provides temporary power to the PCs allowing the system to save the data of the last acquired run before the PCs shut down x D Philips has identified that some internal components of the Phase UPS may fail When this occurs there will be a complete loss of power to the system causing the system to shut down and or not start up The Phase UPS system may fail without warning and prior to any mains power failure | Philips will bypass the 1 Phase UPS in the affected systems (disconnecting all connections of the 1 Phase UPS but not removing it physically). This will render the 1 Phase UPS inactive, preventing the occurrence of the above-described issue. Not having the 1 Phase UPS connected will not affect system usage during normal operation. Once the 1 Phase UPS is bypassed, controlled shutdown and data backup in the event of a hospital mains power interruption will no longer be available, which may cause the last acquired run to be lost when the power failure occurs during the acquisition or shortly after the run was acquired. Data loss from the last run on Azurion systems is not expected, as these systems are equipped with solid state drives (SSD). Philips will contact all affected customers to schedule the execution of this activity (reference: FCO72200594). As of the date of this Urgent Field Safety Notice, Philips expects to start bypassing the 1 Phase UPS by Q4 2025. | 2025-07-31 13:31
Philips Medical Systems Nederland B.V. | AZURION 7 M20 | 12 | - | The Phase UPS is an optional component that can be installed in Allura and Azurion systems to support a controlled shutdown of the PCs the computers that control the system in the event of a mains power failure The Phase UPS provides temporary power to the PCs allowing the system to save the data of the last acquired run before the PCs shut down x D Philips has identified that some internal components of the Phase UPS may fail When this occurs there will be a complete loss of power to the system causing the system to shut down and or not start up The Phase UPS system may fail without warning and prior to any mains power failure | Philips will bypass the 1 Phase UPS in the affected systems (disconnecting all connections of the 1 Phase UPS but not removing it physically). This will render the 1 Phase UPS inactive, preventing the occurrence of the above-described issue. Not having the 1 Phase UPS connected will not affect system usage during normal operation. Once the 1 Phase UPS is bypassed, controlled shutdown and data backup in the event of a hospital mains power interruption will no longer be available, which may cause the last acquired run to be lost when the power failure occurs during the acquisition or shortly after the run was acquired. Data loss from the last run on Azurion systems is not expected, as these systems are equipped with solid state drives (SSD). Philips will contact all affected customers to schedule the execution of this activity (reference: FCO72200594). As of the date of this Urgent Field Safety Notice, Philips expects to start bypassing the 1 Phase UPS by Q4 2025. | 2025-07-31 13:31
Philips Medical Systems Nederland B.V. | Allura Xper FD10 | 12 | - | The Phase UPS is an optional component that can be installed in Allura and Azurion systems to support a controlled shutdown of the PCs the computers that control the system in the event of a mains power failure The Phase UPS provides temporary power to the PCs allowing the system to save the data of the last acquired run before the PCs shut down x D Philips has identified that some internal components of the Phase UPS may fail When this occurs there will be a complete loss of power to the system causing the system to shut down and or not start up The Phase UPS system may fail without warning and prior to any mains power failure | Philips will bypass the 1 Phase UPS in the affected systems (disconnecting all connections of the 1 Phase UPS but not removing it physically). This will render the 1 Phase UPS inactive, preventing the occurrence of the above-described issue. Not having the 1 Phase UPS connected will not affect system usage during normal operation. Once the 1 Phase UPS is bypassed, controlled shutdown and data backup in the event of a hospital mains power interruption will no longer be available, which may cause the last acquired run to be lost when the power failure occurs during the acquisition or shortly after the run was acquired. Data loss from the last run on Azurion systems is not expected, as these systems are equipped with solid state drives (SSD). Philips will contact all affected customers to schedule the execution of this activity (reference: FCO72200594). As of the date of this Urgent Field Safety Notice, Philips expects to start bypassing the 1 Phase UPS by Q4 2025. | 2025-07-31 13:31
Philips Medical Systems Nederland B.V. | Azurion 5 M20 | 12 | - | The Phase UPS is an optional component that can be installed in Allura and Azurion systems to support a controlled shutdown of the PCs the computers that control the system in the event of a mains power failure The Phase UPS provides temporary power to the PCs allowing the system to save the data of the last acquired run before the PCs shut down x D Philips has identified that some internal components of the Phase UPS may fail When this occurs there will be a complete loss of power to the system causing the system to shut down and or not start up The Phase UPS system may fail without warning and prior to any mains power failure | Philips will bypass the 1 Phase UPS in the affected systems (disconnecting all connections of the 1 Phase UPS but not removing it physically). This will render the 1 Phase UPS inactive, preventing the occurrence of the above-described issue. Not having the 1 Phase UPS connected will not affect system usage during normal operation. Once the 1 Phase UPS is bypassed, controlled shutdown and data backup in the event of a hospital mains power interruption will no longer be available, which may cause the last acquired run to be lost when the power failure occurs during the acquisition or shortly after the run was acquired. Data loss from the last run on Azurion systems is not expected, as these systems are equipped with solid state drives (SSD). Philips will contact all affected customers to schedule the execution of this activity (reference: FCO72200594). As of the date of this Urgent Field Safety Notice, Philips expects to start bypassing the 1 Phase UPS by Q4 2025. | 2025-07-31 13:31
MI GWANG CONTACT LENS CO., LTD. | Accez Daily Color Contact Lenses | 5517 | - | Non conformity to the technical standard specification ISO Annex Water Content | • Provide a list of all batches of the product supplied to the Saudi market in the last two years,including their respective quantities. • Withdraw three (3) samples from three (3) different lot numbers (total of nine (9) samples) from the local market. • Re-test the collected samples at a certified reference laboratory approved by the SFDA | 2025-08-03 13:40
Beckman Coulter, Inc. | DxC 500i Clinical Analyzer DxC 500i Access 2 Immunoasssay Module | 10 | - | Beckman Coulter has determined that on the DxC i Clinical Analyzer when ordering a Clinical Chemistry CC combination test simultaneously with an Immunoassay IA test if the IA test is processed between the Clinical Chemistry constituent tests the sample status of the Clinical Chemistry test will remain In Progress with no error events | • This issue only impacts software versions 1.3.3 and 1.3.4. To check your software version, select Menu > About page. • For DxC 500i Clinical Analyzers running software version 1.3.3 or 1.3.4, if this issue occurs, you will need to place a new order for the CC combination test, o Select Home > Sample List task indicator, then select Active in the upper left corner of the Sample List to display samples “In Progress”. o Program the CC combination tests for these “In Progress” samples using a new rack and Sample ID, apply the new barcode, then load the rack onto the analyzer for processing. o Set aside the rack with “In Progress” samples and contact your Beckman Coulter representative to clear them. • Beckman Coulter recommends sharing the content of this letter with your laboratory and/or Medical Director. | 2025-08-03 15:51
Boston Scientific Corp. | WATCHMAN TruSeal™ Access System, WATCHMAN FXD Curve™ Access System | 336 | - | This notification is in response to observations of air embolism complaints received by Boston Scientific related to WATCHMAN procedures The global rate of reported air embolism under any sedation type during the WATCHMAN procedure is associated with death in of cases After a comprehensive investigation it was concluded that these complaints are not associated with the design or manufacture of the WATCHMAN system Based on the review of the complaints Boston Scientific has identified a higher likelihood of air embolism events when procedures are performed without positive pressure controlled ventilation | Customer notification. The WATCHMAN Access Systems Instructions for Use (IFUs) and WATCHMAN physician training will be updated to emphasize instructions related to Access System air management. This update will strengthen the information provided to clinicians regarding the potential for air embolism during WATCHMAN procedures performed under conscious or deep sedation and provide potential mitigation strategies. | 2025-08-04 10:27
NIPRO CORPORATION | Single Patient Dialysis Machine | 641 | - | progressive mechanical degradation of pump motor | replacement of the affected Magnet Pumps | 2025-08-04 12:36
Philips Medical Systems Nederland B.V. | Allura Xper FD20 | 2 | - | Philips has identified that a component air baffle may be potentially missing in some Nehalem FlexVision PCs that are part of the Allura Xper series x D The air baffle directs part of the airflow inside the PC When it is not installed the temperature of the frame grabber card of the FlexVision PC may be outside the specified temperature range which may lead to an increased probability of failure of the frame grabber card If this occurs no image will be shown in the FlexVision monitor of the examination room | Philips will be inspecting affected Nehalem Flexvision-2 PCs and installing the air baffle if it is confirmed that this component is missing. | 2025-08-04 13:05
Aesculap AG | SKULL TRACTION TONG LARGE | 20 | - | It was identified that the Instructions for Use IFU did not include essential information x D regarding contraindications and application details As a result this critical information was not x D made available to the end user | The Instruction for Use (IFU) TA009782 has been updated (Change No. AE0065363) accordingly with the following changes in content: 2.1.2 Intended use The skull traction tong is used for odontoid fractures with non-lethal dislocation. It provides temporary immobilization and/or reduction in the transitional region of an unstable cervical spinal injury, as well as for de-compression of the spinal cord or nerve roots secondary to spinal misalignment. 2.1.4 Absolute contraindications - Patients with open fontanelles - Malformations or known reduction of the bone tissue, in particular temporal bone malformations 2.1.5 Relative contraindications The following conditions, individual or combined, can lead to delayed healing or compromise the success of the operation: - Medical or surgical conditions (e.g., comorbidities) that could hinder the success of the operation. - Children less than 5 years of age - Patients with local sepsis - Patients who are not awake, alert, and cooperative - Skull fractures and particular cervical spine injuries (in particular unstable injuries rostral to the target level, traumatic disc herniations and definitive, complete ligamentous injuries at any cervical spine level). In the presence of relative contraindications, the user decides individually regarding the use of the product. | 2025-08-04 16:06
Ortho Clinical Diagnostics Inc. | LAC Slides | 170 | - | The purpose of this notification is to inform you that QuidelOrtho has confirmed an issue x D involving an increased incidence of calibration failures when attempting to calibrate x D VITROS LAC Slides on VITROS XT Chemistry Systems and VITROS XT x D Integrated Systems | • Upon receipt, load ADD DRV 6339 (or above) on your VITROS XT 3400 and/or VITROS XT 7600 System(s). • Complete the enclosed Confirmation of Receipt form no later than Month DD, 2025. • Save this notification with your User Documentation or post this notification by each VITROS XT 3400/VITROS XT 7600 System in your laboratory until the issue has been resolved. • Please forward this notification if the affected product was distributed outside of your facility. • If your laboratory has experienced the issue with this product and you have not already done so, please report the occurrence to your local Global Services Organization. | 2025-08-06 11:29
Agilent | FLEX Monoclonal Mouse Anti-Human CD20cy, Clone L26, Readyto-Use (Dako Omnis) | 178 | - | Weak staining has been observed in a few cases where GA was used to stain CLL SLL tissue samples x D for CD This weak staining has not been observed in the majority of cases where GA was used to x D test CLL SLL tissue As a reminder the diagnosis of CLL SLL is made in the context of the histologic x D appearance by Hematoxylin and Eosin H E staining with the aid of a panel of immunohistochemical x D markers routinely including other B cell markers such as PAX or CD ancillary studies e g flow x D cytometry on peripheral blood or bone marrow and clinical information | Follow the instructions provided on the next page(s) in detail to assure timely and compliant implementation of this FSCA. An overview of timeframes for the different actions can be found after the instructions under “Timeframes” | 2025-08-06 15:51
Shimadzu Corporation | FDR Visionary Suite | 2 | - | It was found that the contacts of the electromagnetic contactors used inside the cabinet of the X ray high voltage generator were welded and the current might continue to be supplied to the power supply circuit even if the power of the device is cut off As a result it cannot be denied that the circuit in question overheats and damages the surrounding parts It may also cause smoke x D x D x D | Please pay attention to the following points until the work is completed.  For safety reasons, please turn off the breaker of the distribution board to which the power cord of the unit is connected after turning off the power of the unit according to the operation manual.  Do not forget to turn off the breaker of the distribution board until the replacement of the contactor is completed. 5．Actions planned by Shimadzu to correct the problem We are informing customers of this issue through this Urgent Field Safety Notice. Additionally, Therefore, we will take measures to prevent the current from continuing to flow even if the electromagnetic contactor is welded. This Urgent Field Safety Notice letter has been reported to the appropriate Competent Authorities. | 2025-08-10 12:34
Tandem Diabetes Care | t:slim X2 with Basal-IQ technology and t:slim X2 with Control-IQ technology | 687 | - | Reference CAPA two root causes were identified x D The creation of a point of stress concentration with a sharp bend in the wire when the wire is not routed through the gap in the frame and is instead crushed between the speaker housing and cap creates a point of failure when the speaker is cycled over its x D lifetime leading in the wire resulting in a malfunction on east electronics speakers x D The application of excessive epoxy renders the speaker incapable of fitting into the cavity in the housing when it is pressed into the housing by the operator or when the PCBA is installed on top of it the wires are crushed creating stress on the wire | Tandem Diabetes Care will be releasing a software update aimed at enhancing early detection of speaker failure. This update will also introduce persistent vibration alerts to help reduce potential safety risk. VitalAire Arabia (Distributor) will notify affected pump users when the software update becomes available and request pump users to complete the software update of their insulin pump. | 2025-08-10 13:33
Randox Laboratories Ltd | Immunoassay Premium Level 1 | 1424 | - | Randox Laboratories can confirm that the analyte Thyroid Stimulating Hormone TSH in x D Immunoassay Premium Level lots EC and EC in some methods analysers is x D recovering high outside range Please contact technical services randox com if you obtain x D a high TSH result | Action to be taken: Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2025-08-10 15:01
Randox Laboratories Ltd | Immunoassay Premium Level 1 | 1424 | - | Randox Laboratories can confirm that the analyte Thyroid Stimulating Hormone TSH in x D Immunoassay Premium Level lots EC and EC in some methods analysers is x D recovering high outside range Please contact technical services randox com if you obtain x D a high TSH result | Action to be taken: Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. | 2025-08-10 15:01
Implantcast GmbH | MUTARS® Tumor and Revision System | 78 | - | The results of the simulation have prompted implantcast GmbH to stipulate limits of use for some products and sizes as a precautionary measure This decision is not due to a safety problem with the implants Since the aforementioned items were placed on the market no abnormalities in relation to fractures have been found x D x D implantcast GmbH has decided to mark separately the weight restricted stems in the surgical techniques MUTARS proximal femur and MUTARS proximal femur revision in the chapter Implants The system overview also details the individual weight restrictions for the reconstruction lengths A warning has also been added to the IFU x D | Corrective action done by informing the user | 2025-08-10 17:24
Beckman Coulter Inc. | Unicel DxH 900 Automated Hematology Analyzer | 34 | - | Beckman Coulter has identified a potential for erroneously high Hemoglobin HGB results on samples with elevated White Blood Cell WBC counts This occurs on UniCel DxH Series instruments configured with the LED HGB Photometer x D This issue was identified following a complaint where a patient sample with a WBC count of approximately x cells L generated a falsely elevated HGB by g dL and MCHC by g dL x D Falsely elevated HGB results may be observed in specimens with WBC counts x cells L x D The HGB bias increases proportionally with rising WBC concentrations | •For specimens with a WBC count > 30 x 103 and ≤ 90 x 103 cells/μL, any of the following actions may be taken to obtain accurate HGB results: o Process specimens on an instrument not impacted by this recall. o Process specimens using a validated alternative method for HGB and associated parameters. o Dilute specimens to below 30 x 10³ cells/μL using your laboratory protocol or the DxH predilute function. Refer to your instrument’s Instructions for Use (IFU), Chapter 5, Predilute Panel (PREDIx5) for detailed steps. •For specimens with a WBC count > 90 x 103 cells/μL: o Process specimens on an instrument not impacted by this recall. o Process specimens using a validated alternative method for HGB and associated parameters. •To identify WBC counts > 30 x 103 cells/μL, consider implementing decision rules within your Laboratory Information System (LIS), Middleware, or on affected instruments. These rules can alert operators to initiate follow-up actions as described in this letter. •Consult your Medical Director to determine if a retrospective review of results is warranted. •Beckman Coulter recommends posting this letter on or near the affected analyzers | 2025-08-10 19:26
test | test | 1 | - | - | - | 2025-08-12 13:34
Fritz Stephan GmbH | SOPHIE | 4 | - | Unfortunately PC V has been withdrawn due to another software error that will lead to the effect that changes of the alarm limits for FiO will no longer be transmitted to the MC As a result the effective FiO alarm limits remain at the initially configured values even though the display shows the changed values Changes of the FiO or FiO alarm limits will therefore lead to false positive FiO alarms x D However the FiO mixture and measurement are not affected x D Deviations of more than of the set FiO between the set FiO and the measured FiO will still generate an alarm independent from the adjustable alarm limits | Fritz Stephan GmbH has already corrected that issue in Software PC V4.5.5. If you have already implemented FSCA QM285 in one or more of your devices, please update them to PC V4.5.5 immediately. | 2025-08-12 13:48
GS Elektromedizinische Geräte G. Stemple GmbH | Defib Corpuls 3 Touch SLIM | 276 | - | Failure to perform appropriate checks may result in undetected device readiness It has been observed during post market surveillance activities that the Daily Checks of the Device are not performed consistently or the results of the tests are not being reviewed by users as per the recommended procedure This deviation from the IFU can result in undetected device readiness x D issues and may compromise emergency response effectiveness | To address this, the following corrective measures will be implemented: • The IFU will be updated to clearly communicate that execution and verification of the Daily Checks of the Device is mandatory for ensuring the device remainsin a ready-for-use condition. • Users will be instructed that failure to perform and document the daily check constitutes non-compliance with the manufacturer’s operating instructions and may result in delayed or ineffective therapy delivery. In the event of a failed self-test, immediate action must be taken, as described in IFU chapter 11 Procedure in Case of Malfunctions including temporary removal from service, replacement with a backup device, and escalation to your responsible sales and service partner. | 2025-08-12 17:09
GS Elektromedizinische Geräte G. Stemple GmbH | Defib Unit Corpuls3 | 276 | - | Failure to perform appropriate checks may result in undetected device readiness It has been observed during post market surveillance activities that the Daily Checks of the Device are not performed consistently or the results of the tests are not being reviewed by users as per the recommended procedure This deviation from the IFU can result in undetected device readiness x D issues and may compromise emergency response effectiveness | To address this, the following corrective measures will be implemented: • The IFU will be updated to clearly communicate that execution and verification of the Daily Checks of the Device is mandatory for ensuring the device remainsin a ready-for-use condition. • Users will be instructed that failure to perform and document the daily check constitutes non-compliance with the manufacturer’s operating instructions and may result in delayed or ineffective therapy delivery. In the event of a failed self-test, immediate action must be taken, as described in IFU chapter 11 Procedure in Case of Malfunctions including temporary removal from service, replacement with a backup device, and escalation to your responsible sales and service partner. | 2025-08-12 17:09
Alcon Laboratories Inc. | Viscous Fluid Control Pak | 1332 | - | Alcon is conducting this Medical Device Field Safety Corrective Action of Alcon Standalone Vitrectomy Consumables as there is potential for some pouches within impacted lots to have an incomplete seal | We are asking that you locate and dispose of any affected lots of Alcon Standalone Vitrectomy Consumables remaining in your inventory. To comply with this Medical Device Field Safety Corrective Action and request the replacement of any unused product, please take the following steps: 1. Review your inventory to determine if you have any unused affected product within your facility. See table on page 1 for affected Standalone Vitrectomy Consumable lots shipped to Saudi Arabia. 2. Segregate and dispose of any unused affected product from your inventory. 3. Respond to Alcon indicating your understanding of these instructions even if you have zero (0) units remaining in inventory by completing and returning the attached “Response Form” and returning to Alcon via email. 4. Please forward this notification to all departments within your organization who may be in possession of this affected product; and any other organization to which this product may have been transferred. | 2025-08-12 23:22
Alcon Laboratories Inc. | Illuminator | 1332 | - | Alcon is conducting this Medical Device Field Safety Corrective Action of Alcon Standalone Vitrectomy Consumables as there is potential for some pouches within impacted lots to have an incomplete seal | We are asking that you locate and dispose of any affected lots of Alcon Standalone Vitrectomy Consumables remaining in your inventory. To comply with this Medical Device Field Safety Corrective Action and request the replacement of any unused product, please take the following steps: 1. Review your inventory to determine if you have any unused affected product within your facility. See table on page 1 for affected Standalone Vitrectomy Consumable lots shipped to Saudi Arabia. 2. Segregate and dispose of any unused affected product from your inventory. 3. Respond to Alcon indicating your understanding of these instructions even if you have zero (0) units remaining in inventory by completing and returning the attached “Response Form” and returning to Alcon via email. 4. Please forward this notification to all departments within your organization who may be in possession of this affected product; and any other organization to which this product may have been transferred. | 2025-08-12 23:22
CareDx Pty Ltd | AlloSeq® cfDNA Software | 1 | - | Two sample results were received one contaminated and one with suboptimal amplification where the QC mark was green and the results should not have been accepted due to the aforementioned reasons The software IFU states in Chapter Troubleshooting section If all the QC metrics are green passed there should be no need for any troubleshooting of results | CareDx will update the IFU on the CareDx website. | 2025-08-13 10:58
Dexcom, Inc | Dexcom G6 iOS CGM App | 61 | - | Issue Problem Dexcom has identified an issue with the G iOS OUS CGM App versions and where when a user re foreground their app they may experience some delay before the latest incoming EGV reading is displayed on the User interface UI display and trend graph of the screen In this scenario an older estimated glucose value EGV reading is displayed before the UI screen is updated While alerts alarms and notifications are not affected by this issue the older EGV reading displayed to the user before the UI is updated may or may not be perceived by the user If the user does not perceive the delay there is a potential to make treatment decision based on the older EGV reading displayed to make treatment decision ignoring symptoms and or expectations As a result Dexcom considers this risk unacceptable and is initiating a mandatory app update x D x D Root Cause Investigation revealed an error was made with the implementation of observers within the viewWillAppear function rather than the viewDidLoad function This causes an increase in latency due to a build up of new observers and an increase of function calls that are responsible for checking the volume mute status of the phone and preparing to play an audio alert if required causes the runtime of the application to slow over time | Mandatory Customer Action: Upgrading from app version 1.11.2 or 1.12.0 is mandatory and you will not be able to use those app versions after 15 September, 2025. Update your app now to continue using the app: 1. Tap Update App to go to the app store 2. Install the latest G6 app version 3. Open the G6 app If you need assistance, contact Technical Support: https://www.dexcom.com/global Actions to be taken: Dexcom has access to contact all remaining users of the affected App versions through deployment of in-App notifications. Dexcom intends to utilize this in-App messaging and is targeting to begin notifying all remaining users of the affected G6 OUS iOS CGM App versions of this issue at a minimum weekly between 20 August and 29 August 2025. This message will also notify users that the affected app versions will be blocked on 15 September 2025. Users must either update upon receiving this message or may choose to defer the update until closer to the blocking date. Dexcom has the ability to track user app versions to confirm that the affected users have upgraded their App or discontinued use of the affected App version. The blocking completed via the in-app obsolescence message will ensure that all users are no longer able to use the affected version of the App. | 2025-08-14 07:45
medin Medical Innovations GmbH | medin-NC3® | 7 | - | We ve identified that the blower in the medin NC may fail due to overheating which in some cases has led to device failure | replacement of the blower after 4500 operating hours | 2025-08-14 12:47
SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD. | BeneVision N1 | 8 | - | Mindray has identified that when the BeneVision N patient monitor is used as a module with x D other Mindray monitors there is low probability that upon setting an alarm pause the device may x D remain in the alarm pause state without displaying the remaining alarm pause time in this x D situation a prompt indicating the alarm pause state will remain visible on the interface and all x D other monitoring functions and technical alarm functions will continue to operate normally Users x D can cancel the alarm pause at any time This issue only affects BeneVision N software versions x D V and earlier x D Therefore BeneVision N devices with software versions later than V are not affected by x D this issue | Advise on action to be taken by the CUSTOMER / USER: 1. Please ensure that this Field Safety Notice (FSN) is communicated to all relevant personnel within your organization. 2. After receiving this Field Safety Notice, you may continue to use the BeneVision n N1 as a module with other Mindray monitors. If the "Alarm Pause" quick key is pressed on the BeneVision N1 or the associated monitor, please follow the instructions in the manual to confirm that the remaining alarm pause time is displayed in the physiological alarm information area on the main interface of the monitor where the "Alarm Pause" was activated. If only "Alarm Pause" is displayed, press the "Alarm Pause" quick key again to exit the alarm pause state. 3. The Mindray service team or authorized service personnel will contact you as soon as possible to arrange a software upgrade. 4. Please return the Field Safety Notice Acknowledgement Form to Mindray to confirm that this notice has been read and understood. Advise on action to be taken by the distributor: 1. Please ensure that this Field Safety Notice (FSN) is communicated to all relevant personnel within your organization, as well as to any organization that has received the affected BeneVision N1 devices. 2. If your organization has any BeneVision N1 devices on the affected list, please contact the Mindray service team or authorized service personnel as soon as possible to arrange a software upgrade. 3. Kindly assist the Mindray service team or authorized service personnel in completing this FSCA action. 4. Please return the Field Safety Notice Acknowledgement Form to Mindray to confirm that this notice has been read and understood. | 2025-08-14 16:02
Derma Sciences Canada Inc. | Antibacterial Medical Honey | 22500 | - | Inadequate sealing of sterile barrier packaging x D Shipping box does not adequately protect device during transportation x D Tube twist off cap failure x D | - | 2025-08-15 13:35
Derma Sciences, Inc | Antibacterial Wound Gel | 22500 | - | Inadequate sealing of sterile barrier packaging x D Shipping box does not adequately protect device during transportation x D Tube twist off cap failure x D | - | 2025-08-15 13:35
Siemens Healthcare GmbH | MAGNETOM Skyra | 11 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-17 09:55
Siemens Healthcare GmbH | MAGNETOM Vida | 11 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-17 09:55
Siemens Healthcare GmbH | MAGNETOM Lumina | 11 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-17 09:55
Siemens Healthcare GmbH | MAGNETOM Skyra | 11 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-17 09:55
Siemens Healthcare GmbH | MAGNETOM Vida | 11 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-17 09:55
Siemens Healthcare GmbH | MAGNETOM Lumina | 11 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-17 09:55
Siemens Healthcare GmbH | MAGNETOM Skyra | 11 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-17 09:55
Siemens Healthcare GmbH | MAGNETOM Vida | 11 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-17 09:55
Siemens Healthcare GmbH | MAGNETOM Lumina | 11 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-17 09:55
Siemens Healthcare GmbH | MAGNETOM Skyra | 11 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-17 09:56
Siemens Healthcare GmbH | MAGNETOM Vida | 11 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-17 09:56
Siemens Healthcare GmbH | MAGNETOM Lumina | 11 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-17 09:56
Siemens Healthcare GmbH | MAGNETOM Skyra | 11 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-17 09:57
Siemens Healthcare GmbH | MAGNETOM Vida | 11 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-17 09:57
Siemens Healthcare GmbH | MAGNETOM Lumina | 11 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-17 09:57
Tem Innovations GmbH (Werfen) | ROTEM sigma complete - ROTEM sigma complete + hep | 81 | - | Werfen has identified that the above mentioned lots of ROTEM sigma cartridges have a higher than expected likelihood of generating pressure errors during the ROTEM sigma functional checks These functional checks take place before and after the screen notification to add a sample If this pressure error is identified after the sample has already been inserted it will require sample recollection To minimize the waste of patient samples we ask that you follow the below listed procedure to allow the full functional checks to complete before sample insertion | Based on the above, please take the following immediate actions: - If you are using the affected lot numbers identified above in your facility, please review the provided guidance and adapt your workflow accordingly. 1. Press the start button: 2. Follow the instructions on the screen and remember the time on the clock when cartridge was inserted: 3. Enter or scan patient data and do not confirm: (Confirmation will hide the countdown) 4. 90 seconds after the cartridge was inserted confirm patient data: 5. Insert Sample: - Document the acknowledgement on the Customer Reply Form and return the completed and signed form to the fax number or e-mail address listed. - Share this information with your clinical users and laboratory staff and follow your internal procedures. - Forward this notification to all affected locations within your facility. - Report all device-related suspected serious incidents to the manufacturer, distributor, local contact point, and, if appropriate, the National Competent Authority. - Retain a copy of this notification for your records. | 2025-08-17 10:58
LeMaitre Vascular, Inc. | Tuftex 3F Over-The-Wire Embolectomy Catheter | 499 | - | During product testing packages were found to have incomplete seals sterile barrier x D x D Some parts approx could exhibit voids on the seal area where the tubing connects to the boat tray The voids were observed visually against backlight and confirmed upon opening of Tyvek lid Seal integrity testing of other samples with visually detected voids confirmed failures in some of the samples indicating compromised packaging | Customers are to: - Identify and quarantine any impacted product from your inventory. - Contact the sponsor indicating number of affected devices and request an RMA number to return the products. - Return all impacted products to the sponsor. | 2025-08-17 11:09
Merit Medical Systems ,Inc. | Merit Angioplasty Pack™PhD™Hemostasis Valve | 9650 | - | Potential leakage x D | Merit has chosen to remove the units from the market and requests that you immediately stop using or distributing the affected lots and return the units to Merit | 2025-08-17 13:14
Aesculap AG | ELAN 4 FIXED DURAGUARD | - | - | In the FSN | In the FSN | 2025-08-19 10:04
Medtronic, Inc. | DLP® Left Heart Vent Catheter | 680 | - | Since March there has been an increase in complaints reported for the DLP Malleable Left Heart Vent Catheter for the catheters resisting shape retention when being bent The catheters are intended to be malleable and retain a bend in the shaft x D An investigation into the complaints identified a potential root cause traced back to a change at the sub tier supplier to remove the annealing step in the stainless steel wire This wire is within the catheter CFNs and and is intended to be malleable to permit precise shaping of the catheter Annealing is a critical process for shape retention properties of the wire which enables the catheter to retain its shape when bent x D | Medtronic requests that you take the following actions: • Review your inventory for listed product using attachment A. • Immediately identify and quarantine all unused, listed product in your inventory. • Return unused, listed product in your inventory to Medtronic. Your Medtronic representative can assist you in the return of affected product as necessary. • Complete the enclosed Customer Acknowledgement Form and email to Medtronic. This form must be returned even if you do not have any affected product in your possession. • Please share this notification with others in your organization as appropriate. If product listed above has been forwarded to another facility, please notify the facility of this Medtronic Urgent Field Safety Notice. • Please maintain a copy of this communication in your records. | 2025-08-19 10:37
MAQUET Cardiovascular, LLC | Heartstring III Proximal Seal | 109 | - | Failure of the Heartstring Seal to Load This failure applies to any malfunction of the device that occurs x D during the four step loading process using the Heartstring Loading Device the visual inspection of a x D successfully loaded HS III Seal and unlocking the Delivery Device upon successful completion of the HS x D III Seal loading steps x D Failure of the Heartstring Seal to Deploy into the aortotomy This failure applies to any malfunction of x D the HS III Seal and or Delivery Device from the time the Delivery Device is successfully unlocked to once x D complete delivery of the HS III Seal into the aortotomy and all contents of the Proximal Seal the Tether x D Tension Spring Seal Stem Anchor Tab have exited the Delivery Device x D Failure of the deployed Heartstring Seal to provide adequate hemostasis This failure applies to any x D malfunction of a successfully deployed HS III Seal that does not adequately control bleeding at the x D anastomotic to allow the surgeon to perform suturing of the anastomosis despite gentle adjustment of x D the Tension Spring after Seal deployment and or the use of a blower mister suction or saline syringe to x D clear the anastomotic site This failure includes customer identification of a defective or damaged HS III x D Seal following successful deployment of the HS III Seal into the aorta that results or may result in x D inadequate hemostasis that prevents the surgeon to complete the anastomosis without removing the x D defective HS III Seal x D Root cause investigation is ongoing at this time x D | Failure of the Heartstring Seal to Load a. Important reminders: i. Per the device IFU, the HS III Seal is intended to be loaded and removed from the Loading Device, then inspected and adjusted if necessary to confirm successful loading and finally unlocked prior to clearing any adventitia from the planned anastomotic area and creating the aortotomy. Loading the HS III Seal and confirming successful loading prior to creating the aortotomy minimizes risk to patients because the aorta remains untouched until confirming successful loading of the HS III Seal. ii. If the HS III Seal fails to load as intended, the device should be replaced. b. How to avoid/minimize occurrence of this failure i. Review the Heartstring device IFU prior to use, especially if Heartstring is not used on a routine basis. ii. Load the HS III Seal by carefully following the device step-by-step Instructions for Use (IFU). iii. Per Section 4.3-4.5 of the IFU, confirm the HS III Seal has been properly loaded into the Delivery Device by inspecting the loaded HS III Seal once the HS III Seal has been loaded into the Delivery Device and has been removed from the Loading Device. Page: 4 of 5 OT# 1336551 iv. A HS III Seal that is not properly inspected after removal from the Loading Device may fail during HS III Seal Deployment and cause patient harm. (see information relating to Failure to Deploy). v. Complete successful loading of the HS III Seal prior to cleaning the aortic surface adventitia and creating the aortotomy. c. Action to be taken if the HS III Seal fails to load. i. Discard the failed HS III Seal. ii. Always have a second HS III Seal available to replace a failed device. Refer to the device IFU to ensure successful loading. 2. Failure of the Heartstring III Seal to Deploy into the aortotomy a. Important reminders: i. Ensure the HS III Seal has been properly loaded into the Delivery Device as described in the section above. ii. After removing the loaded HS III Seal from the Loading Device, inspect the loaded HS III Seal, according to IFU Sections 4.3-4.5, to confirm the following: 1. The loaded HS III Seal is not cracked. 2. The portion of the HS III Seal that is not loaded into the Delivery Device Tube does not flare wider than the Delivery Device Tube diameter. 3. The Tension Spring ends are aligned such that they will contact both proximal edges of the loaded HS III Seal during Seal delivery into the aortotomy. b. How to avoid/minimize the occurrence of this failure: i. After confirmation of successful HS III Seal loading and creation of the aortotomy, deploy the loaded HS III Seal into the aortotomy by carefully following the steps for HS III Seal deployment documented in the device IFU. ii. If the HS III Seal fails to deploy as intended, the device should be replaced. c. Action to be taken if the HS III Seal fails to successfully deploy into the aorta. i. Occlude the aortotomy to control bleeding using an index finger. ii. Open a second HS III Seal. Load the HS III Seal per the device IFU. Deliver the loaded HS III Seal into the aortotomy and follow the IFU to complete the anastomosis. iii. If delivery of the second HS III Seal into the aortotomy is not successful, or surgeon preference determines the situation does not allow for use of a replacement HS III Seal, place a partial occlusion clamp on the aorta, isolating the aortotomy, and perform the anastomosis per traditional surgical technique. 3. Failure of the deployed Heartstring III Seal to provide adequate hemostasis: a. How to avoid/minimize occurrence of this failure mode: i. Always load and carefully inspect the loaded HS III Seal to confirm successful loading of the HS III Seal into the Delivery Device prior to clearing any aorta adventitia and creating the aortotomy. (See important reminders directly above.) ii. Do not use the Heartstring Proximal Seal System in patients with a thin-walled aorta. iii. When performing multiple anastomosis, ensure all anastomotic sites are at least 1.5 cm apart to ensure hemostasis. iv. Do not soak the HS III Seal in solution prior to deployment. b. Action to be taken if the HS III Seal does not open/provide adequate hemostasis post Seal deployment. Page: 5 of 5 OT# 1336551 i. Cover the aortotomy with an index finger and gently adjust the Tension Spring to ensure the HS III Seal has been fully deployed. ii. If lack of adequate hemostasis persists: 1. Remove the deployed HS III Seal from the aortotomy per device IFU. 2. Load a second HS III Seal and deploy the loaded Seal into the aortotomy per device IFU and complete the anastomosis. 3. If the second HS III Seal is not successful or surgeon preference determines the situation does not allow for use of a replacement HS III Seal, place a partial occlusion clamp on the aorta, isolating the aortotomy, and perform the anastomosis per traditional surgical technique | 2025-08-19 11:32
Biomerieux | VIDAS | 11 | - | An overestimation of FPSA results when using VIDAS FPSA Ref lots listed above was internally detected now identified as resulting from a slow drift of the adjustment panel used in the Quality Control calibration process x D x D The IFU currently states that an FPSA TPSA ratio is more indicative of prostate cancer and a ratio of or greater is more indicative of BPH Taking the overestimated FPSA into account a corrected FPSA TPSA ratio threshold was identified for patient results obtained using the impacted lots of FPSA Ref listed above This will allow you to continue the use of involved products by applying the corrected threshold x D | corrective action at the manufacturing site was already implemented to correct the issue and manufacture new lots within the performance claims. All new manufactured lots to be released will be conform to performance claims. The new conforming lot (Lot 1010713540) will be manufactured in the upcoming weeks and will not require any correction to be applied | 2025-08-20 10:12
Getinge International AB | GETINGE 88-5 Washer disinfector | 35 | - | From our investigation which includes an evaluation of worst case scenarios it was concluded that there is a medium risk that this issue may cause injury such as skin burns or smoke inhalation While the likelihood is low the potential consequences justify immediate corrective action | In the FSN | 2025-08-20 10:38
optimed Medizinische Instrumente GmbH | Tentos4F | 298 | - | Risk of damage due to unauthorized folding or bending during unpacking or preparation | Consider changes and modifications to the instructions for use (IFU) as the manufacturer has changed the IFU and Following the patient's treatment recommendations | 2025-08-20 12:32
Instrumentation Laboratory | HemosIL LMW Heparin Controls; Part Number: 0020300200; | 156 | - | Multiple complaints indicating lower than expected quality control QC results and in some instances QC recovery below the lower limit of the package insert acceptance range when testing is performed with HemosIL Liquid Anti Xa and HemosIL Heparin Calibrator In some complaints failure to pass QC resulted in patient testing delays x D | Requested Customer Actions Please take the following actions: " Follow the instructions in the original Urgent Medical Device Correction letter to update the ACL TOP test parameters with the corrected value assignments for HemosIL Heparin Calibrators (Part No. 0020300600). " For HemosIL UF Heparin Controls (Part No. 0020300300), continue to use the current target values and acceptance ranges as labeled. " For HemosIL LMW Heparin Controls (Part No. 0020300200), use the reassigned target values and acceptance ranges for the Low and High Controls as provided in the table above. " Note: if the 2D Barcode is scanned, customers need to manually override the uploaded values with those provided in this letter. " Follow the instructions below to update the Quality Control Material Definitions with the updated values (same instructions for all Instrument models): " Select Setup from the menu bar and then select Material List from the dropdown menu. " From Material List , double click on the Material Name LMW Hep Low Control and LMW Hep High Control (performed separately) to open Material Definition . " From Material Definition LMW Hep Low Control or LMW Hep High Control , click on Assigned Values icon (picture of bullseye) under the header. An editable table titled Assigned Values for IL Material Definition Quality Control LMW Hep Low Control or LMW Hep High Control will appear with the Anti-Xa Test Code, Active Lot, and Alternative Lot (if applicable). " In the Assigned Values table, select the Active and/or Alternative lot, and change the target values for LMW Hep Low Control and LMW Hep High Control to the reassigned values provided in the table. " After entering the reassigned values for LMW Hep Low Control and LMW Hep High Control , click on OK to save the changes. " Run quality controls per labeled instructions for use. New acceptance ranges for HemosIL L | 2025-08-21 10:18
Siemens Shenzhen Magnetic Resonance Ltd. | MAGNETOM Spectra | 1 | - | Potential error an ice blockage may exist within the magnet venting system x D In the event of a quench helium gas may be unable to escape through the designed vent paths leading to a pressure build up within the helium containment system This pressure build up could ultimately rupture the helium containment system potentially resulting in a helium leak into the scanning room | Siemens Healthineers is preparing an inspection and, if necessary, a correction that will resolve this potential issue. | 2025-08-21 11:42
Jiangsu Qianyuan New Material Technogy Co.,Ltd. | Optical Lens | - | - | Incorrect label on the devices | • Withdraw and correct the label for all affected products in the Saudi market. • Inform all affected customers by the safety alert. | 2025-08-24 08:28
Olympus | ESG-410 Electrosurgical Generator | - | - | Olympus received forty three complaints between July to March from users where the ESG Electrosurgical Generator displayed error message E System Error followed by either automatic single reboots of the device or continuous reboot loops Users reported that these reboots occurred both during and prior to clinical use An investigation by Olympus identified that this issue is caused by a previous change to an electronic component This component change impacted an internal performance check parameter which could then result in the error message being displayed and an automatic system reboot In some cases the system will be in a continuous reboot None of the forty three complaints resulted in an adverse event x D An Olympus representative will reach out to you to arrange for a device update which resolves the cause of this issue | Our records indicate that your facility has received one or more of the affected units. Olympus requests you to take the following actions: 1. Examine your inventory for affected serial numbers (refer to Table 1 for affected serial numberinformation). 2. Customers may continue to use the device whilst they are waiting for the update, ensuring all users followthe product Instructions for Use (IFU). This includes the instruction to prepare a spare electrosurgicalgenerator, or an alternative procedure, to avoid interruption due to an unexpected electrosurgicalgenerator failure during the procedure. If any problems are encountered with the device, including errormessages and automatic reboots, please contact Olympus. 3. Olympus representative will reach out to you to arrange a mutually convenient time to update theaffected device at no charge. 4. Olympus requests that you acknowledge receipt of this letter. Indicate on the Reply that you havereceived and understood this notification by filling out and returning the completed enclosed Reply Formback to your local Olympus representative ra@olympus-mea.com latest by 22.08.2025. 5. Please forward this notice to other users who may have the affected products if you have furtherdistributed it. | 2025-08-24 08:46
Aesculap AG | ELAN 4 FIXED DURAGUARD STANDARD | 6 | - | During the product realization process there was a mix up between the article numbers GB R and GB R GB R was incorrectly labeled as GB R and vice versa Therefore the ELAN FIXED DURAGUARD STANDARD was labeled as Long and the ELAN FIXED DURAGUARD LONG was labeled as Standard | Identify Device and Return Device | 2025-08-26 10:44
GE Medical Systems Information Technologies, Inc | MAC VU360 | 474 | - | GE HealthCare has become aware that if a user incorrectly places the Acquisition Module into the Acquisition Module Holder it can result in excessive bending and wear of the Acquisition Trunk Cable This is visible as a damaged sheath or exposed wires If a damaged Acquisition Trunk Cable is not recognized and removed from service the cable can overheat with continued use In the event of prolonged contact with the heated cable this could cause a partial thickness thermal injury x D x D There have been no injuries reported as a result of this issue x D | GEHC will send FMI kits to all affected systems for customers to install/replace old parts | 2025-08-26 11:15
Insightra | ULTRA IABP Catheter Kit | 11 | - | Voluntarily updating the instructions for use IFU | These IFUs are not provided in printed form but are instead accessible via QR code included in the kit box. All updates have already been implemented and do not impact the safety, market availability, or clarity of use in KSA.The customers have been informed to access updated IFU by scanning QR code on product and notification from manufacturer is shared with them No further corrective action required | 2025-08-27 18:23
Fisher & Paykel Healthcare Limited | Airvo 2 | 2249 | - | This notice is to inform you of an update to the Disinfection Kit User Manual for maintaining your Airvo and myAirvo device throughout its expected service life The Airvo and myAirvo have an audible alert that sounds for at least seconds when the unit is unintentionally disconnected from a power source The Disinfection Kit User Manual has been updated to include a regular test of this alert to be carried out in between patient uses | F&P has updated the Disinfection Kit User Manual with a step to check the audible alert. As below: 7a: Perform 'Power-Out Alarm' test (Airvo 2 and myAirvo 2 only) Ensure steps 1-6 have been carried out, so that the alarm circuit is charged. (Please refer to warning section in IFU page 4 in English ). Unplug the unit from the mains/utility power socket. Check that the audible alarm sounds for at least 120 seconds. The unit is now ready for storage. 7b: Switch off the Airvo Switch off the unit and unplug from the mains/utility power socket. The unit is now ready for storage Prior to use on a new patient, follow the instructions of the updated Disinfection Kit User Manual. Please acknowledge receipt of this notice by following these steps: 1. Enter your organisation and contact information on the response form. 2. Acknowledge and sign that you have read and understood this notice. 3. If you have distributed Airvo 2 or myAirvo 2 devices, notify your customers and include a copy of this notice. | 2025-08-28 14:06
Fisher & Paykel Healthcare Limited | myAirvo 2 | 2249 | - | This notice is to inform you of an update to the Disinfection Kit User Manual for maintaining your Airvo and myAirvo device throughout its expected service life The Airvo and myAirvo have an audible alert that sounds for at least seconds when the unit is unintentionally disconnected from a power source The Disinfection Kit User Manual has been updated to include a regular test of this alert to be carried out in between patient uses | F&P has updated the Disinfection Kit User Manual with a step to check the audible alert. As below: 7a: Perform 'Power-Out Alarm' test (Airvo 2 and myAirvo 2 only) Ensure steps 1-6 have been carried out, so that the alarm circuit is charged. (Please refer to warning section in IFU page 4 in English ). Unplug the unit from the mains/utility power socket. Check that the audible alarm sounds for at least 120 seconds. The unit is now ready for storage. 7b: Switch off the Airvo Switch off the unit and unplug from the mains/utility power socket. The unit is now ready for storage Prior to use on a new patient, follow the instructions of the updated Disinfection Kit User Manual. Please acknowledge receipt of this notice by following these steps: 1. Enter your organisation and contact information on the response form. 2. Acknowledge and sign that you have read and understood this notice. 3. If you have distributed Airvo 2 or myAirvo 2 devices, notify your customers and include a copy of this notice. | 2025-08-28 14:06
FIAB SpA | ZOLL Medical corp | - | - | Delay in delivering therapy due to non immediate recognition of the electrodes or confusion in the user caused by initially incorrect recognition x D Potentially ineffective therapy reduced energy delivered by the defibrillator to the adult patient x D if the adult electrodes are recognized by the AED as pediatric electrodes | Actions by the User ☑ Identify the devices from the affected batches, as listed in section 1.5 of the FSN (complete list) and in particular in the attached Customer Response Form (batches distributed to the individual Customer/User) ☑ Segregate devices and do not use them ☑ Contact your Local Representative, as indicated on the FSN cover page, to arrange the return of the affected devices. | 2025-08-28 15:03
Cartiva Inc. | CARTIVA® Synthetic Cartilage Implant | 0 | - | Stryker has become aware of recently published data and post market reports indicating that patients implanted with Cartiva SCI may experience a higher than expected occurrence rate when compared to data submitted in the FDA PMA of the following documented hazards revision removal implant subsidence displacement persistent or worsening pain nerve damage or fragmentation x D x D Cartiva SCI devices have been observed in some cases to be revised removed at higher rates than previously observed in the initial Cartiva SCI premarket and post approval studies x D x D The labelling IFU and surgical technique includes warnings precautions and potential complications and adverse reactions related to the Cartiva product The hazards listed are known potential hazards associated with these devices and related procedures | Recall letter to customers that may have any affected product in their possession. Customers are to: - Review the Customer Letter and quarantine any affected product - Destroy affected items within their possession - Inform Stryker if affected items has been transferred | 2025-08-28 17:27
Cardiac Pacemakers Inc. a subsidiary of BSC Corporation | ACCOLADE Family of Pacemakers and CRT-Ps | 67 | - | This communication provides an update to the December customer communications SFDA ref no SA regarding the potential for the ACCOLADE family of pacemakers to initiate Safety Mode due to a high battery impedance state Safety Mode provides back up pacing under critical circumstances it is not intended to be a substitute for chronic pacing therapy The non programmable Safety Mode pacing parameters may not provide optimal support of a patient s cardiac condition e g adequacy of underlying escape rhythm the need for AV VV pacing for cardiac synchrony and or the potential for pacing inhibition due to myopotential oversensing x D Boston Scientific is developing a software upgrade designed to prevent pacemakers in the ACCOLADE family from initiating Safety Mode in an ambulatory setting due to a high battery impedance state Boston Scientific anticipates releasing this software in the region in the coming months x D | Customers notification Software upgrade when the software is released. | 2025-08-29 10:32
Drägerwerk AG & Co. KGaA | Atlan A350 | 20 | - | Failure of the ventilator motor preventive replacement of the ventilator x D motor | Phase 1 1. A Safety Notice is sent to the recipients named in the location list. If recipients do not respond in the first round of communication, remind them two more times. Then send a closing letter and archive it in TrackWise. 2. Add the response status to the updated location list (in a separate column). Enter the final numbers in a comment field. As part of Phase 1, a closing letter of the security notice does not preclude the customer from being contacted in Phase 2 of the field action. Note: Phase 1 and phase 2 can be combined in one action within the target date of delivery of the safety notice. Phase 2 Step 1: Dismantling the ventilator Step 2: Checking the encoder date code on the dismantled ventilator, see additional information Step 3: Replace the ventilator if necessary | 2025-09-02 12:10
ZHENJIANG CV OPTICAL CO., LTD | BLUE SPEKTRO | 1 | خطأ بالملصق التعريفي بالمنتج لايؤثر عل سلامة المنتج والمستخدم | - | The product bearing an incorrect identification label has been recalled in order to replace it with the correct label matching the approved product description. | 2025-09-03 00:06
DiaMed GmbH | ID-Anti-S | 1 | - | Dear Valued Customer Channel Partner x D x D This letter is to inform you about a quality issue related to false positive reactions obtained with Bio Rad anti S test serum A safety issue has been identified impacting the below listed products that could pose a potential risk for patients x D x D Reason for the Field Safety Notice x D x D Bio Rad has received complaints of false positive reactions associated with ID Anti S ref An internal investigation confirmed that the test serum ID Anti S was contaminated with an anti Dia antibody The contaminant was identified in a raw material used during the manufacture of the affected lots listed below The presence of the contaminant may lead to false positive results x D x D Kindly review the attached URGENT FIELD SAFETY NOTICE x D | Action(s) to be taken by the Customer: Bio-Rad is requesting that customers affected by this notice take the following action: • Check your inventory for any remaining of the affected lots. • If you find any affected product, dispose following your internal protocol. • Contact Bio-Rad to request replacement for any of the affected products. • If any patients or donors were tested with the affected lots and received a positive result, consult with your medical director to decide if further testing or other steps are needed. Please ensure this notice is passed to all those who need to be aware within your organization or to any organization where the impacted devices have been transferred. Please complete and return the attached response form as soon as possible so that we are assured you have received this important communication. Resolution by Bio-Rad: Bio-Rad takes product quality and safety very seriously, and investigations are ongoing to determine how our quality control can be improved to avoid the recurrence of this issue. The National Competent (Regulatory) Authority has been informed of this field safety notice. Kindly review the attached URGENT FIELD SAFETY NOTICE | 2025-09-04 10:10
DiaMed GmbH | ID-Antigen Profile III | 1 | - | Dear Valued Customer Channel Partner x D x D This letter is to inform you about a quality issue related to false positive reactions obtained with Bio Rad anti S test serum A safety issue has been identified impacting the below listed products that could pose a potential risk for patients x D x D Reason for the Field Safety Notice x D x D Bio Rad has received complaints of false positive reactions associated with ID Anti S ref An internal investigation confirmed that the test serum ID Anti S was contaminated with an anti Dia antibody The contaminant was identified in a raw material used during the manufacture of the affected lots listed below The presence of the contaminant may lead to false positive results x D x D Kindly review the attached URGENT FIELD SAFETY NOTICE x D | Action(s) to be taken by the Customer: Bio-Rad is requesting that customers affected by this notice take the following action: • Check your inventory for any remaining of the affected lots. • If you find any affected product, dispose following your internal protocol. • Contact Bio-Rad to request replacement for any of the affected products. • If any patients or donors were tested with the affected lots and received a positive result, consult with your medical director to decide if further testing or other steps are needed. Please ensure this notice is passed to all those who need to be aware within your organization or to any organization where the impacted devices have been transferred. Please complete and return the attached response form as soon as possible so that we are assured you have received this important communication. Resolution by Bio-Rad: Bio-Rad takes product quality and safety very seriously, and investigations are ongoing to determine how our quality control can be improved to avoid the recurrence of this issue. The National Competent (Regulatory) Authority has been informed of this field safety notice. Kindly review the attached URGENT FIELD SAFETY NOTICE | 2025-09-04 10:10
DiaMed GmbH | ID-DiaPanel | 224 | - | Dear Valued Customer Channel Partner x D The purpose of this letter is to inform you about a quality issue related to Bio Rad ID DiaPanel and ID DiaPanel P x D Reason for the Action x D Following customer feedback and Bio Rad internal investigation we have identified that Cell in the lots listed below may show negative or weak positive reactions when tested with anti D antibodies e g IH QC Instead of the expected clear positive reaction Cell may exhibit reduced or absent reactivity x D Our testing confirms that this behavior is due to a lower density of D antigen on the donor red blood cells used for Cell This is likely associated with a weak or partial D phenotype x D x D Kindly review the attached | Action(s) to be taken by the Customer: Bio-Rad requests that customers affected by this notice take the following actions: • Use the updated Antigen Table already upload online and available for download: [F0617127X2_V.02 4.pdf] • Use the three additional D-positive cells (R1R1, R2R2, Ror) which remain suitable to confirm anti-D reactivity. The product remains safe for use without restrictions; other antigens expressed on Cell 1 are unaffected. Please ensure that this notice is shared with all relevant personnel within your organization and any third-party recipients of the affected products. Please complete and return the attached response form as soon as possible to confirm receipt of this important communication. Resolution by Bio-Rad: Bio-Rad takes product quality very seriously, and we have been diligently investigating all issues raised. Kindly review the attached | 2025-09-04 12:30
DiaMed GmbH | ID-DiaPanel-P papainized | 224 | - | Dear Valued Customer Channel Partner x D The purpose of this letter is to inform you about a quality issue related to Bio Rad ID DiaPanel and ID DiaPanel P x D Reason for the Action x D Following customer feedback and Bio Rad internal investigation we have identified that Cell in the lots listed below may show negative or weak positive reactions when tested with anti D antibodies e g IH QC Instead of the expected clear positive reaction Cell may exhibit reduced or absent reactivity x D Our testing confirms that this behavior is due to a lower density of D antigen on the donor red blood cells used for Cell This is likely associated with a weak or partial D phenotype x D x D Kindly review the attached | Action(s) to be taken by the Customer: Bio-Rad requests that customers affected by this notice take the following actions: • Use the updated Antigen Table already upload online and available for download: [F0617127X2_V.02 4.pdf] • Use the three additional D-positive cells (R1R1, R2R2, Ror) which remain suitable to confirm anti-D reactivity. The product remains safe for use without restrictions; other antigens expressed on Cell 1 are unaffected. Please ensure that this notice is shared with all relevant personnel within your organization and any third-party recipients of the affected products. Please complete and return the attached response form as soon as possible to confirm receipt of this important communication. Resolution by Bio-Rad: Bio-Rad takes product quality very seriously, and we have been diligently investigating all issues raised. Kindly review the attached | 2025-09-04 12:30
Philips Medical Systems Nederland B.V. | Allura Xper FD10 | 98 | - | The Nehalem Host PC s Nehalem IP image processing PC s Nehalem Flexvision PCs Q Geo PCs and Ivybridge Geo PCs that are part of the configuration of the Allura systems Allura Xper Allura CV and Allura Centron have a BIOS battery to preserve the configuration settings and time when the PC is powered off ensuring correct time and configuration at start up x D Per design when the Allura system is switched on PC s get powered and automatically initiate their start sequence which involves a full start of the PC and the Operating System without any user interaction e g no PC start press button keyboard input or PC login credentials x D Philips has identified that the BIOS battery may deplete faster than initially anticipated during the design phase When the battery is depleted the start up process of the Allura system will be halted and the Allura system will not start No user messages are shown before the battery is low on power or depleted | Philips will replace the BIOS battery in all affected systems. Your local Philips representative will contact you to schedule a visit of a Field Service Engineer visit to perform this action that is expected to start as of November 2025. | 2025-09-05 14:28
Philips Medical Systems Nederland B.V. | Allura Xper FD10 OR Table | 98 | - | The Nehalem Host PC s Nehalem IP image processing PC s Nehalem Flexvision PCs Q Geo PCs and Ivybridge Geo PCs that are part of the configuration of the Allura systems Allura Xper Allura CV and Allura Centron have a BIOS battery to preserve the configuration settings and time when the PC is powered off ensuring correct time and configuration at start up x D Per design when the Allura system is switched on PC s get powered and automatically initiate their start sequence which involves a full start of the PC and the Operating System without any user interaction e g no PC start press button keyboard input or PC login credentials x D Philips has identified that the BIOS battery may deplete faster than initially anticipated during the design phase When the battery is depleted the start up process of the Allura system will be halted and the Allura system will not start No user messages are shown before the battery is low on power or depleted | Philips will replace the BIOS battery in all affected systems. Your local Philips representative will contact you to schedule a visit of a Field Service Engineer visit to perform this action that is expected to start as of November 2025. | 2025-09-05 14:28
Philips Medical Systems Nederland B.V. | Allura Xper FD10/10 | 98 | - | The Nehalem Host PC s Nehalem IP image processing PC s Nehalem Flexvision PCs Q Geo PCs and Ivybridge Geo PCs that are part of the configuration of the Allura systems Allura Xper Allura CV and Allura Centron have a BIOS battery to preserve the configuration settings and time when the PC is powered off ensuring correct time and configuration at start up x D Per design when the Allura system is switched on PC s get powered and automatically initiate their start sequence which involves a full start of the PC and the Operating System without any user interaction e g no PC start press button keyboard input or PC login credentials x D Philips has identified that the BIOS battery may deplete faster than initially anticipated during the design phase When the battery is depleted the start up process of the Allura system will be halted and the Allura system will not start No user messages are shown before the battery is low on power or depleted | Philips will replace the BIOS battery in all affected systems. Your local Philips representative will contact you to schedule a visit of a Field Service Engineer visit to perform this action that is expected to start as of November 2025. | 2025-09-05 14:28
Philips Medical Systems Nederland B.V. | Allura Xper FD20 | 98 | - | The Nehalem Host PC s Nehalem IP image processing PC s Nehalem Flexvision PCs Q Geo PCs and Ivybridge Geo PCs that are part of the configuration of the Allura systems Allura Xper Allura CV and Allura Centron have a BIOS battery to preserve the configuration settings and time when the PC is powered off ensuring correct time and configuration at start up x D Per design when the Allura system is switched on PC s get powered and automatically initiate their start sequence which involves a full start of the PC and the Operating System without any user interaction e g no PC start press button keyboard input or PC login credentials x D Philips has identified that the BIOS battery may deplete faster than initially anticipated during the design phase When the battery is depleted the start up process of the Allura system will be halted and the Allura system will not start No user messages are shown before the battery is low on power or depleted | Philips will replace the BIOS battery in all affected systems. Your local Philips representative will contact you to schedule a visit of a Field Service Engineer visit to perform this action that is expected to start as of November 2025. | 2025-09-05 14:28
Philips Medical Systems Nederland B.V. | Allura Xper FD20 OR Table | 98 | - | The Nehalem Host PC s Nehalem IP image processing PC s Nehalem Flexvision PCs Q Geo PCs and Ivybridge Geo PCs that are part of the configuration of the Allura systems Allura Xper Allura CV and Allura Centron have a BIOS battery to preserve the configuration settings and time when the PC is powered off ensuring correct time and configuration at start up x D Per design when the Allura system is switched on PC s get powered and automatically initiate their start sequence which involves a full start of the PC and the Operating System without any user interaction e g no PC start press button keyboard input or PC login credentials x D Philips has identified that the BIOS battery may deplete faster than initially anticipated during the design phase When the battery is depleted the start up process of the Allura system will be halted and the Allura system will not start No user messages are shown before the battery is low on power or depleted | Philips will replace the BIOS battery in all affected systems. Your local Philips representative will contact you to schedule a visit of a Field Service Engineer visit to perform this action that is expected to start as of November 2025. | 2025-09-05 14:28
Philips Medical Systems Nederland B.V. | Allura Xper FD20/10 | 98 | - | The Nehalem Host PC s Nehalem IP image processing PC s Nehalem Flexvision PCs Q Geo PCs and Ivybridge Geo PCs that are part of the configuration of the Allura systems Allura Xper Allura CV and Allura Centron have a BIOS battery to preserve the configuration settings and time when the PC is powered off ensuring correct time and configuration at start up x D Per design when the Allura system is switched on PC s get powered and automatically initiate their start sequence which involves a full start of the PC and the Operating System without any user interaction e g no PC start press button keyboard input or PC login credentials x D Philips has identified that the BIOS battery may deplete faster than initially anticipated during the design phase When the battery is depleted the start up process of the Allura system will be halted and the Allura system will not start No user messages are shown before the battery is low on power or depleted | Philips will replace the BIOS battery in all affected systems. Your local Philips representative will contact you to schedule a visit of a Field Service Engineer visit to perform this action that is expected to start as of November 2025. | 2025-09-05 14:28
Philips Medical Systems Nederland B.V. | Allura Xper FD20/15 | 98 | - | The Nehalem Host PC s Nehalem IP image processing PC s Nehalem Flexvision PCs Q Geo PCs and Ivybridge Geo PCs that are part of the configuration of the Allura systems Allura Xper Allura CV and Allura Centron have a BIOS battery to preserve the configuration settings and time when the PC is powered off ensuring correct time and configuration at start up x D Per design when the Allura system is switched on PC s get powered and automatically initiate their start sequence which involves a full start of the PC and the Operating System without any user interaction e g no PC start press button keyboard input or PC login credentials x D Philips has identified that the BIOS battery may deplete faster than initially anticipated during the design phase When the battery is depleted the start up process of the Allura system will be halted and the Allura system will not start No user messages are shown before the battery is low on power or depleted | Philips will replace the BIOS battery in all affected systems. Your local Philips representative will contact you to schedule a visit of a Field Service Engineer visit to perform this action that is expected to start as of November 2025. | 2025-09-05 14:28
Philips Medical Systems Nederland B.V. | Allura Centron | 98 | - | The Nehalem Host PC s Nehalem IP image processing PC s Nehalem Flexvision PCs Q Geo PCs and Ivybridge Geo PCs that are part of the configuration of the Allura systems Allura Xper Allura CV and Allura Centron have a BIOS battery to preserve the configuration settings and time when the PC is powered off ensuring correct time and configuration at start up x D Per design when the Allura system is switched on PC s get powered and automatically initiate their start sequence which involves a full start of the PC and the Operating System without any user interaction e g no PC start press button keyboard input or PC login credentials x D Philips has identified that the BIOS battery may deplete faster than initially anticipated during the design phase When the battery is depleted the start up process of the Allura system will be halted and the Allura system will not start No user messages are shown before the battery is low on power or depleted | Philips will replace the BIOS battery in all affected systems. Your local Philips representative will contact you to schedule a visit of a Field Service Engineer visit to perform this action that is expected to start as of November 2025. | 2025-09-05 14:28
Philips Medical Systems Nederland B.V. | Allura Xper FD10 | 46 | - | The Geo PC of the model Wolfdale Q generation that controls the geometry movements of Philips Allura Systems may not perform as intended due to the deterioration of certain internal components the CMOS battery hard disk drive and or power supply unit If this deterioration occurs motorized movements of the system will become unavailable and the following message will be displayed to the user WARNING Motorized movement not available Manual movements of the stand longitudinal transversal and rotational for ceiling mounted systems and rotational for floor mounted systems and table longitudinal movements are available Imaging X ray functionality remains available The issue may be intermittent and may be resolved by performing a system cold restart However a system cold restart may not always resolve the issue A system warm restart will not resolve the issue x D x D | Philips will be replacing the Geo PC in all affected systems. Philips expects to start replacing the Geo PC units by September 2025. Your local Philips representative will contact you to schedule a visit to confirm If your system has a model Wolfdale (Q35 generation) Geo PC and to plan the replacement of the Geo PC once available. | 2025-09-05 14:50
Philips Medical Systems Nederland B.V. | Allura Xper FD10 OR Table | 46 | - | The Geo PC of the model Wolfdale Q generation that controls the geometry movements of Philips Allura Systems may not perform as intended due to the deterioration of certain internal components the CMOS battery hard disk drive and or power supply unit If this deterioration occurs motorized movements of the system will become unavailable and the following message will be displayed to the user WARNING Motorized movement not available Manual movements of the stand longitudinal transversal and rotational for ceiling mounted systems and rotational for floor mounted systems and table longitudinal movements are available Imaging X ray functionality remains available The issue may be intermittent and may be resolved by performing a system cold restart However a system cold restart may not always resolve the issue A system warm restart will not resolve the issue x D x D | Philips will be replacing the Geo PC in all affected systems. Philips expects to start replacing the Geo PC units by September 2025. Your local Philips representative will contact you to schedule a visit to confirm If your system has a model Wolfdale (Q35 generation) Geo PC and to plan the replacement of the Geo PC once available. | 2025-09-05 14:50
Philips Medical Systems Nederland B.V. | Allura Xper FD10/10 | 46 | - | The Geo PC of the model Wolfdale Q generation that controls the geometry movements of Philips Allura Systems may not perform as intended due to the deterioration of certain internal components the CMOS battery hard disk drive and or power supply unit If this deterioration occurs motorized movements of the system will become unavailable and the following message will be displayed to the user WARNING Motorized movement not available Manual movements of the stand longitudinal transversal and rotational for ceiling mounted systems and rotational for floor mounted systems and table longitudinal movements are available Imaging X ray functionality remains available The issue may be intermittent and may be resolved by performing a system cold restart However a system cold restart may not always resolve the issue A system warm restart will not resolve the issue x D x D | Philips will be replacing the Geo PC in all affected systems. Philips expects to start replacing the Geo PC units by September 2025. Your local Philips representative will contact you to schedule a visit to confirm If your system has a model Wolfdale (Q35 generation) Geo PC and to plan the replacement of the Geo PC once available. | 2025-09-05 14:50
Philips Medical Systems Nederland B.V. | Allura Xper FD20 | 46 | - | The Geo PC of the model Wolfdale Q generation that controls the geometry movements of Philips Allura Systems may not perform as intended due to the deterioration of certain internal components the CMOS battery hard disk drive and or power supply unit If this deterioration occurs motorized movements of the system will become unavailable and the following message will be displayed to the user WARNING Motorized movement not available Manual movements of the stand longitudinal transversal and rotational for ceiling mounted systems and rotational for floor mounted systems and table longitudinal movements are available Imaging X ray functionality remains available The issue may be intermittent and may be resolved by performing a system cold restart However a system cold restart may not always resolve the issue A system warm restart will not resolve the issue x D x D | Philips will be replacing the Geo PC in all affected systems. Philips expects to start replacing the Geo PC units by September 2025. Your local Philips representative will contact you to schedule a visit to confirm If your system has a model Wolfdale (Q35 generation) Geo PC and to plan the replacement of the Geo PC once available. | 2025-09-05 14:50
Philips Medical Systems Nederland B.V. | Allura Xper FD20/15 | 46 | - | The Geo PC of the model Wolfdale Q generation that controls the geometry movements of Philips Allura Systems may not perform as intended due to the deterioration of certain internal components the CMOS battery hard disk drive and or power supply unit If this deterioration occurs motorized movements of the system will become unavailable and the following message will be displayed to the user WARNING Motorized movement not available Manual movements of the stand longitudinal transversal and rotational for ceiling mounted systems and rotational for floor mounted systems and table longitudinal movements are available Imaging X ray functionality remains available The issue may be intermittent and may be resolved by performing a system cold restart However a system cold restart may not always resolve the issue A system warm restart will not resolve the issue x D x D | Philips will be replacing the Geo PC in all affected systems. Philips expects to start replacing the Geo PC units by September 2025. Your local Philips representative will contact you to schedule a visit to confirm If your system has a model Wolfdale (Q35 generation) Geo PC and to plan the replacement of the Geo PC once available. | 2025-09-05 14:50
Randox laboratories | Liquid Assayed Specific Protein Control Level 1 | - | - | Randox Laboratories can confirm that the Rheumatoid Factor RF concentrations in Liquid x D Assayed Specific Protein Controls Level Lots LPC and LPC have decreased Randox x D have since updated the IFU s and removed the associated targets and ranges as the controls x D are at the sensitivity limit of some methods x D Randox Laboratories can also confirm that Rheumatoid Factor RF target and range values x D have been reassigned in level lots LPC LPC and level Lot LPC Please see table x D below for updated target values | Action to be taken: Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Download updated IFU from www.randox.com and discard previous copies. Please accept our apologies for any inconvenience caused. Thank you for your patience and understanding. If you have any questions or concerns, please contact Randox Technical Services. | 2025-09-07 10:45
Randox laboratories | Liquid Assayed Specific Protein Control Level 2 | - | - | Randox Laboratories can confirm that the Rheumatoid Factor RF concentrations in Liquid x D Assayed Specific Protein Controls Level Lots LPC and LPC have decreased Randox x D have since updated the IFU s and removed the associated targets and ranges as the controls x D are at the sensitivity limit of some methods x D Randox Laboratories can also confirm that Rheumatoid Factor RF target and range values x D have been reassigned in level lots LPC LPC and level Lot LPC Please see table x D below for updated target values | Action to be taken: Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Download updated IFU from www.randox.com and discard previous copies. Please accept our apologies for any inconvenience caused. Thank you for your patience and understanding. If you have any questions or concerns, please contact Randox Technical Services. | 2025-09-07 10:45
Stago | Reagent | 52 | - | Stago conducted investigations and confirmed a potential risk of underestimation of D Dimer D Di levels with the two listed lots of STA Liatest D Di Plus | Discontinue their use and dispose of them in accordance with your local regulations. | 2025-09-07 14:03
Jaeger Medical | Vyntus products with SNIP/RHINO Option | 5 | - | The purpose of this communication is to inform you that Jaeger Medical formerly known as Vyaire x D Medical GmbH is voluntarily notifying customers regarding the Vyntus devices with the optional x D Rhinomanometry and Sniff Nasal Inspiratory Pressure SNIP features x D Rhinomanometry allows for the evaluation of ventilatory function of the nose x D SNIP allows for the non invasive measurement of inspiratory muscle strength x D It has been identified that the Reprocessing Instructions do not contain proper reprocessing x D instructions for the cleaning disinfection and disposal of some components V x D Connection set Elbow with tubing associated with these testing features x D To ensure patient safety and compliance all potentially affected devices should be disposed of x D after single use | Actions to be taken by Jaeger Medical (formerly Vyaire Medical GmbH): • Coordinate with Distributors/End-Users to ensure acknowledgment of the single use SNIP/RHINO components. As part of this process, impacted customers will be required to complete a response form confirming: o All impacted customers received the notification regarding the affected product and need to confirm use of these components as single use only. • Ensure that all impacted customers receive the correct instructions in their respective language. • Track customer responses and take appropriate actions to ensure compliance. Action to be taken by the Distributor/End-User: • Confirm receipt and thoroughly review the contents of the Customer Notification Package (including this notification and the Distributor/End-User Response Form). • Fully complete the attached Distributor/End-User Response Form and return it to GMB-EMEA-FSCA-RDX-INTL@jaegerrdx.com. The email subject line should be labeled “Response Form: FSCA-25-001". • We respectfully request the completed and signed Distributor/End-User Response Form to be returned no later than 01 OCT 2025 or within 30 days of receipt. | 2025-09-08 17:40
Siemens Healthcare Diagnostics Inc. | Atellica® CH UCFP | 241820 | - | Issue Description x D Siemens Healthineers previously issued an Urgent Field Safety Notice ACHC A OUS communicating the potential for falsely depressed patient quality control QC and or calibration results with the Atellica CH UCFP assay The initial Urgent Field Safety Notice included instructions to change the well status setting to Empty in certain scenarios This follow up communication provides a set of simplified instructions in the Customer Actions section below x D All future lots are impacted until further notice Siemens is working to determine root cause and restore the assay performance x D See Appendix for Detailed Customer Instructions x D x D Kindly review the attached | Customer Actions • Moving forward, keep only one (1) Atellica CH UCFP reagent pack onboard the system at a time. • Unload any Atellica CH UCFP packs that are currently onboard the system. • Siemens is instructing customers to perform QC on each well. Follow the instructions provided below: o Enable QC on Pack Change By Assay Type for CH to allow for QC to be processed when switching between wells. (See “Appendix” for detailed instruction) Note: This setting enables QC on Pack Change for all CH assays. All QC Levels for CH assays for which this is not required should be deselected in the QC Master List. • If your lab is using the Atellica to monitor QC: o Enable Patient QC Flagging to allow for impacted patient results to be identified after a QC failure. (See “Appendix” for detailed instruction) Note: All impacted patient results will have a “QC Fail” flag. Ensure that a result with a flag of “QC Fail” is held for review so that it can be rerun after passing QC. o Upon QC failure, manually reorder QC. o Verify that QC results are passing. Once acceptable QC has been obtained, repeat all UCFP testing for all patient results flagged with “QC Fail.” ▪ If QC results are still unacceptable after 2 attempts, unload the reagent pack and replace with a fresh pack. ▪ If you experience repeated QC issues on consecutive wells, follow your laboratory protocol for additional QC troubleshooting to evaluate potential alternative causes for QC failures. • If your lab uses the Atellica Data Manager or other middleware to monitor QC: o After a failed QC, ensure any UCFP samples processed after the failed QC are held for review. o Repeat QC. o Verify QC results are passing and repeat any held UCFP testing for patient results obtained since the failed QC. ▪ If QC results are still unacceptable after 2 attempts, unload the reagent pack and replace with a fresh pack. ▪ If you experience repeated QC issues on consecutive wells, follow your laboratory protocol for additional QC troubleshooting to evaluate potential alternative causes for QC failures. • As a result of these actions, track additional reagent consumption (number of tests) to report to Siemens Healthineers for future reimbursement/credit. • Please review this letter with your Medical Director to determine the appropriate course of action, including for any previously generated results, if applicable. • Complete and return the Field Correction Effectiveness Check Form attached to this letter within 30 days. • Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Kindly review the attached | 2025-09-09 12:15
GE Medical Systems Scs 283 Rue De La Miniere, BUC, 78530, France | AW Server SW | 77 | - | GE HealthCare has become aware of a security vulnerability in AW Server products If a complex password has not been configured for the AW Server and a malicious actor gains access to the network the AW Server is installed on the vulnerability could be exploited If this occurs the malicious actor could potentially compromise the confidentiality integrity and availability of patient data x D x D There have been no injuries reported to GE HealthCare as a result of this issue x D | Customer letter to be sent to all affected customers. Pre-install manual safety warning to be updated. Technical Service Manual instructions to be updated. CMT, pFMEA, MMSR to be updated. AW Configuration tool software to be updated to mandate complex password change & prevent weak passwords. Field Action to be initiated to change all Service password to complex passwords. | 2025-09-09 13:49
Draeger Medical, Inc. | Vista 120 | 1528 | - | Under certain conditions the monitor can freeze without an alarm or restart being triggered | The supplement sheet containing important information is printed out as necessary and sent to affected users together with an accompanying letter. Sending the information via email to the customers is preferred. These are requested to add the supplementary sheet to the existing instructions for use. | 2025-09-10 08:17
Baxter Medical Systems GmbH + Co. KG | Charging unit mounted W/T AC TS7500 | 1 | - | Root Cause x D The investigation identified two root causes x D The charging unit mounted W T AC TS does not fulfill current requirement s regarding protection class II according to IEC AMD AMD Medical electrical equipment Part General requirements for basic safety and essential performance x D The Design History File DHF for the charging unit is not complete missing information | Corrective Action(s): 1. An engineering change will be completed for the charging unit mounted W/T AC TS7500 (housing) to comply to protection class II according to IEC 60601-1:2005/AMD1:2012/AMD2:2020. 2. The DHF for the charging unit will be updated with the missing information. Field Correction: Baxter will replace the metal plate with a plastic plate, which will provide the insulation and protection needed. Actions to be Taken by Customers: 1. If wall-mounted charging units should malfunction or fail to start user should immediately stop using device and disconnect power source. 2. Contact Baxter to discuss service options. A device inspection will be scheduled, and replacement options will be discussed. 3. Complete the enclosed customer reply form and return it to Baxter | 2025-09-10 12:51
SCHILLER AG | ARGUS PB-3000 | 4 | - | Please find attached FSN letter | Please find attached FSN letter | 2025-09-11 12:13
Fritz Stephan GmbH | SOPHIE | 4 | - | Unfortunately PC V has been withdrawn due to another software error that will lead to the effect that changes of the alarm limits for FiO will no longer be transmitted to the MC As a result the effective FiO alarm limits remain at the initially configured values even though the display shows the changed values Changes of the FiO or FiO alarm limits will therefore lead to false positive FiO alarms x D However the FiO mixture and measurement are not affected x D Deviations of more than of the set FiO between the set FiO and the measured FiO will still generate an alarm independent from the adjustable alarm limits | Fritz Stephan GmbH has already corrected that issue in Software PC V4.5.5. If you have already implemented FSCA QM285 in one or more of your devices, please update them to PC V4.5.5 immediately. | 2025-09-11 13:36
SCHILLER Medical SAS | DEFIGARD HD-7 | 12 | - | Please find the attached FSN letter | Please find the attached FSN letter | 2025-09-11 14:41
Beckman Coulter Biotechnology ( Suzhou ) Co., LTD. | DxFLEX Flow Cytometer | 7 | - | Beckman Coulter has determined that the DxFLEX software version may crash due to a code error issue x D While using the DxFLEX Flow Cytometry instrument for data acquisition using user defined acquisition protocol the CytExpert for DxFLEX software version may unexpectedly crash Due to that the data acquisition may not be completed according to the preset conditions i e number of events and or acquisition time When the software crash is not detected promptly the sample may be completely consumed | Follow the steps below when running experiments in the instrument until the software is upgraded to the version that will fix the issue: a. Prepare the experiment file (create sample, set location, etc.) b. Switch to first “Tube” sheet for each task to be acquired c. Save and close the experiment file d. Open the experiment again (DO NOT switch to REPORT sheet after opening experiment) This step will prevent the software from crashing due to this code error issue. e. Perform acquisition f. Analyze data. • Beckman Coulter recommends posting this letter on or near the affected analyzers until the update to the software that addresses this issue is completed | 2025-09-12 23:38
Olympus | OLYMPUS BRONCHOFIBERSCOPE, BRONCHOVIDEOSCOPES | 211 | - | Olympus previously issued Field Corrective Actions in after investigating complaints of endobronchial combustion during therapeutic bronchoscopy Olympus provided details on ablation device compatibility provided additional recommendations regarding patient preparation and reiterated warnings in the Instruction for Use about injury and death resulting from improper device use Since the Field Corrective Action in as a part of continued investigation activities Olympus conducted additional assessment on the use of compatible bronchoscopes with laser argon plasma coagulation and high frequency therapy equipment Through these investigations Olympus has determined that additional IFU updates are necessary to provide further clarification on safe and effective use of the bronchoscopes when used in conjunction with these instruments This Correction supersedes the IFU addendum provided in x D x D | "1.Ensure all personnel are completely knowledgeable and thoroughly aware of the updates to the IFU contained in this letter when using the affected scopes. For products which are actively sold, please obtain the updated IFU through web portal www.olympus- europa.com. All updates to the IFU are enclosed with this letter. Discard any copies of the previous IFU and the addendum issued in 2023. If an updated IFU is not available, include a copy of this letter with your existing IFU. Discard any copies of the addendum issued in 2023. 2.*IFU will not be updated for this product as it is discontinued for Appendix 1 - List of Affected Devices" | 2025-09-14 11:22
Sopha Medical Vision (SMV) 105, Rue Morane Saulnien 78530BucÎle-de-France | Sopha Medical Vision (SMV) | 1 | - | GE HealthCare has become aware that certain Nuclear Medicine systems that are past the End of Guaranteed Service herein referred to as Impacted Systems and listed in the Affected Product Details section of this letter could have been transported or relocated without adequate detector support This could result in excessive stress being applied to the detector mounting mechanisms compromising their integrity If this occurs it can potentially result in a detector fall and lead to life threatening bodily injury x D x D There have been no reports of detector falls or injuries as a result of this potential issue x D x D The Impacted Systems affected by this notice are those that have not been manufactured for several years GE HealthCare has previously communicated End of Guaranteed Service for these Impacted Systems x D | Deploy a customer letter advising the consignees to stop use of their system(s), with no corrections provided by GEHC. | 2025-09-14 16:55
Smiths Medical | CADD-Solis Ambulatory Infusion Pump | 0 | - | update service manual | update service manual | 2025-09-14 17:04
werfen | ROTEM® sigma ROTROL P | 15 | - | x D Issue Description and Impact x D x D Werfen has identified that the above mentioned lot of ROTEM sigma ROTROL P has a higher than expected likelihood of generating quality control out of range results n INTEM C CT or HEPTEM C CT assays in conjunction with cartridge lots numbers between S and S | Based on the above, please take the following immediate actions: - If you are using the affected ROTEM sigma ROTROL P lot number identified above in your facility, please utilize the target range sheet received with this information for all ROTEM sigma cartridge complete and complete plus hep LOTs S240101 and higher for quality control measurements. - Document the acknowledgement on the Customer Reply Form and return the completed and signed form to the fax number or e-mail address listed. - Share this information with your clinical users and laboratory staff and follow your internal procedures. - Forward this notification to all affected locations within your facility. - Report all device-related suspected serious incidents to the manufacturer, distributor, local contact point, and, if appropriate, the National Competent Authority. - Retain a copy of this notification for your records. | 2025-09-15 12:12
Maquet Cardiopulmonary GmbH | Heater Unit HU 35, 230V | 17 | - | HU units are disinfected at the end of production however they are not produced in a sterile environment x D It is important to note that these units are not classified as sterile medical devices Both the Instruction for Use x D IFU and the Service Manual state this information x D Due to the recent relocation of the production site in June a proactive decision was made to manufacture x D additional HU units in advance to mitigate potential supply disruptions As a result the actual storage x D duration for the affected units extended to up to months x D The extended storage period increases the risk of bacterial growth While disinfection at the end of production x D reduces microbial load it may not eliminate all bacteria in the circuitry Bacteria present in the devices may x D proliferate over the prolonged storage time contributing to contamination x D In course of a manufacturer rework initiated on July HU units from stock were disinfected and x D sampled prior to distribution On August MCP became aware of a unit which was tested positive for x D mycobacteria after disinfection Later further units were tested positive for mycobacteria Further MCP has x D received four complaints regarding contaminations of the water circuit of HU devices that were part of the x D pre produced stock with Mycobacteria chimaera | For devices currently used in treatment: Upon termination of treatment, the customer shall take the HU 35 devices affected out of operation at the earliest opportunity and perform: The descaling process of the water circuit with 2% citric acid according to chapter 8.2.1 of the IFU (including subsequent 3 rinses of the device) The high-level disinfection process of the water circuit with a concentration of 5% Chloramine T according to chapter 8.2.3 of the IFU. Further, customers are recommended to perform testing of the units affected for NTM after performing this high-level disinfection, prior to any further patient use. Please note that by using common PCR analysis also dead NTM or DNA fragments of them lead to positive results. This does not necessarily indicate an inadequate disinfection. Therefore, MCP recommends performing cultural detection tests for the presence of NTM. If there are still atypical mycobacteria in the water system detectable, please repeat the high-level disinfection. Afterwards, follow the usual disinfection protocol as per IFU. For devices currently not used in treatment/ not installed: The customer shall immediately perform: The descaling process of the water circuit with 2% citric acid according to chapter 8.2.1 of the IFU (including subsequent 3 rinses of the device) The high-level disinfection process of the water circuit with a concentration of 5% Chloramine T according to chapter 8.2.3 of the IFU. Further, customers are recommended to perform testing of the units affected for NTM after performing this high-level disinfection, prior to any further patient use. Please note that by using common PCR analysis also dead NTM or DNA fragments of them lead to positive results. This does not necessarily indicate an inadequate disinfection. Therefore, MCP recommends performing cultural detection tests for the presence of NTM. If there are still atypical mycobacteria in the water system detectable, please repeat the high-level disinfection. Afterwards, follow the usual disinfection protocol as per IFU. | 2025-09-15 15:14
Abbott | TactiFlex™ Ablation Catheter, Sensor Enabled | - | - | The TactiFlex Ablation Catheter Sensor Enabled family of catheters are packaged in a long straight x D tray with a tip protector to prevent damage to the sensors in the catheter tip Improper removal of the x D catheter from the tip protector can potentially break or damage the tip In order to avoid damage to the x D tip prior to lifting the distal tip out of the tray one must ensure the distal tip has been fully retracted from x D the tip protector The current IFU requires inspection of the packaging and catheter for damage and or x D kinking prior to use | Continue following the product IFU by inspecting the packaging and catheter for damage and/or kinking prior to use.  Review the guidance in Appendix A information on proper removal technique of the TactiFlex™ Ablation Catheter, Sensor Enabled™ from the packaging tray, ensuring to inspect the electrodes and catheter tip carefully for integrity and overall condition prior to use.  Abbott has made changes to the packaging tray design to help mitigate tip damage during catheter removal from the packaging. The deployment of the new packaging has begun | 2025-09-15 16:20
TERUMO | RENZAN™ Peripheral Stent System | - | - | This notification is to inform you that we have initiated a voluntary recall of Terumo Neuro s x D RENZAN Peripheral Stent System We have received complaints relating to excessive x D resistance when attempting to unsheathe the stent causing potential failure of the delivery system x D which may lead to vessel damage and or difficulty removing the entire system | immediately stop using/distributing and quarantine all impacted devices | 2025-09-15 16:49
Ventana Medical Systems, Inc. | CONFIRM anti-Estrogen Receptor (ER) (SP1) Rabbit Monoclonal Primary Antibody | 187 | - | Investigation is ongoing under non conformance NC A preliminary internal investigation revealed a x D reduction in tumor stain intensity and percent tumor cell staining in weak staining breast carcinoma cases when x D testing specific batches of raw material Current raw material batches utilized in the production of finished goods x D show no evidence of light staining A review of the historical data did not identify any discernible patterns or x D consistent trends in titer values | Roche is requesting that customers discontinue the use of and discard any remaining inventory of specific lots (listed in TABLE A below) of CONFIRM anti-Estrogen Receptor (ER) (SP1) Rabbit Monoclonal Primary Antibody (05278414001 / 790-4325) and CONFIRM anti-Estrogen Receptor (ER) (SP1) Rabbit Monoclonal Primary Antibody (05278406001 / 790-4324) (collectively referred to as “CONFIRM anti-Estrogen Receptor (ER) (SP1) Rabbit Monoclonal Primary Antibody”). | 2025-09-16 10:01
Getinge | CARDIOSAVE IABP HYBRID AND RESCUE | 147 | - | Datascope Getinge determined that the exterior of the Cardiosave Hybrid and Rescue IABP x D may be susceptible to fluid ingress at specific locations on the device IABPs contain various x D electronic circuit boards Liquid spills such as saline can create bridges of resistance x D between the circuit components causing the circuit to not function as intended This can x D impact initiation or continuation of counterpulsation therapy | Datascope/Getinge has developed a solution to address this potential issue. Three (3) hardware kits will be released to update the Cardiosave IABP Hospital Cart, IABP Display, and to provide one (1) updated Transport Protective Cover for the system when used in Rescue/Transport configuration. | 2025-09-16 10:47
Getinge | CARDIOSAVE Hybrid and CARDIOSAVE Rescue Intra-Aortic Balloon Pumps | 147 | - | Datascope Getinge received complaints reporting unexpected shutdown of the Cardiosave IABP x D while providing therapy in very rare instances An unexpected shutdown occurs without forewarning x D the screen abruptly turns off no further instruction or status is available to the user and a highpitched x D x D alarm is emitted An internal investigation of the complaints determined an unexpected x D shutdown may be due to blood entering into the Cardiosave IABP when therapy is provided with a x D perforated IAB catheter | A Datascope/Getinge SSU representative will schedule a service visit to impacted customers to install the updated software (D.03), and hardware update to PIM and test the Cardiosave IABP to ensure it functions as intended prior to releasing it for clinical use. Also, the affected customers may contact Getinge to request a service visit to address the hardware and software updates. | 2025-09-16 10:51
Biomerieux | VIDAS® HIV DUO AG/AB | 14 | - | bioM rieux received customers complaints who reported that after launching a calibration of the assay VIDAS HIV DUO AG AB Ref on VIDAS KUBE Ref the instrument did not calculate the result and showed an error Impossible to compute result After this error occurs the customer could not use the assay VIDAS HIV DUO AG AB and there is no available workaround for the customer | • Workarounds are already available to prevent the issue at your level or fix it if it already occure until the new Knowledge Base that solves the issue will be at your disposal by the end of August 2025 | 2025-09-16 11:08
Biomerieux | VIDAS® CK-MB | 104 | - | bioM rieux identified during internal monitoring a drift of the S Relative Fluorescence Value RFV on VIDAS CK MB Reference lots and This can lead to an invalid calibration out of range high S values | The manufacturing of a new raw material has been launched to be able to release conforming lots | 2025-09-16 11:21
Beckman Coulter, Inc. | DxC 500i Clinical Analyzer DxC 500i Access 2 Immunoasssay Module | 20 | - | A sample is presented to the analyzer from the sample handler or x D the sample carousel x D While tests associated with the sample are in progress x D The sample is manually diluted x D Additional tests are ordered through the Test Order Entry page x D and a dilution factor is entered x D The sample ID used for the original test requests is also used x D for the tests ordered with the diluted sample x D When tests ordered for the diluted sample are complete the user x D interface correctly displays the dilution factor that was entered on x D the Test Order Entry page However the analyzer does not apply x D the dilution factor to the result | • When ordering tests for a manually diluted sample through the Test Order Entry page, a unique sample ID is required. • Alternatively, use the Dilute and Rerun feature or an automated dilution to measure diluted samples. Refer to Rerunning a Test with a Diluted Sample in the System Help for more information. • Beckman Coulter suggests sharing this letter with your Medical Director to determine if performing a retrospective review of results is necessary | 2025-09-17 09:09
Beckman Coulter, Inc. | DxI 9000 Access Immunoassay Analyzer | 20 | - | A sample is presented to the analyzer from the sample handler or x D the sample carousel x D While tests associated with the sample are in progress x D The sample is manually diluted x D Additional tests are ordered through the Test Order Entry page x D and a dilution factor is entered x D The sample ID used for the original test requests is also used x D for the tests ordered with the diluted sample x D When tests ordered for the diluted sample are complete the user x D interface correctly displays the dilution factor that was entered on x D the Test Order Entry page However the analyzer does not apply x D the dilution factor to the result | • When ordering tests for a manually diluted sample through the Test Order Entry page, a unique sample ID is required. • Alternatively, use the Dilute and Rerun feature or an automated dilution to measure diluted samples. Refer to Rerunning a Test with a Diluted Sample in the System Help for more information. • Beckman Coulter suggests sharing this letter with your Medical Director to determine if performing a retrospective review of results is necessary | 2025-09-17 09:09
CareFusion 303, Inc. | BD Alaris™ Pump | - | - | BD is providing this notice to inform customers of worst case fluid delivery performance under certain use cases for the BD Alaris Pump Module model pump module when used with a subset of compatible pump infusion sets x D x D The User Manuals for BD Alaris Systems outside North America discloses the large volume pump performance flow rate and bolus accuracy time to alarm occlusions and post occlusion bolus volume in accordance with IEC that include performance of the device delivering infusates through representative infusion sets under standard operating conditions and describe and warn users regarding all operating conditions and factors that can deviate that performance x D x D In North America the more recently issued AAMI TIR standard guides infusion pump vendors on testing and data disclosure requirements This AAMI TIR standard requires characterisation and publication of data also under worst case conditions including testing the edges of the operating conditions with the infusion set that is expected to yield the largest degradation of pump performance x D x D BD recently initiated a notice to customers in North America regarding published worst case performance for a subset of infusion sets that include in line filters back check valves and multiple injections ports Infusion sets with these design features may impact the fluid delivery performance particularly at programmable rates under mL h and at or under mL of volume The most notable performance variations have been observed with the micron in line filter infusion set at flow rates below mL h x D x D Considering the BD Alaris Systems are also distributed outside of North America to ensure consistency across regions BD is providing this performance data and associated warnings in this notification to better inform clinical practice for your region x D x D Note this data does not impact the performance claims published currently in your BD Alaris System User Manual v x D x D Following a review of the sets that have been supplied in your region in the last years shelf life of the sets and based on the data published in North America BD has identified a subset of sets that have the potential to perform per Appendix A x D x D In this letter BD is sharing the latest performance data reiterating existing warnings and mitigations from User Manuals and requesting custom | Actions taken by BD: 1. BD is contacting all impacted customers to notify them of this Field Action. 2. BD is providing a list of alternative sets, referenced under the “Actions for Clinical Users” section of this notice. 3. BD Clinical Consultants can help with identifying alternative sets, help review customer’s infusion data analytics, and quantify types of infusions infused with specific rate types. Customer Actions: • Review the information in Table 1 and in Appendix B to determine if the infusion sets in your possession are impacted. • Read and apply the “Actions for Clinical Users” section above in its entirety. • Refer to Appendix A for the set performance data. For ease of reference, the performance data in Appendix A should be compared to the specifications provided in the administration set User Manuals. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 19th October 2025. • This notice needs to be passed on to all those who need to be aware within your organisation or to any organisation where the potentially affected devices have been transferred. • Please maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. • Please report all device-related incidents to the manufacturer, distributor or local representative, and the national Competent Authority if appropriate, as this provides important feedback. Distributor Actions: • Identify the facilities where you have distributed an Alaris™ Pump Module and infusions sets in Appendix B and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 19th October 2025. o There is no requirement to return your customer response forms to BD, you should maintain these on file at your facility. Return only your final consolidated response form. • Complete and return the Customer Response Form following completion of your reconciliation activities. • Please maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. • Please report all device-related incidents to the manufacturer, or local representative, and the national Competent Authority if appropriate, as this provides important feedback. | 2025-09-21 16:54
Alcon Laboratories Inc. | UltraVit | 15280 | - | Alcon has received complaints of CONSTELLATION ULTRAVIT K and HYPERVIT K probes that unexpectedly failed to actuate and cut during use Globally reports of CONSTELLATION ULTRAVIT K and HYPERVIT K failing to actuate and cut during use represent less than of CONSTELLATION ULTRAVIT and HYPERVIT probes sold from January through August x D Our investigation determined that a portion of ULTRAVIT and HYPERVIT probes were manufactured with a component received from a supplier that did not perform as intended which can lead to increased friction within the probe engine The increased friction may cause premature actuation failure resulting in failure to cut x D x D Alcon detected an increased trend of adverse event AE reports associated with CONSTELLATION ULTRAVIT K probes related to this issue From January through September Alcon has received a global total of complaints associated with CONSTELLATION ULTRAVIT K probes for failure to cut that included adverse events AEs There is no increase in reported adverse events related to CONSTELLATION HYPERVIT K probes | To mitigate the potential risk associated with unexpected cut stoppage during surgery, Alcon is advising customers to take the following precautions: 1. Strictly adhere to the Instructions for Use (IFU) regarding limitations, including maximum use time and single use only. 2. When using a probe model identified as within the scope of the Medical Device Correction: a. Reduce the actuation rate to no more than 5000 actuations per minute b. Adjust other related console settings as needed to accommodate the reduced cutting/actuation rate c. If a reduction of cutting capability or failure to actuate is observed during the surgical procedure, stop immediately. Take all necessary surgical precautions to remove the probe and replace it with an unused probe. | 2025-09-21 17:35
Alcon Laboratories Inc. | TOTALPLUS | 15280 | - | Alcon has received complaints of CONSTELLATION ULTRAVIT K and HYPERVIT K probes that unexpectedly failed to actuate and cut during use Globally reports of CONSTELLATION ULTRAVIT K and HYPERVIT K failing to actuate and cut during use represent less than of CONSTELLATION ULTRAVIT and HYPERVIT probes sold from January through August x D Our investigation determined that a portion of ULTRAVIT and HYPERVIT probes were manufactured with a component received from a supplier that did not perform as intended which can lead to increased friction within the probe engine The increased friction may cause premature actuation failure resulting in failure to cut x D x D Alcon detected an increased trend of adverse event AE reports associated with CONSTELLATION ULTRAVIT K probes related to this issue From January through September Alcon has received a global total of complaints associated with CONSTELLATION ULTRAVIT K probes for failure to cut that included adverse events AEs There is no increase in reported adverse events related to CONSTELLATION HYPERVIT K probes | To mitigate the potential risk associated with unexpected cut stoppage during surgery, Alcon is advising customers to take the following precautions: 1. Strictly adhere to the Instructions for Use (IFU) regarding limitations, including maximum use time and single use only. 2. When using a probe model identified as within the scope of the Medical Device Correction: a. Reduce the actuation rate to no more than 5000 actuations per minute b. Adjust other related console settings as needed to accommodate the reduced cutting/actuation rate c. If a reduction of cutting capability or failure to actuate is observed during the surgical procedure, stop immediately. Take all necessary surgical precautions to remove the probe and replace it with an unused probe. | 2025-09-21 17:35
Alcon Laboratories Inc. | HyperVit | 15280 | - | Alcon has received complaints of CONSTELLATION ULTRAVIT K and HYPERVIT K probes that unexpectedly failed to actuate and cut during use Globally reports of CONSTELLATION ULTRAVIT K and HYPERVIT K failing to actuate and cut during use represent less than of CONSTELLATION ULTRAVIT and HYPERVIT probes sold from January through August x D Our investigation determined that a portion of ULTRAVIT and HYPERVIT probes were manufactured with a component received from a supplier that did not perform as intended which can lead to increased friction within the probe engine The increased friction may cause premature actuation failure resulting in failure to cut x D x D Alcon detected an increased trend of adverse event AE reports associated with CONSTELLATION ULTRAVIT K probes related to this issue From January through September Alcon has received a global total of complaints associated with CONSTELLATION ULTRAVIT K probes for failure to cut that included adverse events AEs There is no increase in reported adverse events related to CONSTELLATION HYPERVIT K probes | To mitigate the potential risk associated with unexpected cut stoppage during surgery, Alcon is advising customers to take the following precautions: 1. Strictly adhere to the Instructions for Use (IFU) regarding limitations, including maximum use time and single use only. 2. When using a probe model identified as within the scope of the Medical Device Correction: a. Reduce the actuation rate to no more than 5000 actuations per minute b. Adjust other related console settings as needed to accommodate the reduced cutting/actuation rate c. If a reduction of cutting capability or failure to actuate is observed during the surgical procedure, stop immediately. Take all necessary surgical precautions to remove the probe and replace it with an unused probe. | 2025-09-21 17:35
Diagnostic Grifols, S.A | Erytra | 14 | - | Incorrect results can be assigned for an autocancelled and rescheduled test A correct result is obtained for the rescheduled test while no results are obtained for the auto cancelled one The interpretation of the rescheduled test is shown in the results detail screen and traceability report However the system shows a wrong result interpretation in the results screen and this result is sent to the LIS x D The issue occurs due to a very rare concurrency problem in the database management x D Three conditions must be met to trigger the incident x D x D A test is automatically cancelled and rescheduled x D x D This test is rescheduled while its status is already RUNNING x D x D The software is not able to successfully discard the cancelled test x D Due to this error the test remains internally linked to two different instances of the technique the cancelled one and the one that is created upon rescheduling The microtube results from the cancelled execution have no information but the software assigns a default interpretation result The information contained in the executed instance is correct However the interpretation result shown in the results screen for this technique is from the cancelled technique instance | If you have installed software version 4.5.0, you are requested to install patch v4.5.1 on your Erytra analyzer, which is already available and can be remotely installed, or to install software version v4.6.1, which will be available for installation in the next weeks. If you have installed software version 4.6.0, you are requested to install software version 4.6.1 on your Erytra analyzer, which will be available for installation in the next weeks. | 2025-09-22 09:50
Beckman Coulter, Inc. | DxC 500i Clinical Analyzer DxC 500i Access 2 Immunoasssay Module | 17 | - | Beckman Coulter has determined that the analyzer may fail to apply a dilution factor to tests associated with a manually diluted sample when the tests are ordered through the Test Order Entry page and a unique sample ID is not used x D This issue only occurs with Access Immunoassay tests Chemistry tests are not impacted x D The issue occurs under the following conditions x D A sample is presented to the analyzer from the sample handler or the sample carousel x D While tests associated with the sample are in progress x D The sample is manually diluted x D Additional tests are ordered through the Test Order Entry page and a dilution factor is entered x D The sample ID used for the original test requests is also used for the tests ordered with the diluted sample x D When tests ordered for the diluted sample are complete the user interface correctly displays the dilution factor that was entered on the Test Order Entry page However the analyzer does not apply the dilution factor to the result x D Automated dilutions and tests ordered using the Dilute and Rerun feature are not affected by this issue | • When ordering tests for a manually diluted sample through the Test Order Entry page, a unique sample ID is required. • Alternatively, use the Dilute and Rerun feature or an automated dilution to measure diluted samples. Refer to Rerunning a Test with a Diluted Sample in the System Help for more information. • Beckman Coulter suggests sharing this letter with your Medical Director to determine if performing a retrospective review of results is necessary. | 2025-09-22 12:05
MicroVention Europe SARL | FRED X | 1 | - | manufacturing issue related to the tantalum length and or tantalum attachment pattern | تم شحن المواد الى الاردن اذن تسيلم رقم 138 و بوليص رقم 5194073220بتاريخ 5/3/2025 | 2025-09-25 11:04
Medtronic Sofamor Danek USA Inc. | Catalyft™ PL Expandable erbody System erbody Cage, PL40 | 1684 | - | An issue was escalated for evaluation to the Spine post market issue escalation triage board regarding reported events of Catalyft PL cages losing height postoperatively x D The postoperative cage loss of lordosis is the result of the device experiencing lack of load as a result of improper loading of the device during implantation or the device experiencing lack of load as a result of subsidence and or migration post implantation x D Evaluated complaints of cage loss of lordosis reported asymptomatic findings detected on routine post operative imaging Some patients have undergone revision surgery due to pain or pseudoarthrosis x D The use of a Catalyft expandable interbody system without following the device s IFU and surgical technique could result in potential adverse events such as loss of lordosis migration subsidence and or neurological injury x D | • Share this notification within your facility with clinical users of the Catalyft™ PL & PL40 system. • Complete the enclosed Customer Acknowledgement Form. | 2025-09-25 16:21
CareDx Pty Ltd. | AlloSeq® cfDNA | 7 | - | The root cause has been identified to be linked to the raw material used in the AlloSeq cfDNA PCR Mix components of these lots The issue has been identified to be isolated to these lots based on shared ingoing components All ingoing components have since been replaced hence the issue is isolated to these lots | Removal of lot 24-5400 from the market | 2025-09-25 16:47
Biomedical Data Solutions Ltd. | Connect 3.2.3.2 | 2 | - | An Assay Rule was created for Profile Anti A that was defined with assay Anti A has a grade between and x D to Hold Pending Review and to Notify x D When the result for Anti A was received with a grade of the result did not trigger the rule as was expected | A resolution to this issue is being provided in a future software update for ORTHO CONNECT™ to permanently fix this issue. | 2025-09-28 10:53
Codman & Shurtleff, Inc | CODMAN® Disposable Perforator | 1031 | - | On April Integra issued a voluntary recall for the mm Codman Disposable Perforator due to an inadequate weld proud weld on specific lots of mm perforator that can potentially cause the product to disassemble break separate x D To date there have been no complaints or adverse events reported in relation to the mm and mm Codman Disposable Perforators disassembly Out of an abundance of caution as the and mm perforators are manufactured using the same equipment and operating parameters that affected the mm Perforator Integra has made the decision to expand the scope of the field action to include the mm and mm perforators as noted in Table | If you have affected product(s) in your warehouse: a. Quarantine them immediately. | 2025-09-28 15:12
Beckman Coulter Inc. | UniCel DxI 600 Access Immunoassay System | 26 | - | An internal investigation has determined that the pick and place PnP gantries in the UniCel DxI analyzers listed in Appendix A may experience increased physical interference due to sub components within the gantry assembly This interference can result in more frequent Z motion errors in the Wash PnP and or Incubator PnP | If Wash PnP and/or Incubator PnP Z-motion errors appear in the analyzer event log, take the following steps: 1. Attempt to reinitialize the analyzer: a. Go to the Diagnostics screen on the User Interface. b. Select Initialize System. 2. If reinitialization is unsuccessful, contact Beckman Coulter Technical Support in your region. Beckman Coulter recommends posting this letter on or near the affected analyzers. | 2025-09-28 16:27
Beckman Coulter Inc. | UniCel DxI 800 Access Immunoassay System | 26 | - | An internal investigation has determined that the pick and place PnP gantries in the UniCel DxI analyzers listed in Appendix A may experience increased physical interference due to sub components within the gantry assembly This interference can result in more frequent Z motion errors in the Wash PnP and or Incubator PnP | If Wash PnP and/or Incubator PnP Z-motion errors appear in the analyzer event log, take the following steps: 1. Attempt to reinitialize the analyzer: a. Go to the Diagnostics screen on the User Interface. b. Select Initialize System. 2. If reinitialization is unsuccessful, contact Beckman Coulter Technical Support in your region. Beckman Coulter recommends posting this letter on or near the affected analyzers. | 2025-09-28 16:27
Elekta Instrument AB | LEKSELL VANTAGE ARC SYSTEM | 3 | - | For the specified serial numbers there is an increased risk that paint flakes from the engravings may separate from the Leksell Vantage Arc System components Separation of paint flakes is likely to occur during cleaning and sterilization If paint flakes are not washed away during the cleaning and sterilization process these could remain on the Leksell Vantage Arc System when used in surgery This poses a risk of paint flakes entering the surgical area This situation has not been observed in clinical practice but the risk should be fully considered and mitigated The paint does not include any metallic components and is not magnetic | Elekta will send an Important Field Safety Notice 100-01-301-011 to all affected customers from 30-SEP-25. The notice informs users of the specific product and version numbers affected by the issue,and any work around that can be used to avoid the issue. • Sites affected will be customers with Arc and Support Arc with serial numbers: SH00001 - SH00338 A copy of the Important Field Safety Notice is to be kept with the most current labelling and all personnel working with the product should be made aware of the content of the letter. The customers are instructed to complete and return the Important Field Safety Notice Acknowledgement form to their local Elekta representative as soon as possible. | 2025-09-29 11:12
Abbott GmbH. & Co. KG | Alinity hq Analyzer | 168 | - | Please find the attached report | Please find the attached report | 2025-09-29 15:06
GE Medical Systems Information Technologies, Inc | CARESCAPE ONE | 196 | - | GE HealthCare has become aware of the potential for loss of monitoring on CARESCAPE B CARESCAPE ONE B x M B x P VSP B x M B x P VSP Patient Monitors and Portrait Vital Signs Monitors when powered by certain batteries Affected batteries can lose their capability to hold charge and provide power see Affected Product Details below If this situation occurs it will cause loss of patient monitoring and potentially delay recognizing patient status changes requiring treatment x D x D There have been no injuries reported to GE HealthCare as a result of this issue x D | Affected battery replacement by GE Healthcare Service Representative. | 2025-09-30 06:14
GE Healthcare Finland Oy | Carescape Monitor B450 | 196 | - | GE HealthCare has become aware of the potential for loss of monitoring on CARESCAPE B CARESCAPE ONE B x M B x P VSP B x M B x P VSP Patient Monitors and Portrait Vital Signs Monitors when powered by certain batteries Affected batteries can lose their capability to hold charge and provide power see Affected Product Details below If this situation occurs it will cause loss of patient monitoring and potentially delay recognizing patient status changes requiring treatment x D x D There have been no injuries reported to GE HealthCare as a result of this issue x D | Affected battery replacement by GE Healthcare Service Representative. | 2025-09-30 06:14
GE Medical Systems Information Technologies, Inc | B1X5M | 196 | - | GE HealthCare has become aware of the potential for loss of monitoring on CARESCAPE B CARESCAPE ONE B x M B x P VSP B x M B x P VSP Patient Monitors and Portrait Vital Signs Monitors when powered by certain batteries Affected batteries can lose their capability to hold charge and provide power see Affected Product Details below If this situation occurs it will cause loss of patient monitoring and potentially delay recognizing patient status changes requiring treatment x D x D There have been no injuries reported to GE HealthCare as a result of this issue x D | Affected battery replacement by GE Healthcare Service Representative. | 2025-09-30 06:14
GE Medical Systems Information Technologies, Inc | Portrait VSM | 196 | - | GE HealthCare has become aware of the potential for loss of monitoring on CARESCAPE B CARESCAPE ONE B x M B x P VSP B x M B x P VSP Patient Monitors and Portrait Vital Signs Monitors when powered by certain batteries Affected batteries can lose their capability to hold charge and provide power see Affected Product Details below If this situation occurs it will cause loss of patient monitoring and potentially delay recognizing patient status changes requiring treatment x D x D There have been no injuries reported to GE HealthCare as a result of this issue x D | Affected battery replacement by GE Healthcare Service Representative. | 2025-09-30 06:14
GE Medical Systems Information Technologies, Inc | CARESCAPE ONE | 196 | - | GE HealthCare has become aware of the potential for loss of monitoring on CARESCAPE B CARESCAPE ONE B x M B x P VSP B x M B x P VSP Patient Monitors and Portrait Vital Signs Monitors when powered by certain batteries Affected batteries can lose their capability to hold charge and provide power see Affected Product Details below If this situation occurs it will cause loss of patient monitoring and potentially delay recognizing patient status changes requiring treatment x D x D There have been no injuries reported to GE HealthCare as a result of this issue x D | Affected battery replacement by GE Healthcare Service Representative. | 2025-09-30 06:14
GE Healthcare Finland Oy | Carescape Monitor B450 | 196 | - | GE HealthCare has become aware of the potential for loss of monitoring on CARESCAPE B CARESCAPE ONE B x M B x P VSP B x M B x P VSP Patient Monitors and Portrait Vital Signs Monitors when powered by certain batteries Affected batteries can lose their capability to hold charge and provide power see Affected Product Details below If this situation occurs it will cause loss of patient monitoring and potentially delay recognizing patient status changes requiring treatment x D x D There have been no injuries reported to GE HealthCare as a result of this issue x D | Affected battery replacement by GE Healthcare Service Representative. | 2025-09-30 06:14
GE Medical Systems Information Technologies, Inc | B1X5M | 196 | - | GE HealthCare has become aware of the potential for loss of monitoring on CARESCAPE B CARESCAPE ONE B x M B x P VSP B x M B x P VSP Patient Monitors and Portrait Vital Signs Monitors when powered by certain batteries Affected batteries can lose their capability to hold charge and provide power see Affected Product Details below If this situation occurs it will cause loss of patient monitoring and potentially delay recognizing patient status changes requiring treatment x D x D There have been no injuries reported to GE HealthCare as a result of this issue x D | Affected battery replacement by GE Healthcare Service Representative. | 2025-09-30 06:14
GE Medical Systems Information Technologies, Inc | Portrait VSM | 196 | - | GE HealthCare has become aware of the potential for loss of monitoring on CARESCAPE B CARESCAPE ONE B x M B x P VSP B x M B x P VSP Patient Monitors and Portrait Vital Signs Monitors when powered by certain batteries Affected batteries can lose their capability to hold charge and provide power see Affected Product Details below If this situation occurs it will cause loss of patient monitoring and potentially delay recognizing patient status changes requiring treatment x D x D There have been no injuries reported to GE HealthCare as a result of this issue x D | Affected battery replacement by GE Healthcare Service Representative. | 2025-09-30 06:14
Abiomed Inc. | Automated Impella Controller (AIC) | 40 | - | Abiomed identified a design change to correct the AIC purge retainer fixture that will be implemented through device servicing at the Abiomed Service Center This design change involves the addition of a third mounting screw for the purge retainer improving fixation The product is not being removed and hospital inventory may continue to be used x D | "• Hospital inventory can continue to be used to ensure continuity of care. • Upon contact from Abiomed’s field servicing team, customers are asked to please work with them to return the identified device(s) for the change to be implemented. Customers are referred to Customer Reply form for list of impacted AIC serial numbers. • Customers are asked to review this notice carefully, and forward to anyone in their facility that needs to be informed (i.e., those who manage, transport, store, stock, or use the subject products). • If any of the subject products have been forwarded to another facility, customers should contact that facility and provide them with this notice." | 2025-09-30 17:26
SAMSUNG ELECTRONICS Co., Ltd. | SAMSUNG DIGITAL DIAGNOSTIC MOBILE X-RAY SYSTEM | 3 | - | this may be caused by metal fatigue resulting from accumulated stress over a long period of use | Samsung Electronics will carry out the field safety corrective action for all affected systems produced between October 2016 to July 2018. The corrective action is to reinforce the arm with a high tensile bolt to eliminate the risk of the THU falling and preventing any potential risk of injury. Samsung service engineers will correct all affected systems free of charge and contact you to arrange for the correction. | 2025-10-02 09:17
Boston Scientific Corporation | AVVIGO™+ Multi-Modality Guidance System | 2 | - | Boston Scientific has received reports indicating the grid mark overlay on the cross sectional view did not present correctly when an OptiCross Peripheral Imaging Catheter was connected to the AVVIGO Multi Modality Guidance System while in either Live or Record mode The Boston Scientific investigation determined that due to a software anomaly a grid mark overlay is incorrectly displayed during Live mode when AVVIGO is connected to an OptiCross Peripheral Imaging Catheter Note that the correct overlay display for this catheter is the grid mark overlay which allows for a reference depth setting of mm for peripheral catheters x D This software anomaly is limited to Live or Record mode In Review mode the correct grid mark overlay is displayed and the vessel image is correctly scaled to the appropriate size | Software update | 2025-10-02 10:29
Philips Medical Systems DMC GmbH | CombiDiagnost R90 | 9 | - | Philips has identified that some ceiling suspension telescopic carriages manufactured from July through June may be difficult to move vertically i e additional force physical exertion required for manual movement due to potential manufacturing issues | A Philips representative will contact you to schedule a time for a Field Service Engineer (FSE) to visit your site to inspect the ceiling suspension and implement a follow-up hardware correction, if required (reference 2025-PD-DXR-008) | 2025-10-02 12:27
Philips Medical Systems DMC GmbH | DigitalDiagnost C90 | 9 | - | Philips has identified that some ceiling suspension telescopic carriages manufactured from July through June may be difficult to move vertically i e additional force physical exertion required for manual movement due to potential manufacturing issues | A Philips representative will contact you to schedule a time for a Field Service Engineer (FSE) to visit your site to inspect the ceiling suspension and implement a follow-up hardware correction, if required (reference 2025-PD-DXR-008) | 2025-10-02 12:27
Medlstim ASA | Medistim MiraQ System | 1 | - | Medistim has identified an issue with the printer support module in certain MiraQ Systems that may cause the device to shut down unexpectedly This problem is limited to the models and serialnumbers listed above x D | Change the Printer support module at gulf medical warehouse | 2025-10-02 13:17
Randox Laboratories Ltd | Evidence Investigator Analyser | 19 | - | Randox Laboratories have updated the Evidence Investigator Operator Manual to include a section on Cyber Security informing customers of security and cybersecurity recommendations Please see attached Technical Bulletin and updated manual | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to support@randoxtoxicology.com within five working days. If you have any questions or concerns, please contact Randox Technical Services. | 2025-10-06 10:09
Becton Dickinson and Company | BD BACTEC™ MGIT™ 960 STR 4.0 Kit | 1 | - | BD has identified internally that the BD BACTEC MGIT Streptomycin Solution vial material JAA batch supplied as a component in the BD BACTEC MGIT Streptomycin Kit listed in Table has been labeled with an incorrect expiration date of April The correct expiration date for this product is May While the kit carton displays the correct expiration date the individual vial label does not | Clinical User Actions 1. Customers should immediately discard the affected vial of lyophilized streptomycin from the BD BACTEC™ MGIT™ 960 Streptomycin 4.0 Kit. Note: No other components in the kit are required to be discarded. 2. If this product has already been used, no additional patient follow-up actions are required. BD Actions: BD has identified the root cause and is taking corrective actions to prevent recurrence of this issue. To date, BD does not plan to initiate any further advice or information in a follow-up FSN. Customer Actions: 1. Cease use of any unused affected vials of lyophilized streptomycin from the BD BACTEC™ MGIT™ 960 Streptomycin 4.0 Kit as listed in Table 1. 2. Identify and quarantine all unused affected BD BACTEC™ MGIT™ 960 Streptomycin 4.0 Kit. 3. Make a note of the lot numbers and destroy all unused affected vials of lyophilized streptomycin. 4. Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 02nd November 2025. 5. This notice needs to be passed on to all those who need to be aware within your organisation or to any organisation where the potentially affected devices have been transferred. 6. Please transfer this notice to other organisations on which this action has an impact. 7. Please maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. 8. Please report all device-related incidents to the manufacturer, distributor or local representative, and the national Competent Authority if appropriate, as this provides important feedback. Distributor Actions: • Cease distribution. • Identify, quarantine, making a note of the lot numbers then destroy all undistributed affected BD BACTEC™ MGIT™ 960 Streptomycin 4.0 Kits. • This notice needs to be passed on to all those who need to be aware within your organisation or to any organisation where the potentially affected devices have been transferred. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 02nd November 2025 o There is no requirement to return your customer response forms to BD, you should maintain these on file at your facility. Return only your final consolidated response form. • Complete and return the Customer Response Form following completion of your reconciliation activities. • Please maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. • Please report all device-related incidents to the manufacturer, or local representative, and the national Competent Authority if appropriate, as this provides important feedback. | 2025-10-06 10:25
Randox Laboratories Ltd | RX Monaco | 5 | - | Reason for Notification x D Randox Laboratories can confirm that the RX Monaco RX RX R has now been discontinued x D Randox will continue to provide technical support service engineering support and spare parts for x D existing devices until st December | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. If you have any questions or concerns, please contact Randox Technical Services. | 2025-10-06 10:59
Randox Laboratories Ltd | RX Daytona + with ISE | 103 | - | Randox Laboratories have updated the following analyser manuals to include further information on maintenance and prevention of cross contamination carryover avoidance Please refer to the technical bulletins listed below for the complete details | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Once received, the updated pages from the manuals will be shared. The compete updated Rx manuals are also available in electronic format on request. If you have any questions or concerns, please contact Randox Technical Services. | 2025-10-06 11:07
Randox Laboratories Ltd | RX Daytona + without ISE | 103 | - | Randox Laboratories have updated the following analyser manuals to include further information on maintenance and prevention of cross contamination carryover avoidance Please refer to the technical bulletins listed below for the complete details | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Once received, the updated pages from the manuals will be shared. The compete updated Rx manuals are also available in electronic format on request. If you have any questions or concerns, please contact Randox Technical Services. | 2025-10-06 11:07
Randox Laboratories Ltd | RX Imola | 103 | - | Randox Laboratories have updated the following analyser manuals to include further information on maintenance and prevention of cross contamination carryover avoidance Please refer to the technical bulletins listed below for the complete details | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Once received, the updated pages from the manuals will be shared. The compete updated Rx manuals are also available in electronic format on request. If you have any questions or concerns, please contact Randox Technical Services. | 2025-10-06 11:07
Randox Laboratories Ltd | RX Modena | 103 | - | Randox Laboratories have updated the following analyser manuals to include further information on maintenance and prevention of cross contamination carryover avoidance Please refer to the technical bulletins listed below for the complete details | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Once received, the updated pages from the manuals will be shared. The compete updated Rx manuals are also available in electronic format on request. If you have any questions or concerns, please contact Randox Technical Services. | 2025-10-06 11:07
Becton Dickinson and Company | Product Cybersecurity - Expansion | 2 | - | Through our cybersecurity monitoring tools and processes BD recently identified unauthorised access to a limited portion of its information technology IT environment BD terminated the unauthorised access and applied additional security measures After a thorough investigation BD confirmed that product service credentials intended for use by BD technical support teams for certain BD products were accessed by an unauthorised actor Until these product service credentials are updated there is a risk of unauthorised access that may impact the confidentiality integrity and or availability of the relevant products and associated system or data x D x D For those BD Diagnostic Solutions products listed in the customer letter they are only vulnerable if a threat actor is either physically at the device and or has breached your hospital or laboratory s local network x D x D To date BD has not been made aware of any unauthorised use of these product service credentials and has received no reports of these credentials being used for unauthorised access to any BD device x D | BD Actions: BD has implemented additional security measures to further strengthen its IT environment. Actions to be taken by BD: BD Regional Support teams will be working to contact affected customers to discuss the next steps to be taken, which may include updating credentials and/or software. Customer Actions: • Review the products listed in Table 1 to determine if the products in your possession are impacted. • Complete and return the Customer Response Form even if you no longer use/have the impacted product as listed in Table 1 in your facility by 21st October 2025. • Circulate this notification to all those who need to be aware within your organisation or to any organisation where the potentially affected product has been transferred. • If you experience any issues, please report as a complaint as per your normal process. Distributor Actions: • Identify the facilities where you have distributed affected product and notify them immediately of this notification. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 21st October 2025. • Complete and return the Customer Response Form following completion of your reconciliation activities. • If you experience any issues, please report as a complaint as per your normal process. | 2025-10-06 12:21
Cardinal Health 200, LLC | Salem Sump™ | - | - | Cardinal Health received complaints regarding the Anti Reflux Valve ARV breaking due to use x D error The ARV failures are attributed to excessive force during use Based on this Cardinal x D Health is providing additional guidance when using these products | 1. REVIEW your inventory for the affected product codes. All lots of inventory are impacted. 2. COMMUNICATE the risks and appropriate usage of the ARV with all personnel. 3. POST a copy of this notification in your storeroom and clinical areas. 4. NOTIFY any customers to whom you may have distributed/forwarded affected product (or to whom you intend to distribute/forward product) about this Field Safety Corrective Action and share a copy of this notice. 5. RETURN the enclosed acknowledgment form by following the directions on the Form, whether you have affected product or not. | 2025-10-07 06:30
Cardinal Health 200, LLC | Salem Sump™ | - | - | Cardinal Health received complaints regarding the Anti Reflux Valve ARV breaking due to use x D error The ARV failures are attributed to excessive force during use Based on this Cardinal Health is providing additional guidance when using these products | 1. REVIEW your inventory for the affected product codes. All lots of inventory are impacted. 2. COMMUNICATE the risks and appropriate usage of the ARV with all personnel. 3. POST a copy of this notification in your storeroom and clinical areas. 4. NOTIFY any customers to whom you may have distributed/forwarded affected product (or to whom you intend to distribute/forward product) about this Field Safety Corrective Action and share a copy of this notice. 5. RETURN the enclosed acknowledgment form by following the directions on the Form, whether you have affected product or not. | 2025-10-07 06:46
Abbott Ireland Diagnostics Division | FreeGo | - | - | Through our post market surveillance process it has come to our attention that some customers are using the FreeGo Enteral Feeding Sets in a manner inconsistent with the labeling which has the potential to pose a health risk | Reinforcement of Instructions for Use. | 2025-10-08 12:18
Beckman Coulter Ireland, Inc | BICARBONATE | 279 | - | Beckman Coulter has received complaints for elevated results generated using Bicarbonate reagent when there are high levels of Lactate Dehydrogenase LDH in the test sample x D Beckman Coulter has investigated and confirmed that the lots listed above of Bicarbonate reagent may generate Bicarbonate results that are biased high due to interference with LDH in the test sample x D The level of LDH interference is not consistent across lots of Bicarbonate reagents The magnitude of the impact varies depending on Bicarbonate lot in use and the level of LDH in the test sample | Discontinue use and dispose all Bicarbonate OSR6x37/OSR6x90 lots listed in Table A in the ISSUE section of this letter, following your laboratory protocols and local regulations. Contact your local Beckman Coulter representative for replacement or re-imbursement. | 2025-10-10 22:50
BD Switzerland Sàrl | BD Alaris™ Syringe Pumps | 8 | - | Based on customer feedback BD has identified that a potential manufacturing defect affecting certain BD Alaris Syringe Pumps related to keypad functionality It has been observed that specific lots of the keypad components may exhibit unintended behaviour impacting the On Off Hold and Run keys x D x D Specifically the affected pumps may experience one or more of the following failure modes x D x D The pump may enter Tech Mode autonomously during power up x D Keypad buttons may become unresponsive preventing normal operation x D A KY error may be displayed indicating a stuck key see Picture below x D x D x D x D x D x D x D x D x D x D x D x D Picture KY Stuck Key Error Displayed x D x D The specific lots of the keypad components were potentially installed by BD either during the manufacturing or the servicing of the specific BD Alaris Syringe Pumps listed in Table webpage link or shipped to service organisations within the spare parts kits between February and th August as listed in Table x D x D Note The webpage link referenced below lists the pumps manufactured with the affected keypad components or pumps serviced directly by BD only If your Service Organisation is not BD they will contact you if any additional pumps were fitted with the impacted parts and therefore may not appear in the webpage link x D x D | Actions for Clinical Users • Organize for the affected Alaris™ Syringe pumps serial numbers to be made available to BD, your Biomedical Engineering team or to your Service Organisation as applicable for keypad replacements. o Clinical users can identify impacted pumps by reviewing the list of Impacted Pumps, found on the webpage: https://bdx.my.site.com/CC360/s/impactedproducts?rn=MMS-25-5342%20GLOBAL o If your Service Organisation is not BD, they will contact you if any additional pumps were fitted with the impacted parts. • Continue to use the Alaris™ Infusion pumps with their existing labeling and Directions for Use. In case you observe the pump presenting with the described issue, immediately remove it from service for inspection by Qualified Service Personnel. Use an alternative syringe pump in the meantime. o Reminder, as stated in the DFU for PK pumps (SKUs: PKnexus1, 8005PK201 or 8005TIG03) “if a surgical procedure is interrupted during a TCI (Target Controlled Infusion), all current pharmacokinetic / pharmacodynamic model information will be lost. Under such circumstances, switching the pump off and on and restarting the infusion whilst the patient contains a significant residual drug dose, could result in an over-infusion and therefore the pump should not be restarted in TCI mode”. • If you experience any issues with the Alaris™ Infusion pumps, please report as a complaint as per your normal process. • Per good clinical practice, o Ensure spare infusion pumps are available in clinical areas delivering critical or life sustaining medications. o If performing a patient transfer, assess if additional equipment, or alternative methods of medication administration, is required in case of system alarms. Actions for Biomedical Engineers / Service Organisations BD is issuing a Technical Service Bulletin (BDIN00346) to all trained Biomedical Engineers, service organisations to provide instructions on: • Actions to take with any impacted spare part kits • How to remediate impacted pumps BD Actions: BD has investigated and identified the root cause of the issue. Corrective actions have been implemented. To date, BD does not plan to initiate any further advice or information in a follow-up FSN. Customer Actions: • Review the information in Table 1/ webpage link to determine if the BD Alaris™ Syringe Pumps in your possession are impacted. • BD distributed spare part kits between February 2024 and 29th August 2025 (see Table 2) and if you serviced pumps with these parts, you may have additional impacted pumps. Review your service records and determine if pumps were serviced with impacted spare parts. Technical Service Bulletin (BDIN00346) is being issued to all Service Organisations to provide instructions on how to identify impacted spare parts and what actions to take. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 13th November 2025. • This notice needs to be passed on all those who need to be aware within your organisation or to any organisation where the potentially affected devices have been transferred. • Transfer this notice to other organisations on which this action has an impact. • Maintain awareness of this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. • Report all device-related incidents to the manufacturer, distributor or local representative, and the national Competent Authority if appropriate, as this provides important feedback. Distributor Actions: • Review the information in Table 1/ webpage link and determine if the BD Alaris™ Syringe Pumps in your possession are impacted. • BD distributed spare part kits between February 2024 and 29th August 2025 (see Table 2) and if you serviced pumps with these parts, you may have additional impacted pumps. Review your service records and determine if pumps were serviced with impacted spare parts. Technical Service Bulletin (BDIN00346) is being issued to all Service Organisations to provide instructions on how to identify impacted spare parts and what actions to take. Notify these customers of this notice and replace with new, unaffected spare parts. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 13th November 2025. o There is no requirement to return your customer response forms to BD, you should maintain these on file at your facility. Return only your final consolidated response form. • Complete and return the Customer Response Form following completion of your reconciliation activities. • Maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. • Report all device-related incidents to the manufacturer, or local representative, and the national Competent Authority if appropriate, as this provides important feedback. | 2025-10-12 12:53
Abiomed Inc. | Automated Impella Controller (AIC) | 52 | - | Abiomed identified cybersecurity vulnerabilities in the Automated Impella Controller AIC operating system that could be exploited through network or physical access No incidents or patient harm have been reported Abiomed is developing security updates including a USB driver isolation patch and physical access controls with the software patch expected by December The AIC can continue to be used as intended x D | -Keep the AIC in a secure environment with restricted access whether in clinical use or not. -Abiomed is working on security updates and measures to address these cybersecurity vulnerabilities. To mitigate the vulnerability risk, Abiomed field representatives will contact customers to implement a USB driver isolation software patch and physical controls to eliminate vulnerability or, if requested by the customer, to arrange disabling the AIC’s network capabilities. The AIC may continue to be used as intended. Timing for the availability of the patch is currently estimated as December, 2025." | 2025-10-12 17:22
GE Healthcare | All 1202 Flex Focus variant using UA1214 - Mobile keyboard dock with battery or UA1814 Mobile touch dock with battery. All bkSpecto , bkActiv, bk3000, bk3500 & bk5000 with battery | 43 | - | GE HealthCare has become aware that the Instructions for Use within the user guide and service manual do not adequately state the required battery replacement intervals end of life handling procedures or general battery safety precautions This lack of guidance could result in delayed battery replacement improper handling or continued use of degraded batteries potentially increasing the risk of a system malfunction or in rare cases smoke or fire x D There have been no injuries reported as a result of this issue x D | Customer Letter with link to updated addendum to all affected customers | 2025-10-14 12:08
SHENZHEN MINDRAY BIO-MEDICAL ELECTRONICS CO., LTD | V90 electronic vaporizer (for used with A9 Anesthesia system) | 29 | - | Mindray has identified that the V electronic vaporizer used with the A anesthesia x D system may have leakage issue x D During the process of filling the anesthetic agent into the vaporizer frequently applying x D excessive downward pressure on the anesthetic agent bottle may damage the internal x D structure of the filling port This damage may lead to improper sealing of the filling port after x D the removal of the anesthetic agent bottle and may result in anesthetic leakage x D While leakage occurs medical staff will smell the odor of anesthetic and can address the x D leakage issue | Hardware replacement | 2025-10-14 14:16
CellaVision AB | AUTOMATED DIGITAL CELL MORPHOLOGY ANALYZER (DI-60) | 32 | - | A malfunction has been identified with the DI analyzer relating to slide identification x D DI has a shuttle which transports slides from the Sysmex CF into the DI for analysis and back to the Sysmex CF after analysis The shuttle has a fetch position for slides coming into the DI for analysis and a return position for slides going out from the DI after analysis At the occasion when the barcode reader reads the incoming slide the incoming and the outgoing slides are located next to each other in the shuttle If the analyzer fails to read the barcode of the incoming slide there is a risk that it instead reads the barcode of the outgoing slide The incoming slide is thus assigned the same identifier as the outgoing slide x D This malfunction may lead to misattribution of patient results Potential clinical impact includes incorrect or delayed diagnosis initiation of inappropriate treatment or delay in necessary treatment any of which could result in serious health consequences depending on the patient condition and test type x D Although no adverse health outcomes have been confirmed to date the malfunction represents a potential for serious indirect harm | A malfunction has been identified with the DI-60 analyzer relating to slide identification. DI-60 has a shuttle which transports slides from the Sysmex CF into the DI-60 for analysis, and back to the Sysmex CF after analysis. The shuttle has a fetch position (for slides coming into the DI-60 for analysis) and a return position (for slides going out from the DI-60 after analysis). At the occasion when the barcode reader reads the incoming slide, the incoming and the outgoing slides are located next to each other in the shuttle. If the analyzer fails to read the barcode of the incoming slide there is a risk that it instead reads the barcode of the outgoing slide. The incoming slide is thus assigned the same identifier as the outgoing slide. This malfunction may lead to misattribution of patient results. Potential clinical impact includes incorrect or delayed diagnosis, initiation of inappropriate treatment, or delay in necessary treatment, any of which could result in serious health consequences depending on the patient condition and test type. Although no adverse health outcomes have been confirmed to date, the malfunction represents a potential for serious indirect harm. For this reason, a Field Safety Corrective Action (FSCA) is being initiated to mitigate the risk and ensure continued patient safety. Internal investigation has confirmed that the malfunction is caused by conditions in the CCU firmware in combination with newer barcode readers (Honeywell Vuquest and Jadak FM-5). Specifically, the barcode reader has a field of view that is too large, allowing capture of the barcode on the return slide (located next to the fetch slide). The barcode reader continues to attempt barcode reading after the gripper has released the return slide, thereby exposing its barcode. To eliminate the root causes, corrective actions have been implemented in the CCU firmware (version 4.2.7), which includes: - Reducing the effective field of view of the barcode reader. - Preventing additional barcode reading attempts once the gripper has released the return slide. Firmware version 4.2.7 has already been introduced in manufacturing. As part of this FSCA, all affected DI-60 instruments in the field will be updated with the corrected firmware via Distributor's service engineers. Distributors are required to confirm completion of the update. These actions are justified to prevent recurrence of the malfunction, to mitigate the risk of slide misidentification, and to ensure patient safety. 1. Action To Be Taken by the User ☒ Identify Device ☒ On-site device modification / inspection The following actions shall be taken by the user of the device: 1. Identify whether any of your DI-60 analyzer(s) are affected by verifying the serial number(s) against the list of affected units provided in Table 1. 2. If your unit(s) are affected, please implement the following interim measure to mitigate the risk of sample misattribution: a. Before signing each order, confirm that the number of slides in the order matches the number of slides ordered. No order shall contain more slides than prescribed. b. Await contact from Sysmex service personnel, who will arrange installation of the updated CCU firmware. 3. You may continue to use the system(s) in accordance with its intended use, provided the above instructions are followed 4. Please ensure that this Field Safety Notice is circulated to all users of the device so that they are made aware of the potential issue. 5. Retain this letter until the firmware correction has been implemented. Ensure that the notice is kept in a visible and accessible location. 6. Fill in the customer response form and send via e-mail to whomever you received the FSN from | 2025-10-14 16:08
Medical Components, Inc. (dba Medcomp) | Hemo-Cath Silicone Double Lumen Catheter Set | 1 | - | x D The root cause cannot be determined at this stage as the affected device and supporting photos have not been received from the hospital and there is no direct means of communication with the facility where the incident occurred The manufacturer will not be able to perform a technical investigation until the device or additional information becomes available | Please note that we currently have no direct means of communication with the hospital where the incident occurred. The affected device and any supporting photos have not been received. Therefore, the manufacturer is unable to begin a technical investigation or determine the root cause at this stage. The manufacturer (Medcomp) has opened a complaint file and requested the return of the affected device for evaluation. Once the device or additional information is obtained from the hospital, the investigation will proceed, and a final report will be submitted to SFDA as soon as possible. | 2025-10-15 14:30
Abbott | CentriMag Blood Pump | 21 | - | Failure to properly lock the Blood Pump into the x D Motor could lead to particulate or increased heat while using the system | Ensure the CentriMag Blood Pump is securely locked into the motor, following the IFU. If you are unable to fully rotate and lock a blood pump into a motor, exchange the pump for a new one, and return the original pump to Abbott for a full refund or credit | 2025-10-15 15:17
Abbott | HeartMate II and HeartMate 3 System Controllers | 412 | - | when performing x D Backup Battery installation or replacement clinicians should minimize excessive handling movement x D stress and pulling at the interface between the ribbon cable and controller | remind patients and caregivers to follow the Patient Handbook: If System Controller Backup Battery Fault Alarm is active, patients should ﬁrst ensure that no other alarms are active, and then they should call their hospital contact as soon as possible for diagnosis and instructions. Patients and caregivers should not immediately attempt to exchange the System Controller without clinical guidance. | 2025-10-15 15:36
Olympus | SnareMaster Plus | 885 | - | During an internal review of the SD Snares sterile packaging process Olympus discovered holes in a small number of device pouches Olympus s investigation determined that small holes may occur during packaging of the device pouches into the shelf carton and or during transportation of the shelf carton Olympus has not received any complaints related to this issue | 1. Examine your inventory and quarantine any SD-400 Snares with the lot numbers listed above. 2. Cease usage of the product immediately. | 2025-10-16 19:19
BK Medical Aps | bk medical | - | - | The Instructions for Use within the user guide and service manual do not adequately state the required battery replacement intervals end of life handling procedures or general battery safety precautions x D This lack of adequate guidance in the IFU and Service Manual led to potential use of aged or degraded batteries without proper replacement intervals | GE HealthCare is providing updated Instructions for Use addendums and updated Service Manuals with specific instructions regarding battery safety.” “You can continue to use your device by following the instructions below: Review and follow the applicable addendums to the Instructions for Use and the updated Service Manuals. The Addendums and Service Manuals are available for download at: https://www.gehealthcare.com/support/manuals Follow the conditions defined within the addendum and Service Manual applicable to your product for further instructions. Place the applicable addendum and Service Manual with your existing product Instructions for Use.” | 2025-10-19 10:33
DiaMed GmbH | IH-500 | 141 | - | The purpose of this letter is to advise you of important information regarding an observed increase in the rate of Air Gap Failure messages flagged by the IH instrument when using ID Cards x D Reason for the Action x D Bio Rad has become aware of an increase in complaints related to Air Gap Failure messages occurring on the IH system when using ID Cards x D Impact on IH Workflow x D When an Air Gap failure is detected results are automatically flagged in IH Com and the system disables automatic result validation Users must therefore manually review and validate the affected results regardless of their standard IH Com verification settings x D x D Kindly review the attached | Action(s) to be taken by the Customer: Bio-Rad is recommending that customers affected by this notice take the following action: • Continue to review results carefully when an Air Gap Failure flag is displayed. • For results flagged as air gap failure, manual validation is required. Please follow established rules for interpretation and repeat testing if necessary. • Please follow the instructions in the flag management section of the IH-Com User Manual. Please ensure that this notice is shared with all relevant personnel within your organization and any third-party recipients of the affected products. Please complete and return the attached response form as soon as possible so that we are assured you have received this important communication. Resolution by Bio-Rad: Bio-Rad takes product quality and safety very seriously. Corrective actions are currently in progress to prevent the recurrence of this issue. Kindly review the attached | 2025-10-19 14:48
Olympus | Evis Exera III Duodenovideoscope, Evis Lucera Elite Duodenovideoscope, Duodenovideoscope | 15 | - | Olympus s ongoing assessment of TJF duodenoscope reprocessing to address positive cultures and infections has identified updates to the reprocessing materials detailed in the Summary of Changes section below to minimize potential deviations in TJF duodenoscope reprocessing Olympus is informing users of these changes and is reminding users to closely follow the operation and reprocessing instructions Customers are required to acknowledge reviewing all such materials which can be accessed through the link provided in the Actions Required section below | 1. Carefully read the content of this notification. 2. Access the Olympus education portal and complete the required training course. "Complete the TJF Training and Labeling course, which requires you to download and review the following materials: 1) Introduction to Duodenoscope Reprocessing updated training module 2) Updated Operation Manual (Operation IFU) 3) Reprocessing Manual (Reprocessing IFU) 4) Visual Inspection Guide 5) OnTrack Reprocessing In-Service/Customer Competency form" 3-Olympus requests that you acknowledge receipt of this letter by filling out and returning the completed enclosed Reply Form back to your local Olympus representative 4-Keep a copy of this notification with your current Operation Manual. Olympus requests that you replace any existing copies of the Operation Manual with the updated Operation Manual and include a copy of the Visual Inspection Guide, obtained through the link above in Step 2 5-Olympus recommends that you receive on-site TJF duodenoscope reprocessing training at your facility. | 2025-10-19 23:27
SCHILLER AG | Easy pulse | 10 | - | Please find the attached FSN letter | Please find the attached FSN letter | 2025-10-20 11:44
Covidien llc | Signia | 66 | - | Medtronic has identified an internal manufacturing issue with the Small Diameter Stapling Reload s Cover Tube Press step which can produce parts with a malformed press This press in the external cover tube creates an interference hold of internal components within the stapling reload s shaft which enables adequate transfer of force through the reload to fire the stapler and maintain jaw orientation during firing A malformed cover tube press can lead to internal components movement during firing leading to the potential for uncontrolled motion x D The issue can occur using either the Endo GIA Ultra Universal Stapler or Signia Stapling System x D While using Endo GIA Ultra Universal Stapler or Signia Stapling System uncontrolled articulation in either direction can occur during firing as well as after firing if the user continues to squeeze the handle or press the fire button after the firing of the reload has been completed x D Uncontrolled articulation is also possible if the reload is being used outside of the indicated tissue thickness range x D Update October x D Medtronic has identified an internal manufacturing issue with the Small Diameter Stapling Reload s Cover Tube Press step which can produce parts with a malformed press The tube press on the reloads were measured and found to be out of specification Engineering functionally tested one device and found that the reload de articulated during firing Production and releases were suspended x D The malformed cover tube press occurred from an incorrect parameter setting for long cover tube assembly in the Programmable Logic Controller PLC program for mm CFNs Parameters were reviewed for the mm PLC program and found to be correct The parameter for mm was validated to be mm for long cover tubes however a line of code in the logic allowed for this parameter to be reverted to the previously validated setting of mm Although a similar line of code exists for short cover tubes there are additional aspects to this line of code that prevented it from being activated The scope of this issue has been bounded to all long cover tube lots built since the implementation of the new parameters under CDP totaling four lots N F Y N F Y N D UY N B UY x D | • Identify and quarantine all unused impacted product listed above. • See Attachment A for guidance on identifying affected lots. • Return all quarantined devices to Medtronic. Please contact your Medtronic Representative in case you need assistance with the return of the affected devices. o Credit for the returned affected devices will be issued based on the RGA number. o If purchased from a distributor, contact your distributor directly to arrange for the return of the devices back to your distributor. • Pass on this notice to all those who need to be aware within your organization or to any organization where the listed affected lots may have been transferred or distributed. | 2025-10-20 14:12
Spacelabs Healthcare Inc | DVI display cables | 1752 | - | The issue arises from inadequate electromagnetic shielding in the DVI cable design During routine IEC electromagnetic compatibility EMC testing it was found that the cable s shielding does not sufficiently block external electromagnetic interference This design or manufacturing deficiency allows unwanted electrical signals to escape or enter through the cable leading to potential interference with the operation of nearby sensitive medical devices x D | 1-Acknowledge receipt and confirm understanding of the Urgent Medical Device Correction Notification. -Request replacement cables for the affected DVI cables no later than 31 December 2025. 2-The acknowledgment can be completed by either: Submitting the Acknowledgment Form (Appendix A) by mail, or Completing the online acknowledgment form provided by Spacelabs. 3-These steps ensure that affected cables are identified and replaced to prevent potential electromagnetic interference risks. | 2025-10-21 18:29
Olympus | ShockPulse SE Transducer, ShockPulse SE Lithotriptor | 13 | - | Olympus received customer feedback that the ShockPulse Lithotripsy Transducer may either fail to start up or the transducer may start briefly and then stop accompanied by an error light on the generator In addition the body of the transducer handpiece may gradually increase in temperature during clinical use The investigation has shown that the transducer may fail in the field before reaching its expected reprocessing cycles x D x D | 1-Read this Field Safety Notice carefully. 2-Ensure all staff are informed. This is not a product recall; the device can still be used as instructed, with a backup transducer and probe sterilized and ready. 3-Confirm receipt and understanding by completing and returning the enclosed Reply Form to Gulf Medical representative. | 2025-10-21 23:07
Abbott Molecular Inc. | Alinity m HIV-1 AMP Kit | 3 | - | Please find attached FSN letter | Please find attached FSN letter | 2025-10-22 15:44
Philips Healthcare (Suzhou) Co., Ltd. | Multiva 1.5T | 1 | - | The Quadrature Body Coil QBC seal adhesive may fail creating sharp edges that may come in contact with patients The QBC seal may become loose as the patient table travels in a horizontal motion in and out of the system bore The QBC seal Figure is a rubber seal that is glued between the cone cover and QBC cover and functions to prevent sharp edges of the QBC cover from contacting patients during an examination | Philips will contact you to schedule a time for a Field Service Engineer (FSE) to visit your site to inspect your system’s QBC Seal and replace the QBC Seal if necessary (reference FCO78100616 | 2025-10-23 16:38
Cardinal Health 200, LLC | Kendall SCD™ 700 Sequential Compression System | 4985 | - | Non OEM power cords may lack proper electrical compatibility safety certifications and mechanical fit increasing the risk of device failure and safety standards Their use often unnoticed by clinical staff is further compromised by environmental stress in hospital settings | 1. REVIEW your inventory for the above-mentioned product codes to ensure that an I-Sheng OEM power cord is used. 2. COMMUNICATE the risks of using non-OEM power cords. 3. POST a copy of this notification in your storeroom and clinical areas. 4. NOTIFY any customers to whom you may have distributed/forwarded affected product (or to whom you intend to distribute/forward product) about this medical device product advisory and share a copy of this notice. 5. RETURN the enclosed acknowledgment form by following the directions on the form. | 2025-10-24 06:47
Philips Healthcare (Suzhou) Co., Ltd. | Incisive CT | 17 | - | Philips has identified an issue that may occur with the Philips Incisive CT CT and CT systems if the screws of tube heat exchanger on the rotating scanner are not tightened well after replacing This component may become detached and make contact with other components located within the Philips CT systems during rotation Other components could be damaged due to the contact with the tube heat exchanger x D x D If this issue occurs a loud noise will be emitted and the system will shut down x D The top right gantry cover may become broken if the following sequence occurs x D Tube heat exchanger screws come loose x D The tube heat exchanger is detached during rotating and contacts other components in the system and cause them to break x D A broken component makes contact with the top right gantry cover x D x D The broken top right gantry cover could create a gap allowing a fragment of a damaged component to be expelled at a lower kinetic energy and land on the right side of the gantry | Philips will contact you to schedule a time for a Field Service Engineer (FSE) to visit your site to inspect the tube heat exchanger screws and correct the loosen screw if necessary (reference FCO72800842). | 2025-10-24 13:53
Intersurgical Ltd. | Various Clear-Therm™ Mini HMEF | - | - | The two housings of some HMEF devices have been found to separate when a force is applied during handling of these products or movement and torsion due to repositioning of the patient x D The reported separation of the HMEF would cause gross leakage from the device which in turn would result in leakage of gas from the breathing system x D | Identify Device Quarantine Device the following actions are taken in addition to those detailed in the IFU provided with the device. 1. Identify any potentially affected products from the affected codes and lot numbers listed above. 2. Immediately before use unpack the device and carry out the check as described below. 3. All users must check the HMEF housing is securely welded by holding both connection tapers at each end and apply a downward breaking force by flexing the tapers in the direction shown by the arrows below. 4. Retain any affected sample(s) identified, and please report to us immediately | 2025-10-24 22:01
CareFusion 303, Inc. | Alaris™ Pump | 11067 | - | BD was made aware via customer complaints that dropping the Alaris Pump Module may cause damage to internal components that may not be apparent or readily visible More specifically drops or severe jarring may damage the Alaris Pump module bezel assembly which provides the mechanical foundation for critical pumping components This damage can cause under infusion over infusion unregulated flow or failure of the Alaris Pump module to calibrate x D x D Please note that this issue is separate from worldwide field action MMS FA of Alaris Pump Modules LVP for units built and service kits manufactured with FR resin FR between April and June x D | Clinical User Actions 1. Customers should NOT USE a device that has been dropped or severely jarred. Devices that have been dropped or severely jarred should be segregated and sent to biomedical engineering for inspection, as directed in the BD Alaris™ System with Guardrails™ Suite MX User Manual: “If a device or accessory is dropped or severely jarred, take the device out of use immediately. Send the device to biomedical engineering for inspection to ensure the device is functioning properly before reuse.” 2. Prior to use, inspect the Alaris™ Pump Module for visual damage. If the device has physical damage, remove from service and send to biomedical engineering for further inspection and testing, as directed in the BD Alaris™ Infusion System User Manual: “To ensure that the BD Alaris™ Infusion System remains in good operating condition, visually inspect the system before each use. Check all visible surfaces and moving parts on the devices. If you observe damage in any way or find the device does not function as expected, return it to biomedical engineering for repair.” Biomedical User Actions 1. Perform the testing as specified in the Level of Testing Guidelines in your Technical Service Manual for a dropped device, including rate accuracy verification testing. If any test fails, perform the applicable steps in the troubleshooting section of your Technical Service Manual, which may include replacement of the pumping mechanism. See Technical Service Manual Section. 2. Continue to perform annual preventive maintenance, including rate accuracy verification (as discussed in the previous action item), on all of your Alaris™ Pump modules, as specified in the Alaris System Maintenance Guide. Preventive maintenance is important in ensuring regular testing of any potentially dropped or damaged device that was not immediately removed from use. BD Actions: 1. BD is providing a Tip Sheet (Appendix A) to be distributed to all clinical users, environmental services staff, transport staff, biomedical engineering and any other key personnel who may handle Alaris™ Pump Modules. The Tip Sheet is intended to remind all personnel who handle the Alaris™ Pump Modules that dropped/damaged units should be removed from service and inspected. 2. BD will be adding reinforcements to applicable labeling to reiterate that if devices are dropped or severely jarred, they should be removed from service and inspected. These warnings currently exist in your Alaris™ System with Guardrails™ Suite MX User Manual. 3. BD continues to investigate long-term product enhancements related to this issue. Customer Actions: • Review the information in Table 1 and conduct the actions as per this notification. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 30th November 2025. • This notice needs to be passed on to all those who need to be aware within your organisation or to any organisation where the potentially affected devices have been transferred. This includes clinical and non-clinical users (e.g., cleaning and transport staff) are trained to remove any dropped or severely jarred devices from use and sequester the device for inspection and testing. • Please transfer this notice to other organisations on which this action has an impact. • Please maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. • Please report all device-related incidents to the manufacturer, distributor or local representative, and the national Competent Authority if appropriate, as this provides important feedback. Distributor Actions: • Review the information in Table 1 and conduct the actions as per this notification. • This notice needs to be passed on to all those who need to be aware within your organisation or to any organisation where the potentially affected devices have been transferred. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. • Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 30th November 2025 • There is no requirement to return your customer response forms to BD, you should maintain these on file at your facility. Return only your final consolidated response form. • Complete and return the Customer Response Form following completion of your reconciliation activities. • Please maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. • Please report all device-related incidents to the manufacturer, or local representative, and the national Competent Authority if appropriate, as this provides important feedback. | 2025-10-26 15:59
CareFusion 303, Inc. | Alaris™ Pump | 11067 | - | BD was made aware via customer complaints that dropping the Alaris Pump Module may cause damage to internal components that may not be apparent or readily visible More specifically drops or severe jarring may damage the Alaris Pump module bezel assembly which provides the mechanical foundation for critical pumping components This damage can cause under infusion over infusion unregulated flow or failure of the Alaris Pump module to calibrate x D x D Please note that this issue is separate from worldwide field action MMS FA of Alaris Pump Modules LVP for units built and service kits manufactured with FR resin FR between April and June x D | Clinical User Actions 1. Customers should NOT USE a device that has been dropped or severely jarred. Devices that have been dropped or severely jarred should be segregated and sent to biomedical engineering for inspection, as directed in the BD Alaris™ System with Guardrails™ Suite MX User Manual: “If a device or accessory is dropped or severely jarred, take the device out of use immediately. Send the device to biomedical engineering for inspection to ensure the device is functioning properly before reuse.” 2. Prior to use, inspect the Alaris™ Pump Module for visual damage. If the device has physical damage, remove from service and send to biomedical engineering for further inspection and testing, as directed in the BD Alaris™ Infusion System User Manual: “To ensure that the BD Alaris™ Infusion System remains in good operating condition, visually inspect the system before each use. Check all visible surfaces and moving parts on the devices. If you observe damage in any way or find the device does not function as expected, return it to biomedical engineering for repair.” Biomedical User Actions 1. Perform the testing as specified in the Level of Testing Guidelines in your Technical Service Manual for a dropped device, including rate accuracy verification testing. If any test fails, perform the applicable steps in the troubleshooting section of your Technical Service Manual, which may include replacement of the pumping mechanism. See Technical Service Manual Section. 2. Continue to perform annual preventive maintenance, including rate accuracy verification (as discussed in the previous action item), on all of your Alaris™ Pump modules, as specified in the Alaris System Maintenance Guide. Preventive maintenance is important in ensuring regular testing of any potentially dropped or damaged device that was not immediately removed from use. BD Actions: 1. BD is providing a Tip Sheet (Appendix A) to be distributed to all clinical users, environmental services staff, transport staff, biomedical engineering and any other key personnel who may handle Alaris™ Pump Modules. The Tip Sheet is intended to remind all personnel who handle the Alaris™ Pump Modules that dropped/damaged units should be removed from service and inspected. 2. BD will be adding reinforcements to applicable labeling to reiterate that if devices are dropped or severely jarred, they should be removed from service and inspected. These warnings currently exist in your Alaris™ System with Guardrails™ Suite MX User Manual. 3. BD continues to investigate long-term product enhancements related to this issue. Customer Actions: • Review the information in Table 1 and conduct the actions as per this notification. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 30th November 2025. • This notice needs to be passed on to all those who need to be aware within your organisation or to any organisation where the potentially affected devices have been transferred. This includes clinical and non-clinical users (e.g., cleaning and transport staff) are trained to remove any dropped or severely jarred devices from use and sequester the device for inspection and testing. • Please transfer this notice to other organisations on which this action has an impact. • Please maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. • Please report all device-related incidents to the manufacturer, distributor or local representative, and the national Competent Authority if appropriate, as this provides important feedback. Distributor Actions: • Review the information in Table 1 and conduct the actions as per this notification. • This notice needs to be passed on to all those who need to be aware within your organisation or to any organisation where the potentially affected devices have been transferred. • Identify the facilities where you have distributed affected product and notify them immediately of this notice. • Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 30th November 2025 • There is no requirement to return your customer response forms to BD, you should maintain these on file at your facility. Return only your final consolidated response form. • Complete and return the Customer Response Form following completion of your reconciliation activities. • Please maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. • Please report all device-related incidents to the manufacturer, or local representative, and the national Competent Authority if appropriate, as this provides important feedback. | 2025-10-26 15:59
Draeger Medical, Inc. | ErgoStar | 5250 | - | The smoothbore hoses of the Catheter Mounts may be too thin and thus may tear under stress what could lead to leakages | 1.Dispatch of an Important Safety Notice to all customers by 15.11.2025. 2. Delivery of replacement if requested by the customer by 30.02.2026. | 2025-10-27 14:38
Werfen - Biokit, S.A. | BIO-FLASH | 8 | - | To assess the regulatory impact of the release of software version of BIO FLASH P N | BIO-FLASH software version 3.3.2 will include the following changes: 1. Operating System a. Compatibility with Windows 10 IoT LTSC 2021, supported until 2031. b. The installation process now includes a validation step to ensure the application is only installed on the supported operating system version. 2. Cybersecurity Enhancements a. Increased password complexity requirements. b. User accounts are temporarily locked for 5 minutes after multiple failed login attempts. 3. Logging Improvements a. Added a button to print the Instrument Log. b. Enhanced logging performance: logs are now written in parallel during active runs, improving overall system responsiveness. 4. Specimen Lifetime Configuration a. The Specimen Lifetime setting (in hours) defines how long a specimen remains in the database after its first entry. b. Acceptable values range from 0 to 999 hours (default: 6 hours). A value of 0 means the specimen is never removed. 5. Cartridge Compatibility: added support for 200-test cartridges (availability to be confirmed). 6. Results Screen Optimization a. Added "Expand" option in the submenu for Results and Results Archive. b. Replaced results archiving with a pagination feature. c. Improved automatic results printing functionality. d. Enhanced sorting options in the results list report (e.g., by barcode, assay). e. Improved results filtering capabilities. 7. Performance Optimization: general application performance improvements, enabling safe operation with large panels and high result volumes. 8. User Interface Updates: updated the System Cleaning Solution wizard completion message. 9. LIS Demographics Handling: a. Optional display and storage of demographic data received from the LIS. b. When enabled, a corresponding entry is recorded in the audit log indicating the user’s selection. 10. Audit Log Enhancements: additional logging of user actions, including deletion of sensitive data such as results. 11. Special Fluids (QUANTA Flash Special Wash) Stability Management: improved handling of onboard stability for special fluids. 12. QC Module Enhancements: a. Support for simultaneous QC of multiple reagent cartridges. b. Improved Levey-Jennings chart display. c. Enhanced QC screen usability and performance, including filtering by assay and graphical display of out-of-range data points. 13. Reagent Carousel Thermal Regulation: when the instrument is idle, the reagent carousel now rotates periodically to maintain optimal temperature. 14. Rounding | 2025-10-28 09:11
SEBIA | CAPI 3 IMMUNOTYPING | 16 | - | These abnormal subtractions lead to difficulty in interpreting the light chain which can appear x D x D Either completely subtracted in Kappa and Lambda simultaneously in the case of IgM Kappa x D Or partially subtracted in Lambda x D x D x D This makes interpretation impossible and in some cases may lead to errors x D | Our traceability indicates that you received one of this lot number, therefore, we kindly ask you to: 1. Destruct all remaining kits of CAP! 3 IT of these incriminated lots in your possession. 2. Check the results obtained with these kits (start date of delivery of these lots: 04/09/2025). 3. Fill-in the enclosed forms (Information certificate and destruction form), precisely indicate the number of kits received and the number of kits destroyed. Send back all the forms by e-mail to your local SEBIA contact. | 2025-10-28 10:42
Medtronic MiniMed | CareLink Clinic | 52 | - | In September Medtronic noticed that the CareLink Clinic also known as ProWeb Dashboard was displaying incorrect data Through internal testing anomalies were found x D The SG graph incorrectly displays hypoglycemic patterns from both day and day ranges Cause Faulty JavaScript logic handling date range toggling x D After navigating away and returning to the Dashboard tab the date range resets to days When re selecting days the graph still shows days of data Cause Reference chart fails to update due to tab navigation and date range toggle combination x D | Your Recommended Actions: • If you made individual diabetes treatment recommendations to patients based on the hypoglycemic patterns or on the 30-day sensor glucose data displayed on the 24-hour Sensor Glucose Overview graph on the Patient Dashboard between September 11th, 2025 and October 4th, 2025 (e.g. if you increased a user’s active insulin time and/or decreased a user’s insulin-to-carbohydrate ratio based on the incorrectly displayed data), please revisit those recommendations utilizing patients’ individual CareLink reports, as they were not impacted by this issue. | 2025-10-28 15:03
Medtronic Inc. | External Drainage and Monitoring Kit | 44 | - | Medtronic has received customer complaints reporting cracks and leaks on the stopcock of Becker and Exacta EDMS devices Complaint analysis shows a higher complaint rate associated with the p n Four Way Stopcock used in Becker and Exacta EDMS devices | Please complete and return the enclosed Customer Acknowledgment Form to acknowledge that you have read and understand this letter | 2025-10-28 18:44
Medtronic Inc. | Aurora EV-ICD MRI SureScan | 69 | - | An event was reported by a Medtronic field representative that during routine defibrillation threshold testing at x D an EV ICD implant Ventricular fibrillation VF detection was appropriately met and the capacitors were charged x D to deliver a shock however following charge end the device was no longer showing sense markers and therapy x D was aborted The Medtronic representative stated that it was difficult to tell if the device lost sensing telemetry or other interaction Sensing resumed after seconds and redetection occurred but an external defibrillation shock was manually delivered to convert the rhythm prior to fulfilling device confirmation and therapy delivery A second induction attempt through the device resulted in a shock being delivered after charge end that successfully terminated the arrhythmia Medtronic internal review of the device data determined that the absence of sensing markers occurred due to a firmware timing interaction between the charge end and a ventricular sense VS event coincident with a drop in device measured R wave amplitudes x D x D When an EV ICD detects a ventricular arrhythmia and determines therapy is required the device charges the capacitors and delivers the shock If the end of the capacitor charge interrupts an in progress R wave amplitude measurement there is a potential for the dynamic sensing algorithm to remain at its first plateau of of the prior R wave measurement amplitude This may be impactful and lead to a delay in high voltage therapy if the amplitudes decrease at the time of the interrupt x D x D The root cause of the delayed therapy was determined to be VF undersensing due to a firmware timing x D interaction with the charge end event and a VS event coincident with a significant decrease in the amplitude of x D the VF signals that was not large enough to be sensed at the initial automatic sensing threshold The sensitivity x D remained at the initial automatic sensing threshold of of the prior sensed event amplitude without automatic x D adjustment due to charge end CE interrupting an in progress amplitude measurement When the amplitude of x D the subsequent signal falls below this sensing threshold as seen in the reported complaints delayed x D tachyarrhythmia therapy may result This happens because a non committed shock cannot reconfirm and sync to x D a ventricular sensed beat Non committed shocks include all cardioversion CV therapies in the ventricular x D tachycardia VT and fast ventricular tachycardia FVT zones as well as the first shock in the ventricular fibrillation x D VF zone Syncing to an R wave to deliver a shock is standard practice for transvenous ICDs and external x D defibrillators and is not unique to EV ICD This particular automatic sensing threshold issue is unique and isolated x D to EV ICD due to sensing thresholds being controlled by firmware Sensing behavior is defined in hardware for x D transvenous ICDs x D | Kindly refer to the letter | 2025-10-29 18:20
Medtronic Inc. | Aurora EV-ICD MRI SureScan | - | - | An event was reported by a Medtronic field representative that during routine defibrillation threshold testing at x D an EV ICD implant Ventricular fibrillation VF detection was appropriately met and the capacitors were charged x D to deliver a shock however following charge end the device was no longer showing sense markers and therapy x D was aborted The Medtronic representative stated that it was difficult to tell if the device lost sensing telemetry or other interaction Sensing resumed after seconds and redetection occurred but an external defibrillation shock was manually delivered to convert the rhythm prior to fulfilling device confirmation and therapy delivery A second induction attempt through the device resulted in a shock being delivered after charge end that successfully terminated the arrhythmia Medtronic internal review of the device data determined that the absence of sensing markers occurred due to a firmware timing interaction between the charge end and a ventricular sense VS event coincident with a drop in device measured R wave amplitudes x D x D When an EV ICD detects a ventricular arrhythmia and determines therapy is required the device charges the capacitors and delivers the shock If the end of the capacitor charge interrupts an in progress R wave amplitude measurement there is a potential for the dynamic sensing algorithm to remain at its first plateau of of the prior R wave measurement amplitude This may be impactful and lead to a delay in high voltage therapy if the amplitudes decrease at the time of the interrupt x D x D The root cause of the delayed therapy was determined to be VF undersensing due to a firmware timing x D interaction with the charge end event and a VS event coincident with a significant decrease in the amplitude of x D the VF signals that was not large enough to be sensed at the initial automatic sensing threshold The sensitivity x D remained at the initial automatic sensing threshold of of the prior sensed event amplitude without automatic x D adjustment due to charge end CE interrupting an in progress amplitude measurement When the amplitude of x D the subsequent signal falls below this sensing threshold as seen in the reported complaints delayed x D tachyarrhythmia therapy may result This happens because a non committed shock cannot reconfirm and sync to x D a ventricular sensed beat Non committed shocks include all cardioversion CV therapies in the ventricular x D tachycardia VT and fast ventricular tachycardia FVT zones as well as the first shock in the ventricular fibrillation x D VF zone Syncing to an R wave to deliver a shock is standard practice for transvenous ICDs and external x D defibrillators and is not unique to EV ICD This particular automatic sensing threshold issue is unique and isolated x D to EV ICD due to sensing thresholds being controlled by firmware Sensing behavior is defined in hardware for x D transvenous ICDs x D | Kindly refer to the letter | 2025-10-29 18:21
Illumina, Inc. | NextSeq 550Dx High Output Reagent Kit v2.5 (75 cycles) | 1 | - | Illumina has received several customer complaints describing leak events in NextSeq Dx sequencers where liquid was found x D under the flow cells Based on field feedback received the leaks were intermittent occurrences at the start middle and end of x D the sequencing runs x D After an investigation Illumina determined that this issue was due to loose debris found on flow cells gaskets received from the x D third party supplier of the flow cells The quality issue was related to a change of the third party supplier s manufacturing site x D The products impacted are the flow cells components of the NextSeq Dx sequencing reagents kits which are class A In vitro x D Diagnostic Medical device IVD intended for sequencing of DNA libraries on the NextSeq Dx Instrument | Illumina Actions We have completed a root cause investigation and implemented preventative action to prevent further recurrence. All impacted products in our inventory have been identified and contained. Illumina has notified the pertinent local and international regulatory bodies, including the Competent Authorities, of this issue. Required Customer Actions If you received reagent lots listed in the Appendix, dispose of these reagents per local regulations. See the NextSeq 550Dx Instrument Reference Guide for additional instructions on proper disposal. Please ensure proper personal protective equipment (PPE) is utilized including, but not limited to, eye protection, gloves, and a laboratory coat. If you have or suspect you have come in contact - directly or indirectly via skin or aerosol - with the reagents in the flow cell, please follow your laboratory's standard safety and first aid procedures immediately. Please complete the Verification Form to acknowledge that you have carried out the Required Customer Actions above. We kindly ask that you e-mail the completed form to Illumina Technical Support within 5 business days. Alternatively, you may call or e-mail Illumina Technical Support to provide the information requested on the form. | 2025-10-29 18:24
Illumina, Inc. | NextSeq 550Dx High Output Reagent Kit v2.5 (300 cycles) | 1 | - | Illumina has received several customer complaints describing leak events in NextSeq Dx sequencers where liquid was found x D under the flow cells Based on field feedback received the leaks were intermittent occurrences at the start middle and end of x D the sequencing runs x D After an investigation Illumina determined that this issue was due to loose debris found on flow cells gaskets received from the x D third party supplier of the flow cells The quality issue was related to a change of the third party supplier s manufacturing site x D The products impacted are the flow cells components of the NextSeq Dx sequencing reagents kits which are class A In vitro x D Diagnostic Medical device IVD intended for sequencing of DNA libraries on the NextSeq Dx Instrument | Illumina Actions We have completed a root cause investigation and implemented preventative action to prevent further recurrence. All impacted products in our inventory have been identified and contained. Illumina has notified the pertinent local and international regulatory bodies, including the Competent Authorities, of this issue. Required Customer Actions If you received reagent lots listed in the Appendix, dispose of these reagents per local regulations. See the NextSeq 550Dx Instrument Reference Guide for additional instructions on proper disposal. Please ensure proper personal protective equipment (PPE) is utilized including, but not limited to, eye protection, gloves, and a laboratory coat. If you have or suspect you have come in contact - directly or indirectly via skin or aerosol - with the reagents in the flow cell, please follow your laboratory's standard safety and first aid procedures immediately. Please complete the Verification Form to acknowledge that you have carried out the Required Customer Actions above. We kindly ask that you e-mail the completed form to Illumina Technical Support within 5 business days. Alternatively, you may call or e-mail Illumina Technical Support to provide the information requested on the form. | 2025-10-29 18:24
Getinge | CardioSave IABP | 148 | - | issues please refer to the letter for details | 4 issues, please refer to the letter for details | 2025-10-30 11:26
Smith & Nephew Inc. | LEGION TIBIAL OFFSET BUSHING | 2 | - | Smith Nephew Inc has initiated a Field Action to voluntarily remove one batch of LGN TIBIAL OFFSET BUSHING MM due to a manufacturing error A batch of LGN TIBIAL OFFSET BUSHING MM was laser etched incorrectly with MM instead of MM | 1. Ensure that the contents of this Field Safety Notice are read and understood by those within your organisation who may use LGN TIBIAL OFFSET BUSHING 4MM. 2. Locate and quarantine affected devices immediately. If you have further distributed the product to other organisations, please inform them at once of this Field Action and provide to them a copy of this letter. 3. Please complete the Customer Response form and email or fax it to your national Smith+Nephew agency/distributor. 4. Return quarantined product to your national Smith+Nephew agency/distributor. 5. Please maintain awareness of this notice and resulting action for an appropriate period to ensure the effectiveness of the corrective action. | 2025-10-30 16:50
Cepheid | Xpert CT/NG | 1430 | - | Cepheid has become aware of a rare genetic mutation event between Neisseria gonorrhoeae and x D Neisseria meningitidis that caused some Neisseria gonorrhoeae strains to lack target sites for the x D Cepheid Xpert CT NG test leading to false negative results | Cepheid recommends that laboratories and clinics using the Xpert® CT/NG test as their molecular method for the detection of N. gonorrhoeae monitor for decreases in the positivity rate and ensure that discordant laboratory results (e.g. testing negative on the Xpert® CT/NG test but positive for NG by culture/microscopy or testing negative on Xpert® CT/NG test but positive for NG using anther NAAT be investigated further and reported to Cepheid. Cepheid also asks that you acknowledge receipt of this Urgent Field Safety Notice by completing the attached Customer Response Form and either emailing the completed form to CFQ@cepheid.com or faxing it to +1 (408) 716-3143. Cepheid will not send a reply or confirmation of received response forms. | 2025-10-30 17:10
Beckman Coulter Inc. | Access Immunoassay System Reaction Vessels | 217 | - | Beckman Coulter has determined that Access RVs lots listed above may contain manufacturing deformities The bottom round end of the RV may have a small protrusion or hair like deformity | Check your inventory of Access 2 RVs for the lots listed at the top of this letter. • If you have the lots listed above contact Beckman coulter for replacement RVs. • You may continue to use the lots listed above until you have received a sufficient supply of replacement RVs. Once replacement RVs arrive, then dispose of the lots listed above according to your local regulations. • If you do not have the RV lots listed above, then no further action is needed. Retrospective review of results is not necessary because this issue does not affect individual patient results. | 2025-11-02 15:59
SLE Limited | SLE1500 Gas Blender | 40 | - | SLE have identified an issue with the SLE Interflow Gas Blenders that can potentially cause the device to enter a repeating cycle of resets | SLE are investigating the root cause of this issue and will provide further updates in the coming weeks once a solution has been identified. The decision has been made to initiate this FSCA in order to communicate this issue to the field immediately while investigation into the root cause continues. Further Corrective Action will be communicated via an update form. | 2025-11-03 13:06
SLE Limited | SLE1500 Gas Blender | 41 | - | SLE have identified an issue with the SLE Interflow Gas Blenders that can potentially cause the device to enter a repeating cycle of resets | SLE are investigating the root cause of this issue and will provide further updates in the coming weeks once a solution has been identified. The decision has been made to initiate this FSCA in order to communicate this issue to the field immediately while investigation into the root cause continues. Further Corrective Action will be communicated via an update form. | 2025-11-03 13:06
Philips Medizin Systeme Boeblingen GmbH | IntelliVue MX800 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:19
Philips Medizin Systeme Boeblingen GmbH | IntelliVue MX700 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:19
Philips Medizin Systeme Boeblingen GmbH | MX500 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:19
Philips Medizin Systeme Boeblingen GmbH | MX550 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:19
Philips Medizin Systeme Boeblingen GmbH | MX400 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:19
Philips Medizin Systeme Boeblingen GmbH | MX450 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:19
Philips Medizin Systeme Boeblingen GmbH | MX750 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:19
Philips Medizin Systeme Boeblingen GmbH | IntelliVue MP20 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:19
Philips Medizin Systeme Boeblingen GmbH | IntelliVue MP30 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:19
Philips Medizin Systeme Boeblingen GmbH | IntelliVue MP40 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:19
Philips Medizin Systeme Boeblingen GmbH | IntelliVue MP5 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:19
Philips Medizin Systeme Boeblingen GmbH | IntelliVue MP50 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:19
Philips Medizin Systeme Boeblingen GmbH | IntelliVue MP60, IntelliVue MP70, IntelliVue MP80, IntelliVue MP90 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:19
Philips Medizin Systeme Boeblingen GmbH | IntelliVue MX800 | 24197 | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:30
Philips Medizin Systeme Boeblingen GmbH | IntelliVue Patient Monitor MX750, MX700, MX550, MX500, MX400, MX450 | 24197 | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:30
Philips Medizin Systeme Boeblingen GmbH | IntelliVue MP20 | 24197 | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:30
Philips Medizin Systeme Boeblingen GmbH | IntelliVue MP30 IntelliVue MP40 IntelliVue MP5 IntelliVue MP50 IntelliVue MP60 IntelliVue MP70 IntelliVue MP80 IntelliVue MP90 | 24197 | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:30
Philips Medizin Systeme Boeblingen GmbH | IntelliVue MX800 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:31
Philips Medizin Systeme Boeblingen GmbH | IntelliVue Patient Monitor MX750, MX700, MX550, MX500, MX400, MX450 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:31
Philips Medizin Systeme Boeblingen GmbH | IntelliVue MP20 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:31
Philips Medizin Systeme Boeblingen GmbH | IntelliVue MP30 IntelliVue MP40 IntelliVue MP5 IntelliVue MP50 IntelliVue MP60 IntelliVue MP70 IntelliVue MP80 IntelliVue MP90 | - | - | Philips IntelliVue Patient Monitors provide a feature for pausing all alarms for a configured timeframe or for switching all alarms off for an infinite period until they are switched back on by the user x D The infinite Alarms Off feature is an important functionality in specific clinical environments such as the Operating Room Cath lab or during resuscitations to reduce the audible alarms which may present unnecessary distractions during closely monitored procedures or interventions x D x D Philips has received complaints that the IntelliVue monitors did not alarm In some cases these complaints alleged delayed patient treatment due to the monitors not alarming and in some cases there were reports of severe patient harm or death It was determined that in these instances the Alarms had been switched off for an infinite period and the user may have been unaware that the alarms were off x D | Philips is informing customers of the observed errors and related hazards of enabling infinite Alarms Off, including providing instructions on how to configure the device alarm settings. Philips recommends reviewing the configuration of the monitors with infinite Alarms Off functionality enabled and reconfigure settings to require additional confirmation to ensure users are aware of their action when activating infinite Alarms Off. For IntelliVue MP5 models only: Philips is developing a software update which will enforce the additional confirmation step when the user enables infinite Alarms Off. To update your MP5 software to include this feature, please reach out to your local Philips representative. Philips is also providing an educational video on this topic. Scan the QR code or visit learn.philips.com to access the educational video on how to configure and use the infinite alarms feature in the Philips Learning Center (PLC). Enter your PLC username and password or create an account and then enter this enrollment key to access the video: | 2025-11-03 13:31
Philips Medical Systems Nederland B.V. | Zenition 10 | 4 | - | Philips has identified that during maintenance activities excessive bending forces may be applied to the ethernet cable connector transmitting image data from the detector to the system This may cause the cable to be pulled out of the connector s strain relief potentially leading to intermittent disconnections between the internal wires and the connector pins If one of the connecting pins loses connection live X ray image will no longer be displayed and the error message Doselink Error X ray aborted will be shown on the user interface In some instances a restart of the C arm stand may temporarily resolve the issue | Philips will be replacing the C-Arc Cable in the affected systems. Your local Philips representative will contact you to schedule a visit to replace the C-Arc Cable in the affected systems. Philips will start replacement of the C-arc cables in Q2 2026 (reference FCO71800116). | 2025-11-03 13:50
Philips Medical Systems Nederland B.V. | Zenition 30 | 4 | - | Philips has identified that during maintenance activities excessive bending forces may be applied to the ethernet cable connector transmitting image data from the detector to the system This may cause the cable to be pulled out of the connector s strain relief potentially leading to intermittent disconnections between the internal wires and the connector pins If one of the connecting pins loses connection live X ray image will no longer be displayed and the error message Doselink Error X ray aborted will be shown on the user interface In some instances a restart of the C arm stand may temporarily resolve the issue | Philips will be replacing the C-Arc Cable in the affected systems. Your local Philips representative will contact you to schedule a visit to replace the C-Arc Cable in the affected systems. Philips will start replacement of the C-arc cables in Q2 2026 (reference FCO71800116). | 2025-11-03 13:50
Groupe Lépine | THREADED TROCAR PIN Ø3.2 NR | 8 | - | absence of mechanical stop | recall and distract it all items are in our warehouse | 2025-11-03 14:01
Groupe Lépine | TROCAR PIN Ø3.2 NR | 8 | - | absence of mechanical stop | recall and distract it all items are in our warehouse | 2025-11-03 14:01
BD | BD BACTEC™ MGIT™ 960 PZA Kit | - | - | Description of the product problem x D x D BD is pleased to announce that we have resumed production of a modified version of the BD BACTEC MGIT PZA Kit which includes modified inoculation methods and reduced shelf life x D x D BD has conducted a thorough review of the performance from all supported inoculum sources using the synthetic raw material and the following change has been implemented for the BD BACTEC MGIT PZA Kit x D x D Only inoculum prepared from MGIT tubes days post positivity can be currently supported x D Inoculum prepared from MGIT tubes days past instrument positivity are currently not supported x D Inoculum prepared from solid media are currently not supported x D x D Additionally the shelf life of the product has been adjusted from months to months to reflect the latest internal supporting data x D | BD Actions: 1. BD will continue to work with global regulatory authorities to reintroduce BD BACTEC™ MGIT™ 960 PZA Kit. BD will resume receiving and fulfilling customer orders for this product in markets where permitted, based upon regulatory requirements while maintaining ongoing quality improvements and performance monitoring. 2. BD continues to investigate long-term solutions and will actively work with global regulatory authorities to reintroduce the BD BACTEC™ MGIT™ 960 PZA Kit with the original claims. To date, BD does not plan to initiate any further advice or information in a follow-up FSN. Customer Actions: • Review the information in this notice. • Complete and return the Customer Response Form even if you no longer have any inventory remaining in your facility by 1st December 2025. • This notice needs to be passed on to all those who need to be aware within your organisation or to any organisation which this action has an impact. • Please maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. • Please report all device-related incidents to the manufacturer, distributor or local representative, and the national Competent Authority if appropriate, as this provides important feedback. Distributor Actions: • Review the information in this notice. • This notice needs to be passed on to all those who need to be aware within your organisation or to any organisation which this action has an impact. • Identify the facilities where you have distributed the BD BACTEC™ MGIT™ 960 PZA Kit and notify them immediately of this notice. o Have your customers complete and return the Customer Response form to your organisation for reconciliation purposes by 1st December 2025. o There is no requirement to return your customer response forms to BD, you should maintain these on file at your facility. Return only your final consolidated response form. • Complete and return the Customer Response Form following completion of your reconciliation activities. • Please maintain awareness on this notice and resulting action for an appropriate period to ensure effectiveness of the corrective action. • Please report all device-related incidents to the manufacturer, or local representative, and the national Competent Authority if appropriate, as this provides important feedback. | 2025-11-04 08:27
Osatu S.Coop. | BEXEN CARDIO REANIBEX 300 | 18 | - | Hazard Device may fail to turn on preventing defibrillation use x D Probability Low based on field data x D Potential harm Delay or failure in delivering life saving therapy possibly leading to patient death x D | We have reached out to all affected customers and attached received responses. | 2025-11-05 10:00
Osatu S.Coop. | BEXEN CARDIO REANIBEX Serie 500 | 18 | - | Hazard Device may fail to turn on preventing defibrillation use x D Probability Low based on field data x D Potential harm Delay or failure in delivering life saving therapy possibly leading to patient death x D | We have reached out to all affected customers and attached received responses. | 2025-11-05 10:00
LivaNova USA, Inc. | Symmetry | 6 | - | The purpose of this letter is to share information regarding the Model Symmetry VNS Therapy Generator which is designed to deliver Vagus Nerve Stimulation for chronic or recurrent depression x D As of September LivaNova has received two reports indicating that the battery status may temporarily display a full battery level i e indicator when a range would be expected Importantly this behavior does not affect the delivery of VNS Therapy The generator will accurately deliver therapy as programmed and the overall longevity of the generator battery is unaffected x D The reported impact has been limited to user dissatisfaction with the battery status display There have been no reports of patient harm and therapy delivery remains unaffected The generator battery longevity is unaffected by this issue All clinically actionable battery indicators low battery indicators such as Intensified Follow up Indicator IFI Near End of Service NEOS and End of Service EOS are accurately presented The temporary fluctuation does not interfere with the clinical use and will self correct i e resolve prior to display of the IFI NEOS and EOS indicators x D The Model battery status indicator may x D x D Inaccurately show depletion at an accelerated rate involving the and battery indicators early in product life x D x D Fluctuate and temporarily display full generator battery status i e indicator when a range would be expected This issue presents in the middle of product life x D x D As a reminder product labeling recommends the following actions for the lower battery indicators and importantly each of these indicators is accurately presented by Model x D x D Remaining i e indicator Battery level is within normal operating range and no action is required x D x D Intensified Follow up Indicator IFI Battery has depleted to a level where more frequent clinical monitoring is recommended x D x D Near End of Service NEOS Generator should be replaced as soon as possible x D x D End of Service EOS Generator is no longer supplying stimulation and immediate replacement is recommended If the generator is not replaced it will eventually lose the ability to communicate with the software x D x D This behavior is limited to Model generators listed in Attachment of this letter No other VNS Therapy generator models are affected x D | Update the systems | 2025-11-05 14:47
Rocket Medical PLC | Rocket Seldinger | 148 | - | The Instructions for Use IFU ZDOCK for the devices listed above contains the instruction The chest tube can now be secured to the skin by using the fixation device R PG FX or suture or both and is ready for connection to a suitable sealed drainage system such as the Rocket Blue Bottle or Rocket Ambulatory Bag The fixation device mentioned is no longer supplied by Rocket Medical and Rocket Medical determines that securing the device by suture is a mandatory action to prevent risk of disconnection | The IFU has been updated to include revised guidance regarding the need for a suture, in steps 14 and 15. The updated text now states: "The chest tube must now be secured to the skin by suture and is ready for connection to a suitable sealed drainage system such as the Rocket Blue Bottle or Rocket Ambulatory Bag. Further support to the catheter can be provided by the fixation device." Customers will be contacted to inform them of the update to the instructions and provided with an updated IFU. | 2025-11-06 08:08
HeartSine Technologies Ltd | HeartSine Pad-Pak | 734 | - | The connection issues that may arise from improper insertion of the Pad Pak are not always obvious to the user when the Pad Pak is inserted into the HeartSine samaritan PAD device If the Pad Pak is not properly inserted into the HeartSine samaritan PAD device or if the Pad Pak locator pins are bent the device may fail to deliver the intended therapy during use potentially leading to a delay in treatment or no treatment being delivered during use A delay in treatment or no treatment may result in serious injury or patient expiry | Stryker is notifying all customers who have received affected Pad-Paks to perform the actions outlined above by 30-10-2025. On behalf of Stryker, we thank you sincerely for your response and regret any inconvenience this may have caused. In line with the recommendations of the Medical Device Coordination Group Guidance document Ref MDCG 2023-3 and EU MDR 2017/745, we can confirm that this FSCA has been notified appropriately to the National Competent Authority for your country. | 2025-11-09 10:50
HeartSine Technologies Ltd | HeartSine Pediatric-Pak | 734 | - | The connection issues that may arise from improper insertion of the Pad Pak are not always obvious to the user when the Pad Pak is inserted into the HeartSine samaritan PAD device If the Pad Pak is not properly inserted into the HeartSine samaritan PAD device or if the Pad Pak locator pins are bent the device may fail to deliver the intended therapy during use potentially leading to a delay in treatment or no treatment being delivered during use A delay in treatment or no treatment may result in serious injury or patient expiry | Stryker is notifying all customers who have received affected Pad-Paks to perform the actions outlined above by 30-10-2025. On behalf of Stryker, we thank you sincerely for your response and regret any inconvenience this may have caused. In line with the recommendations of the Medical Device Coordination Group Guidance document Ref MDCG 2023-3 and EU MDR 2017/745, we can confirm that this FSCA has been notified appropriately to the National Competent Authority for your country. | 2025-11-09 10:50
Draeger Medical, Inc. | Vapor 2000 | 16 | - | During the course of our production controls Dr ger detected a certain x D component of the Vapor and Vapor was not delivered by the x D relevant supplier in accordance with the specification Impurities resulting x D from the production process at the supplier site were found inside a certain x D production lot of the Vaporizers | Phase 1: Distribution of Safety Notice until 2025-11-19 Phase 2: Replacement of the affected product(s) until 2026-02-15 | 2025-11-09 15:25
Medtrade products Ltd | Celox PPH | 390 | - | Medtrade Products has produced a revised IFU contains the following changes this FSN is to provide the information prior to the new IFU being available x D Warning x D This product contains chitosan derived from shellfish clinicians should exercise caution with individuals with known shellfish allergies x D Postpartum haemorrhage PPH is a severe and multifaceted condition requiring the coordinated use of various medical interventions in very rare instances patients have reported transient disturbances in vision and auditory function While no definitive causal relationship to specific treatments has been established any occurrence of these symptoms warrants prompt medical evaluation and management | Action To Be Taken by the User >Take note of amendment/reinforcement of Instructions for Use (IFU) -Continue using the product as intended in accordance with IFU with the additional warnings identified in the FSN until the updated IFU becomes available -Monitor for any sign and symptoms related to the reported effect and report any suspected adverse events. Action Being Taken by the Manufacturer -IFU or labelling change The IFU has been amended and product made available for distribution, this FSN is to identify the changes prior to product containing the revised IFU being available to each user at each hospital/clinic | 2025-11-10 11:22
Elekta Solutions AB | Elekta Evo Elekta Harmony Pro Versa HD Elekta Infinity Elekta Harmony Elekta Synergy | 11 | - | The root cause of the issue has been traced to the absence of clear work instructions for the crimping process combined with insufficient training In the absence of proper guidance workers at the supplier s site have been crimping the cable in the incorrect orientation which has directly led to the observed failures | Elekta will send an Important Field Safety Notice 200-01-103-089 to all affected customers from 14-NOV-25. The notice informs users of the specific product and version numbers affected by the issue, and any work around that can be used to avoid the issue. (If applicable)  Sites affected will be all digital Linear Accelerators (Synergy, Infinity, Versa HD, Elekta Evo) up to and including SN 157644, Harmony & Harmony Pro up to and including SN 451054. A copy of the Important Field Safety Notice is to be kept with the most current labelling and all personnel working with the product should be made aware of the content of the letter. The customers are instructed to complete and return the Important Field Safety Notice Acknowledgement form to their local Elekta representative as soon as possible but no later than 30 days from receipt. Elekta will release an Important Field Safety Modification (IFSM) to identify potentially faulty discharge rod crimping and provide further instructions. The modification will include procedures for verifying the reliability of the discharge rod crimping and replacing affected components. The target date for release is March 2026. | 2025-11-10 15:21
Beckman Coulter Ireland, Inc | BICARBONATE | 283 | - | Beckman Coulter has received complaints for elevated results generated using Bicarbonate reagent when there are high levels of Lactate Dehydrogenase LDH in the test sample x D Beckman Coulter has investigated and confirmed that the lots listed above of Bicarbonate reagent may generate Bicarbonate results that are biased high due to interference with LDH in the test sample x D The level of LDH interference is not consistent across lots of Bicarbonate reagents The magnitude of the impact varies depending on Bicarbonate lot in use and the level of LDH in the test sample | • Discontinue use and dispose all Bicarbonate OSR6x37/OSR6x90 lots listed in Table A in the IMPACT section of this letter, following your laboratory protocols and local regulations. Contact your local Beckman Coulter representative for replacement or re-imbursement. • For Bicarbonate OSR6x37/OSR6x90 lots listed in Table B in the IMPACT section, if elevated LDH levels are suspected, please assess results in conjunction with patient’s LDH values. • Share the content of this letter with your laboratory and/or Medical Director to evaluate the requirement for a retrospective review of Bicarbonate results for neonatal patients and patients with elevated LDH levels. | 2025-11-10 23:32
Ortho Clinical Diagnostics Inc. | Performance Verifier I | - | - | QuidelOrtho received complaints regarding results below the published Range of Means ROM when attempting to perform QC using Lot G of VITROS PV I on lots from GENs of VITROS Na Slides x D We investigated the issue and determined that the ROM published for GENs of VITROS Na Slides on the Control Assay Sheet for VITROS PV I Lot G were incorrect We have performed additional testing and revised the ROM for GENs x D Customers are advised to discontinue use and discard the previous Control Assay Sheet for VITROS PV I Lot G which contains the incorrect ROM for GENs of VITROS Na Slides and download the revised Control Assay Sheet for VITROS PV I Lot G which contains the correct ROM for GENs of VITROS Na Slides | • Discontinue use and discard the previous Control Assay Sheet for VITROS PV I, Lot G2905, containing the incorrect ROM for GENs 81-84 of VITROS Na+ Slides. • Download the revised Control Assay Sheet for VITROS PV I, Lot G2905, containing the correct ROM for GENs 81-84 of VITROS Na+ Slides. • Complete the enclosed Confirmation of Receipt form no later than Month DD, 2025. • Save this notification with your User Documentation. • Please forward this notification if the affected product was distributed outside of your facility. • If your laboratory has experienced the issue described in this notification and you have not already done so, please report the occurrence to your local Technical Support team. | 2025-11-11 10:02
Ortho Clinical Diagnostics Inc. | Performance Verifier I | - | - | QuidelOrtho received complaints regarding results below the published Range of Means ROM when attempting to perform QC using Lot G of VITROS PV I on lots from GENs of VITROS Na Slides x D We investigated the issue and determined that the ROM published for GENs of VITROS Na Slides on the Control Assay Sheet for VITROS PV I Lot G were incorrect We have performed additional testing and revised the ROM for GENs x D Customers are advised to discontinue use and discard the previous Control Assay Sheet for VITROS PV I Lot G which contains the incorrect ROM for GENs of VITROS Na Slides and download the revised Control Assay Sheet for VITROS PV I Lot G which contains the correct ROM for GENs of VITROS Na Slides | • Discontinue use and discard the previous Control Assay Sheet for VITROS PV I, Lot G2905, containing the incorrect ROM for GENs 81-84 of VITROS Na+ Slides. • Download the revised Control Assay Sheet for VITROS PV I, Lot G2905, containing the correct ROM for GENs 81-84 of VITROS Na+ Slides. • Complete the enclosed Confirmation of Receipt form no later than Month DD, 2025. • Save this notification with your User Documentation. • Please forward this notification if the affected product was distributed outside of your facility. • If your laboratory has experienced the issue described in this notification and you have not already done so, please report the occurrence to your local Technical Support team. | 2025-11-11 10:02
Ortho Clinical Diagnostics Inc. | Performance Verifier I | - | - | QuidelOrtho received complaints regarding results below the published Range of Means ROM when attempting to perform QC using Lot G of VITROS PV I on lots from GENs of VITROS Na Slides x D We investigated the issue and determined that the ROM published for GENs of VITROS Na Slides on the Control Assay Sheet for VITROS PV I Lot G were incorrect We have performed additional testing and revised the ROM for GENs x D Customers are advised to discontinue use and discard the previous Control Assay Sheet for VITROS PV I Lot G which contains the incorrect ROM for GENs of VITROS Na Slides and download the revised Control Assay Sheet for VITROS PV I Lot G which contains the correct ROM for GENs of VITROS Na Slides | • Discontinue use and discard the previous Control Assay Sheet for VITROS PV I, Lot G2905, containing the incorrect ROM for GENs 81-84 of VITROS Na+ Slides. • Download the revised Control Assay Sheet for VITROS PV I, Lot G2905, containing the correct ROM for GENs 81-84 of VITROS Na+ Slides. • Complete the enclosed Confirmation of Receipt form no later than Month DD, 2025. • Save this notification with your User Documentation. • Please forward this notification if the affected product was distributed outside of your facility. • If your laboratory has experienced the issue described in this notification and you have not already done so, please report the occurrence to your local Technical Support team. | 2025-11-11 10:02
Olympus | SINGLE USE LIGATING DEVICE POLYLOOP | 869 | - | The Single Use Ligating Device PolyLoop HX U may fail to release or detach as expected during use causing the ligation loop to unintentionally remain anchored around patient anatomy x D Olympus received reports of serious injuries due to the loop not releasing x D The failures were caused by x D Movement of the yellow tube joint cylinder away from the device handle x D Forceful advancement of the slider when resistance is encountered x D | Read and follow the Field Safety Notice and IFU warnings. Always have Olympus loop cutter ready when using the PolyLoop. Avoid moving the yellow tube joint away from the handle or forcing the slider. Train all personnel and acknowledge receipt of the notice. Forward the notice to customers if the device was redistributed and report complaints to omea-businesssupport@olympus.com. | 2025-11-12 12:06
Ortho Clinical Diagnostics | ORTHO VISION Analyzer | - | - | QuidelOrtho s internal investigation identified that the regional setting on a small population of ORTHO VISION BioVue Analyzers was set to OCD an analyzer configuration intended for internal use only The regional setting is applied during analyzer installation based on the location where the installation is being performed This may also occur if the analyzer is reformatted or reloaded with new software If left with OCD all tests may be processed irrespective of your region s intended accessibility x D x D For example the setting allows the ORTHO Sera Anti N test to be run on the BioVue Analyzer although it has not been validated for use on the ORTHO VISION BioVue automated platform ORTHO Sera Anti N has been validated for use with manual BioVue cassette testing ORTHO Sera Anti N must not be processed on an ORTHO VISION BioVue Analyzer unless processed in a User Defined Protocol x D | • Do not process the Ortho Sera Anti-N assay on Vision BioVue Analyzers unless it is part of a User Defined Protocol (UDP) - per labeling on the product insert. • Complete the enclosed Confirmation of Receipt form no later than Month DD, 2025. | 2025-11-12 14:47
Ortho Clinical Diagnostics | ORTHO VISION Max Analyzer | - | - | QuidelOrtho s internal investigation identified that the regional setting on a small population of ORTHO VISION BioVue Analyzers was set to OCD an analyzer configuration intended for internal use only The regional setting is applied during analyzer installation based on the location where the installation is being performed This may also occur if the analyzer is reformatted or reloaded with new software If left with OCD all tests may be processed irrespective of your region s intended accessibility x D x D For example the setting allows the ORTHO Sera Anti N test to be run on the BioVue Analyzer although it has not been validated for use on the ORTHO VISION BioVue automated platform ORTHO Sera Anti N has been validated for use with manual BioVue cassette testing ORTHO Sera Anti N must not be processed on an ORTHO VISION BioVue Analyzer unless processed in a User Defined Protocol x D | • Do not process the Ortho Sera Anti-N assay on Vision BioVue Analyzers unless it is part of a User Defined Protocol (UDP) - per labeling on the product insert. • Complete the enclosed Confirmation of Receipt form no later than Month DD, 2025. | 2025-11-12 14:47
Breas Medical Ltd | Clearway 2 | 8 | - | A firmware defect in firmware | Update the firmware | 2025-11-13 10:49
Saudi Mais for Medical Products | Sterile Connectors for Blood Management | 133727 | - | observing hair in the medical product | The product Sterile Connectors for Blood Management with reference 215-0000 and lot number 221204 delivered shall be returned to the manufacturer. | 2025-11-13 13:40
Saudi Mais for Medical Products | Sterile Connectors for Blood Management | 133727 | - | observing hair in the medical product | The product Sterile Connectors for Blood Management with reference 215-0000 and lot number 221204 delivered shall be returned to the manufacturer. | 2025-11-13 13:40
SCHILLER Medical SAS | Defigard touch 7 | 41 | - | As per attached FSN | As per attached FSN | 2025-11-13 14:12
Datex-Ohmeda, Inc. | Carestation 650c A1 | 11 | - | GE HealthCare has become aware of the potential unexpected shutdown of Carestation and Series Anesthesia Systems containing certain power management boards see Affected Product Details below if the AC mains power is unplugged or in the event of an AC mains power failure x D x D Anesthesia systems only operate on battery power in a rare event that AC mains power is lost and there is no continuous backup emergency power If AC power is interrupted the system will not automatically switch over to the battery supply mode and will reboot If this issue occurs a temporary disruption of mechanical ventilation manual ventilation and volatile agent delivery may occur Following the reboot the system will not return to the previous ventilation settings x D x D If this situation is not identified and addressed by the user the loss of ventilation may be life threatening x D x D There have been no injuries reported to GE HealthCare as a result of this issue x D | Device Correction by GEHC Service Representative – To be updated for Phase 2 | 2025-11-17 10:10
Datex-Ohmeda, Inc. | Carestation 750 A1 | 11 | - | GE HealthCare has become aware of the potential unexpected shutdown of Carestation and Series Anesthesia Systems containing certain power management boards see Affected Product Details below if the AC mains power is unplugged or in the event of an AC mains power failure x D x D Anesthesia systems only operate on battery power in a rare event that AC mains power is lost and there is no continuous backup emergency power If AC power is interrupted the system will not automatically switch over to the battery supply mode and will reboot If this issue occurs a temporary disruption of mechanical ventilation manual ventilation and volatile agent delivery may occur Following the reboot the system will not return to the previous ventilation settings x D x D If this situation is not identified and addressed by the user the loss of ventilation may be life threatening x D x D There have been no injuries reported to GE HealthCare as a result of this issue x D | Device Correction by GEHC Service Representative – To be updated for Phase 2 | 2025-11-17 10:10
Natejah Industrial Company | NATEJAH One Step Malaria (Pf/Pan) Rapid Test | - | - | A manufacturing defect is causing inaccurate results in the test | Withdraw all affected products. | 2025-11-18 12:42
Natejah Industrial Company | NATEJAH One Step Malaria (Pf/Pan) Rapid Test | 113120 | - | The false positive PAN line was caused by a variation in the raw material supplied for NATEJAH One Step Malaria Pf Pan Rapid Test No deviations were found in manufacturing equipment or storage conditions | 1. Immediately stop the use and distribution of all affected lots. 2. Quarantine any remaining stock from the listed lots. 3. Inform all end-users, sub-distributors, hospitals, and laboratories who have received these lots. 4. Arrange the return of all affected kits to Natejah Industrial Company for replacement or proper disposal. | 2025-11-18 16:14
Roche Diagnostics | Elecsys Anti-TSHR assay | - | - | The root cause for the variations in the Elecsys Anti TSHR assay was an insufficiency of the assay s routine x D x D production process particularly for results around the IU L Medical Decision Point For the Platform to x D Platform deviation cobas e vs cobas e e differences arose due to an insufficient alignment concept x D x D between the platforms These variations were further influenced by updates to the routine production process for x D this assay For the Lot to Lot deviation on e measures taken to enhance robustness led to additional shifts | Stop using the following lots of Elecsys Anti-TSHR assay and switch to the lots as indicated in the table below: Affected Systems Affected lots (stop using) New unaffected lots (switch to) cobas e 411 Lot 840177 Lot 906872 cobas e 402 cobas e 801 Lot 840183 and Lot 874011 Lot 908953 | 2025-11-19 09:39
Roche Diagnostics | Elecsys Anti-TSHR assay | 90 | - | The root cause for the variations in the Elecsys Anti TSHR assay was an insufficiency of the assay s routine x D x D production process particularly for results around the IU L Medical Decision Point For the Platform to x D Platform deviation cobas e vs cobas e e differences arose due to an insufficient alignment concept x D x D between the platforms These variations were further influenced by updates to the routine production process for x D this assay For the Lot to Lot deviation on e measures taken to enhance robustness led to additional shifts | Stop using the following lots of Elecsys Anti-TSHR assay and switch to the lots as indicated in the table below: Affected Systems Affected lots (stop using) New unaffected lots (switch to) cobas e 411 Lot 840177 Lot 906872 cobas e 402 cobas e 801 Lot 840183 and Lot 874011 Lot 908953 Note that cobas e 601 and cobas e 602 are not affected and can be used with all currently available lots and with the upcoming lots without restrictions. | 2025-11-19 09:42
Philips Medical Systems Nederland B.V. | Achieva 1.5T | 55 | - | Philips has identified that MRI images may display an incorrect orientation where the patient s hands are reversed showing the left hand as the right and the right hand as the left This issue may occur if both of the following conditions are met x D The patient is head first head facing the bore of the MRI scanner in the supermanprone or superman supine position with both hands raised over their head and placed inside a coil x D The exam is set up for prone position laterality set to both anatomic region set to hand arm Philips has identified additional software defects that have no associated safety impact but may impact clinical workflow | Philips will contact you to schedule time for a Field Service Engineer (FSE) to install a software upgrade to resolve the software issues (FCO78100618, FCO78100619, FCO78100626, FCO78100627, FCO78100628, FCO78100629, FCO78100635). | 2025-11-19 14:45
Philips Medical Systems Nederland B.V. | Ingenia 1.5T, Ingenia 1.5T CX, Ingenia 1.5T S, Ingenia 3.0T, Ingenia Ambition S, Ingenia Ambition X, Ingenia Elition X | 55 | - | Philips has identified that MRI images may display an incorrect orientation where the patient s hands are reversed showing the left hand as the right and the right hand as the left This issue may occur if both of the following conditions are met x D The patient is head first head facing the bore of the MRI scanner in the supermanprone or superman supine position with both hands raised over their head and placed inside a coil x D The exam is set up for prone position laterality set to both anatomic region set to hand arm Philips has identified additional software defects that have no associated safety impact but may impact clinical workflow | Philips will contact you to schedule time for a Field Service Engineer (FSE) to install a software upgrade to resolve the software issues (FCO78100618, FCO78100619, FCO78100626, FCO78100627, FCO78100628, FCO78100629, FCO78100635). | 2025-11-19 14:45
Philips Medical Systems Nederland B.V. | MR 5300 | 55 | - | Philips has identified that MRI images may display an incorrect orientation where the patient s hands are reversed showing the left hand as the right and the right hand as the left This issue may occur if both of the following conditions are met x D The patient is head first head facing the bore of the MRI scanner in the supermanprone or superman supine position with both hands raised over their head and placed inside a coil x D The exam is set up for prone position laterality set to both anatomic region set to hand arm Philips has identified additional software defects that have no associated safety impact but may impact clinical workflow | Philips will contact you to schedule time for a Field Service Engineer (FSE) to install a software upgrade to resolve the software issues (FCO78100618, FCO78100619, FCO78100626, FCO78100627, FCO78100628, FCO78100629, FCO78100635). | 2025-11-19 14:45
Philips Medical Systems Nederland B.V. | SmartPath to dStream for 1.5T | 55 | - | Philips has identified that MRI images may display an incorrect orientation where the patient s hands are reversed showing the left hand as the right and the right hand as the left This issue may occur if both of the following conditions are met x D The patient is head first head facing the bore of the MRI scanner in the supermanprone or superman supine position with both hands raised over their head and placed inside a coil x D The exam is set up for prone position laterality set to both anatomic region set to hand arm Philips has identified additional software defects that have no associated safety impact but may impact clinical workflow | Philips will contact you to schedule time for a Field Service Engineer (FSE) to install a software upgrade to resolve the software issues (FCO78100618, FCO78100619, FCO78100626, FCO78100627, FCO78100628, FCO78100629, FCO78100635). | 2025-11-19 14:45
Cepheid | Xpert BCR-ABL Ultra | 1531 | - | Using Xpert BCR ABL Ultra test kits with specimens that have markedly elevated white blood x D cell count cells L may potentially cause inaccurate results | Cepheid recommends that laboratories and clinical facilities utilizing the Xpert® BCR-ABL Ultra assay for quantification of BCR::ABL1 and ABL1 mRNA transcript levels in peripheral blood specimens from t(9;22)-positive chronic myeloid leukemia (CML) patients undergoing tyrosine kinase inhibitor (TKI) therapy, assess white blood cell count (WBCC) prior to testing in cases where elevated white blood cell count is suspected. Specifically, if the WBCC exceeds 30,000 cells/μL, the specimen should be processed in accordance with the “Retest Procedure for ERROR (Code 2008) or INVALID (Type 2)” as detailed in the “Retests” section of the IFU for the following scenarios: - Cases where INVALID test results are observed and are associated with an elevated WBCC. - Cases of valid results in which an elevated WBCC has been observed. This recommendation is intended to mitigate the possibility that viscous samples could yield a falsely low or false negative result. | 2025-11-20 14:38
Beckman Coulter Inc. | UniCel DxI 600 Access Immunoassay System | 185 | - | Beckman Coulter has determined that due to a potential communication x D issue the analyzers listed above may not successfully restart following a x D standard shutdown even after multiple attempts If this issue occurs the x D instrument will be unable to complete initialization and transition to Ready x D mode | • Follow the restart process documented in Section 2 of the "Unicel DxI Access Immunoassay System Instructions for Use”. • If your analyzer continues to experience this issue, please contact Beckman Coulter Technical Support. • Beckman Coulter recommends posting this letter on or near the affected analyzers. | 2025-11-23 23:45
Beckman Coulter Inc. | UniCel DxI 800 Access Immunoassay System | 185 | - | Beckman Coulter has determined that due to a potential communication x D issue the analyzers listed above may not successfully restart following a x D standard shutdown even after multiple attempts If this issue occurs the x D instrument will be unable to complete initialization and transition to Ready x D mode | • Follow the restart process documented in Section 2 of the "Unicel DxI Access Immunoassay System Instructions for Use”. • If your analyzer continues to experience this issue, please contact Beckman Coulter Technical Support. • Beckman Coulter recommends posting this letter on or near the affected analyzers. | 2025-11-23 23:45
SAUDI MAIS CO FOR MEDICAL PRODUCTS | Central Venous Catheter Set (Double Lumen) | 1888 | - | The label issue occurred due to incorrect printing | According to the investigation , we made recall for the remaining quantity with the Customer. | 2025-11-24 10:56
HILL-ROM S.A.S | Accella Therapy | 66 | - | reports of technical box cracks that can cause air leaks If this issue occurs the mattress pressure can stay in a higher pressure phase or a non functioning Alternating Low Pressure ALP mode The mattress will provide visible and audible alerts Error codes or | 1. If the care provider experiences the error codes listed above on an affected Accella mattress (or a mattress used with Spare Parts associated with a Progressa frame), please discontinue use of the product until a technical service representative can replace it. If the error code alarms while a patient is on the mattress, please move the patient to an alternative surface. If the error codes do not occur, the mattress is safe to continue use. 2. If the care provider experiences a service indicator light on a stand-alone affected mattress, the light could be one of the error codes. Out of caution, please discontinue use of the product until a technical service representative can troubleshoot the service indicator error. If the service indicator light alarms while a patient is on the mattress, please move the patient to an alternative surface. If the service indicator light does not occur, the mattress is safe to continue use. 3. To correct the issue, Baxter will be sending personnel onsite to replace the technical box of all impacted Accella therapy air mattresses (and Spare Parts associated with Progressa Frames). A Baxter representative will contact your facility to schedule a time to complete the correction. | 2025-11-24 12:39
COPAN | SL SOLUTION, 1ML | 42 | - | SLsolution reagent includes dithiothreitol DTT which is oxidized when it gets in contact with air and or metals potentially losing its functionality in fluidification of respiratory specimens x D SLsolution reagent must show a colourless appearance On the other hand when the DTT reagent is deteriorated x D the SLsolution reagent shows a pink or yellow discoloration device failure x D its functionality in respiratory specimens fluidification cannot be guaranteed x D | • COPAN Italia recalling all lots of SLsolution™ devices reference number 0U019N • Please inspect your inventory and destroy any remaining stock of lot listed page 1 of this letter. Please make entry on the attached Acknowledgement Form regarding the number of kits destroyed. • Complete the attached Acknowledgement Form to confirm receipt of, and compliance to, this notice and return of Acknowledgement completion to BioMérieux. | 2025-11-24 16:10
3M Health Care | 3M Cavilon™ Durable Barrier Cream | - | - | There is no instruction for use in Arabic | The corrective action will be applied solely to the products currently available in each facility’s stock. We will provide the Arabic IFU (as per the current approved registration application) and ensure it is communicated to all affected facilities. Each facility will be required to sign the FSN once it is finalized by the manufacturer. | 2025-11-25 08:31
Abbott GmbH. & Co. KG | Alinity ci-series System Control Module | 327 | - | As per FSN attached | As per FSN attached | 2025-11-25 12:58
DiaMed GmbH | DiaClon Anti-P1 | 101 | - | The purpose of this letter is to inform you about a quality issue related to DiaClon Anti P instruction for use IFU x D Reason for the Action x D Following customer feedback and an internal investigation Bio Rad has identified an error on e IFU code number B v DiaClon Anti P Instruction For Use Result Interpretation Guidance Specifically the results for Well Reaction Grade that should be classified as Positive are incorrectly described as Non interpretable x D The discrepancy in the incorrect version affects result interpretation guidance within the IFU Specifically the incorrect version reclassified weak agglutination grades e g from positive to non interpretable x D x D Kindly review the attached | Action(s) to be taken by the Customer: Bio-Rad is recommending that customers affected by this notice take the following action: • Discard e-IFU code number B007212 v.02 2024-08-03 • Replace with the new e-IFU code number B007212 v.02.1 2025-11-04. This new version can be downloaded from the “Download.bio-rad.com” website as indicated on the box label: Please ensure that this notice is shared with all relevant personnel within your organization and any third-party recipients of the affected products. Please complete and return the attached response form as soon as possible so that we are assured you have received this important communication. Resolution by Bio-Rad: The corrected Instruction For Use (IFU) version has been uploaded with a new e-IFU code: B007212 v.02.1 2025-11-04 and is available for download on our website listed above. The changes are highlighted in grey for clarity. Kindly review the attached | 2025-11-26 13:37
Philips | Tempus Pro Monitor | 4 | - | The Tempus Pro Monitor User Operator Manual uses the term Invasive Pressure and the device software allows users to label pressure channels as ICP or BDR x D ICP and BDR measurements on the Tempus Pro Monitor have not been tested or validated for accuracy x D The device is only cleared for Invasive Blood Pressure IBP monitoring | Philips is currently investigating this and will send a follow-up communication regarding further actions. | 2025-11-27 13:45
Getinge | OTESUS OR table column, stationary | 1 | - | The new LEDs of the Optic and contact carrier module can cause interfering signals Wider beam angle between the OR Table Columns and the OR Table Tops These interfering signals can cause the following issues with the Otesus OR Table System x D Issue If the head down feet down trend function is moved to the end position and the button on the control device is kept pressed it may happen that the OR Table Top moves in the opposite trend direction until the button is released x D Issue OR Table Top cannot be moved to the zero position With Button and Leg Back function buttons Then there is no leg back movement possible x D If one of the errors occurs and no key on an input device is pressed for at least seconds the system goes into standby If the desired command is pressed again on an input device afterwards there is a possibility that it will be executed before one of the described errors occurs again Refer to the FSN for a detailed description of this issue x D | The device can be used in accordance with the instructions for use. • Please ensure this message is forwarded to any individuals that need notification within your organization or any organization where the affected devices have been transferred. • Please complete and send the Customer Response Form 2025-020 – OT-1394552. • Your Customer Service Contact will contact you to schedule the correction as swiftly as possible. • Getinge will perform a corrective action to solve the issue. | 2025-11-27 17:05
United Medical Industries Limited Co. L.L.C | UNIMED Procedure Pack Cath Lab Angiography Pack | 737 | - | We United Medical Industries Limited Co L L C received a FSN from our supplier MedNet GmbH Borkstrasse M nster Germany who is the original manufacturer of the Colour Coded Syringe ml LL Polycarbonate Blue stating that x D MedNet GmbH is issuing this urgent field safety notice after contamination with particles was identified in some syringes Particles may not be detectable before use due to the potential for the particle to adhere to the syringe seal therefore in the worst case scenario their use on patients could lead to an embolism | 1. Action on Quantity Distributed: • Send a Field Safety Notice / Field Safety Corrective Action to the end user. 2. Action on Quantity Available: • Return the quantities to United Medical Industries Limited Co. L.L.C. or • Quarantine the available quantity • Wait for the individually packed sterile “Colour Coded Syringe, 10ml, LL, Polycarbonate, Blue” to be supplied by United Medical Industries Limited Co. L.L.C • Piggy bag the supplied individually packed syringes with the procedure pack. • Supply instructions not to use the “Colour Coded Syringe, 10ml, LL, Polycarbonate, Blue” present inside the pack and use only the individually packed sterile “Colour Coded Syringe, 10ml, LL, Polycarbonate, Blue” as piggy bag, discard the unused syringes as per waste management policy of the institution. | 2025-12-03 09:04
United Medical Industries Limited Co. L.L.C | Unimed Procedure Pack Cardiothoracic Pack Cardiac Catheterization | 737 | - | We United Medical Industries Limited Co L L C received a FSN from our supplier MedNet GmbH Borkstrasse M nster Germany who is the original manufacturer of the Colour Coded Syringe ml LL Polycarbonate Blue stating that x D MedNet GmbH is issuing this urgent field safety notice after contamination with particles was identified in some syringes Particles may not be detectable before use due to the potential for the particle to adhere to the syringe seal therefore in the worst case scenario their use on patients could lead to an embolism | 1. Action on Quantity Distributed: • Send a Field Safety Notice / Field Safety Corrective Action to the end user. 2. Action on Quantity Available: • Return the quantities to United Medical Industries Limited Co. L.L.C. or • Quarantine the available quantity • Wait for the individually packed sterile “Colour Coded Syringe, 10ml, LL, Polycarbonate, Blue” to be supplied by United Medical Industries Limited Co. L.L.C • Piggy bag the supplied individually packed syringes with the procedure pack. • Supply instructions not to use the “Colour Coded Syringe, 10ml, LL, Polycarbonate, Blue” present inside the pack and use only the individually packed sterile “Colour Coded Syringe, 10ml, LL, Polycarbonate, Blue” as piggy bag, discard the unused syringes as per waste management policy of the institution. | 2025-12-03 09:04
United Medical Industries Limited Co. L.L.C | UNIMED Procedure Pack Interventional Radiology-IR Pack ANGIO SET OR PACK | 737 | - | We United Medical Industries Limited Co L L C received a FSN from our supplier MedNet GmbH Borkstrasse M nster Germany who is the original manufacturer of the Colour Coded Syringe ml LL Polycarbonate Blue stating that x D MedNet GmbH is issuing this urgent field safety notice after contamination with particles was identified in some syringes Particles may not be detectable before use due to the potential for the particle to adhere to the syringe seal therefore in the worst case scenario their use on patients could lead to an embolism | 1. Action on Quantity Distributed: • Send a Field Safety Notice / Field Safety Corrective Action to the end user. 2. Action on Quantity Available: • Return the quantities to United Medical Industries Limited Co. L.L.C. or • Quarantine the available quantity • Wait for the individually packed sterile “Colour Coded Syringe, 10ml, LL, Polycarbonate, Blue” to be supplied by United Medical Industries Limited Co. L.L.C • Piggy bag the supplied individually packed syringes with the procedure pack. • Supply instructions not to use the “Colour Coded Syringe, 10ml, LL, Polycarbonate, Blue” present inside the pack and use only the individually packed sterile “Colour Coded Syringe, 10ml, LL, Polycarbonate, Blue” as piggy bag, discard the unused syringes as per waste management policy of the institution. | 2025-12-03 09:04
Cook Incorporated | Flexor® Check-Flo® Introducer | 5 | - | Cook Medical identified that products supplied in the affected device lots may have been manufactured out of specification Impacted products may have been cut to the incorrect length not correctly trimmed and not inspected properly The issue was identified internally during observation of the manufacturing process where operators were found deviating from the established procedure | Identify Device(s), Quarantine Device(s), Return Device(s) to Cook Medical | 2025-12-03 16:03
test | test | - | - | test | test | 2025-12-04 13:16
Roche Diagnostics | cobas c 703 Analytical Unit | 7 | - | To prevent occurrences of premature deterioration to the Reaction Cells Wash tubing Roche Diagnostics Service x D Team will schedule additional service visits to replace tubing that is older than months x D A Corrective and Preventive Action CAPA investigation has been initiated and the root cause investigation x D continues Once the root cause analysis is complete appropriate corrective and preventive measures will be defined x D and communicated as needed | Customers can continue to use the affected device, but should perform a visual check of the rinse tubing immediately. ● In addition, please perform a visual check of the Reaction Cells Wash tubing when executing the monthly operator maintenance action “Replace Reaction Cells”. Please refer to the following attachment to FSN version1. ● In case you identify any damage on the tubing, please report the damage to your local Roche Service Representative ● No general recommendations with respect to the review of previous results can be given by using the cobas c703 analytical unit. Customers should follow their standard laboratory operating procedures. Any specific questions raised by the users should be addressed individually, considering all relevant clinical information. | 2025-12-07 15:57
Aesculap AG | MINOP TROCAR 150MM 4 WKING CHANNEL 6.0MM | - | - | Aesculap AG has identified a deviation in the length tolerances of the long shaft of the affected MINOP trocars When combined with the corresponding neuroendoscopes this deviation may restrict the surgical field of view When a MINOP neuroendoscope article number PE A is connected a minimal obstruction of the image on the monitor can be detected see Figure In particular when a MINOP neuroendoscope article number PE A is connected the image on the monitor is obscured in the upper quarter see Figure | Confirm receipt and understanding of this Safety Information. Take note of the potential failure mode. In the event that this issue occurs during surgery and a full 360° view is required, be aware that it may be necessary to rotate the trocar and the endoscope to compensate for the restricted view. | 2025-12-07 21:04
Micro Therapeutics, Inc. | Riptide™ Large Bore Aspiration Tubing | 53 | - | Medtronic received five from United Kingdom complaints between September and October regarding incorrect expiration date on the pouch label for Riptide Tubing lot This product lot was manufactured on October making the true expiration date of the product October which was reflected correctly on the carton label However the pouch label indicated July For one of the five complaints the product was used and the issue was detected after the procedure There was no consequence to the patient since the product was within the true use by date For the remaining complaints there was no patient involvement as the issue was detected prior to use Additional information was received from a Medtronic distribution center that a Riptide tubing unit displayed the incorrect expiration date of July on the carton label in this instance instead of the pouch label This unit belonged to lot all units of lot are in Medtronic control x D Based on supplier investigation till date a newly trained operator with limited experience in the labeling process was identified as a potential factor contributing to this labeling non conformance A horizontal assessment was performed for all available MAT lots which indicated that there are four other products lots which may have the same issue From these additional lots lot has been released to Medtronic and a portion of it is commercially distributed The remaining product lots are all fully contained at the distribution center x D | 1. Please note the true use-by dates for the impacted lots that are provided in Table 1 above. The products from the impacted lots are safe to use within the true use-by date and the customers may continue to use them until their true use-by date. 2. Complete the Customer Acknowledgment Form enclosed with this letter, acknowledging that you have received this information. 3. This notice should be distributed to all others in your organization who should be aware, or to any organization where the potentially affected devices have been transferred. Please maintain a copy of this notice in your records and retain it with the affected product. | 2025-12-08 13:55
Philips Medical Systems Nederland B.V. | AZURION 3 M15 | 13 | - | Issue Loss of X ray after Pedal Tap x D When the foot switch is pressed and released rapidly without generating X rays due to a short duration of contact the system may enter a state where X ray generation is not available When this occurs the system does not show any message to the user indicating that X ray generation is not possible x D A warm restart may resolve the situation in some cases A cold restart will always resolve the situation x D x D Issue Loss of X ray after Phase Fault x D Philips Azurion systems are designed to automatically recover from phase faults being an abnormal condition in which one or more phase voltages drop to near zero When a phase fault occurs the X ray generator software initiates an automatic recovery process to restore full system functionality during which imaging X ray functionality is not available for approximately seconds x D x D Due to a software issue during this recovery the communication with the system software may be lost resulting in the system not recognizing that the generator has recovered and is ready for operation In this situation the system shows the user message Generator is busy starting up no X ray possible If the phase fault is detected by the generator during acquisition the system also shows the message Run Aborted Tube Problem x D x D A cold restart of the system will always restore system functionality x D x D Issue Missing live X ray on the FlexVision monitor x D When viewing high resolution images i e resolution above x pixels visible on screen as Live K the FlexVision monitor may not display the X ray live window blank screen when one of the following specific sequence of actions is followed x D a Changing the live viewset up by maximizing another view on the FlexVision monitor or by explicitly removing the Live X ray application from the screen or hiding it only possible if X ray is disabled or by activating the screensaver or by starting comfort Themes followed by a warm restart of the system x D x D b Performing a warm restart to the system and then changing the FlexVision screen layout from applications to less than applications x D A warm restart will resolve the issue However the issue will re occur if the sequence of actions described in scenario a or b are again performed x D A cold restart will resolve this issue unless afterwards a warm restart is performed for any reason in combination with the sequence of actions described in scenario a or b x D x D x D Issue System keeps restarting x D x D When the Patient database in the Azurion system s Suite PC which stores metadata about patients studies procedures and series objects contains more than series objects system functionality becomes unavailable x D x D To address this the Azurion system initiates a recovery restart However because the number of series objects remains unchanged the issue persists causing the system to repeatedly perform recovery restarts No message is displayed to indicate this situation x D x D To resolve this issue a reinstallation of the Suite PC software is required x D x D Issue AMC triple drive x D Due to a firmware issue in the Advanced Motion Controller AMC Axis drive integrated electronic and mechanical component responsible for the precise movement and positioning of the Azurion s stand x D x D a Movements of the stand might be slow when this occurs the message Adjustment of frontal stand is required is displayed to the user When this occurs D scan functionality is not available The message Rotational scanning is unavailable Reselect the X ray protocol is displayed to the user To resolve this issue a calibration of the system by service is required x D x D b Motorized movements of the stand might become unavailable when this occurs the message Adjustment of frontal stand is required is shown to the user A cold restart of the system may temporarily resolve the issue in some cases x D Manual movements of the stand remain available Table movements are not impacted by this issue x D x D x D Issue C Partition of Suite PC running out of free space x D The C drive of the Azurion Suite PC may gradually lose available space due to the accumulation of printer spool files x D The Windows operating system creates a spool file for each print job and deletes it once the job is completed successfully x D If a print job fails these spool files are not removed and accumulate after each system re start when the Windows operating system retries the print job consuming disk space x D When the system re starts and the C drive is full motorized movements are disabled The following messages appear on the screen Some geometry movements are not available and The geometry is starting Do not change SID x D x D x D Issue System remains in continuous restart mode after the start up x D x D Due to a software issue the License Manager which verifies installed licenses and enables optional functionality in the Azurion system based on the available licenses expands the Windows registry database at every system start or restart x D Over time the registry may grow beyond the maximum size that the License Manager can load When this happens the License Manager cannot start and the system will enter a continuous reboot loop The Azurion system will not complete booting and the message X ray system is booting remains displayed on the screen x D x D x D Issue Misalignment of Marker Tool Overlay Applicable only x D to Azurion systems with the Marker Tool option x D x D The Marker Tool allows clinical users to draw freehand markers on a previously acquired X ray image to add guidance or indicate a region of interest Markers are displayed on all images within the series and on any images acquired after they are created x D When geometry movement exceeds predefined thresholds e g angulation or rotation differences greater than or a Beam IsoCenter to Patient deviation greater than mm markers are hidden x D x D When the geometry returns within the threshold markers reappear However they may not be restored to their exact original position and can deviate by several millimeters x D x D If the Marker Tool is used during a clinical procedure to indicate the correct placement of a device this deviation could lead to incorrect placement of the device x D x D Issue Longitudinal Position Error Applicable only to Azurion x D systems with Poly G frontal stand x D When requesting a movement of the frontal stand in longitudinal direction the Azurion system may detect a mismatch between the expected longitudinal set position and the measured actual position of the stand This mismatch is caused by a glitch in the longitudinal position value provided by the position sensing potentiometer x D x D When the mismatch is detected the motion power to the frontal stand is automatically cut off The message Some stand movements are not available is displayed to the user x D x D In this situation X ray imaging and table movements will continue to function x D x D This issue can be resolved with a cold restart of the system or by restarting the geometry x D | Philips will address the identified issues by implementing software updates in all affected systems. Philips expects to have software update 3.1.5 (FCO72200613), that resolves issues 1 - 7, released by Q1 2026 (subject to regulatory clearance). The Misalignment of Marker Tool Overlay (issue 8) and Longitudinal Position Error (issue 9) will not be resolved with software update 3.1.5. The solution for those issues is planned for Q4 2026 (subject to regulatory clearance) with software update 3.1.15 (FC072200684). Your local Philips representative will contact you to schedule visits to install the software updates once available. | 2025-12-08 14:14
Philips Medical Systems Nederland B.V. | AZURION 7 M12 | 13 | - | Issue Loss of X ray after Pedal Tap x D When the foot switch is pressed and released rapidly without generating X rays due to a short duration of contact the system may enter a state where X ray generation is not available When this occurs the system does not show any message to the user indicating that X ray generation is not possible x D A warm restart may resolve the situation in some cases A cold restart will always resolve the situation x D x D Issue Loss of X ray after Phase Fault x D Philips Azurion systems are designed to automatically recover from phase faults being an abnormal condition in which one or more phase voltages drop to near zero When a phase fault occurs the X ray generator software initiates an automatic recovery process to restore full system functionality during which imaging X ray functionality is not available for approximately seconds x D x D Due to a software issue during this recovery the communication with the system software may be lost resulting in the system not recognizing that the generator has recovered and is ready for operation In this situation the system shows the user message Generator is busy starting up no X ray possible If the phase fault is detected by the generator during acquisition the system also shows the message Run Aborted Tube Problem x D x D A cold restart of the system will always restore system functionality x D x D Issue Missing live X ray on the FlexVision monitor x D When viewing high resolution images i e resolution above x pixels visible on screen as Live K the FlexVision monitor may not display the X ray live window blank screen when one of the following specific sequence of actions is followed x D a Changing the live viewset up by maximizing another view on the FlexVision monitor or by explicitly removing the Live X ray application from the screen or hiding it only possible if X ray is disabled or by activating the screensaver or by starting comfort Themes followed by a warm restart of the system x D x D b Performing a warm restart to the system and then changing the FlexVision screen layout from applications to less than applications x D A warm restart will resolve the issue However the issue will re occur if the sequence of actions described in scenario a or b are again performed x D A cold restart will resolve this issue unless afterwards a warm restart is performed for any reason in combination with the sequence of actions described in scenario a or b x D x D x D Issue System keeps restarting x D x D When the Patient database in the Azurion system s Suite PC which stores metadata about patients studies procedures and series objects contains more than series objects system functionality becomes unavailable x D x D To address this the Azurion system initiates a recovery restart However because the number of series objects remains unchanged the issue persists causing the system to repeatedly perform recovery restarts No message is displayed to indicate this situation x D x D To resolve this issue a reinstallation of the Suite PC software is required x D x D Issue AMC triple drive x D Due to a firmware issue in the Advanced Motion Controller AMC Axis drive integrated electronic and mechanical component responsible for the precise movement and positioning of the Azurion s stand x D x D a Movements of the stand might be slow when this occurs the message Adjustment of frontal stand is required is displayed to the user When this occurs D scan functionality is not available The message Rotational scanning is unavailable Reselect the X ray protocol is displayed to the user To resolve this issue a calibration of the system by service is required x D x D b Motorized movements of the stand might become unavailable when this occurs the message Adjustment of frontal stand is required is shown to the user A cold restart of the system may temporarily resolve the issue in some cases x D Manual movements of the stand remain available Table movements are not impacted by this issue x D x D x D Issue C Partition of Suite PC running out of free space x D The C drive of the Azurion Suite PC may gradually lose available space due to the accumulation of printer spool files x D The Windows operating system creates a spool file for each print job and deletes it once the job is completed successfully x D If a print job fails these spool files are not removed and accumulate after each system re start when the Windows operating system retries the print job consuming disk space x D When the system re starts and the C drive is full motorized movements are disabled The following messages appear on the screen Some geometry movements are not available and The geometry is starting Do not change SID x D x D x D Issue System remains in continuous restart mode after the start up x D x D Due to a software issue the License Manager which verifies installed licenses and enables optional functionality in the Azurion system based on the available licenses expands the Windows registry database at every system start or restart x D Over time the registry may grow beyond the maximum size that the License Manager can load When this happens the License Manager cannot start and the system will enter a continuous reboot loop The Azurion system will not complete booting and the message X ray system is booting remains displayed on the screen x D x D x D Issue Misalignment of Marker Tool Overlay Applicable only x D to Azurion systems with the Marker Tool option x D x D The Marker Tool allows clinical users to draw freehand markers on a previously acquired X ray image to add guidance or indicate a region of interest Markers are displayed on all images within the series and on any images acquired after they are created x D When geometry movement exceeds predefined thresholds e g angulation or rotation differences greater than or a Beam IsoCenter to Patient deviation greater than mm markers are hidden x D x D When the geometry returns within the threshold markers reappear However they may not be restored to their exact original position and can deviate by several millimeters x D x D If the Marker Tool is used during a clinical procedure to indicate the correct placement of a device this deviation could lead to incorrect placement of the device x D x D Issue Longitudinal Position Error Applicable only to Azurion x D systems with Poly G frontal stand x D When requesting a movement of the frontal stand in longitudinal direction the Azurion system may detect a mismatch between the expected longitudinal set position and the measured actual position of the stand This mismatch is caused by a glitch in the longitudinal position value provided by the position sensing potentiometer x D x D When the mismatch is detected the motion power to the frontal stand is automatically cut off The message Some stand movements are not available is displayed to the user x D x D In this situation X ray imaging and table movements will continue to function x D x D This issue can be resolved with a cold restart of the system or by restarting the geometry x D | Philips will address the identified issues by implementing software updates in all affected systems. Philips expects to have software update 3.1.5 (FCO72200613), that resolves issues 1 - 7, released by Q1 2026 (subject to regulatory clearance). The Misalignment of Marker Tool Overlay (issue 8) and Longitudinal Position Error (issue 9) will not be resolved with software update 3.1.5. The solution for those issues is planned for Q4 2026 (subject to regulatory clearance) with software update 3.1.15 (FC072200684). Your local Philips representative will contact you to schedule visits to install the software updates once available. | 2025-12-08 14:14
Philips Medical Systems Nederland B.V. | AZURION 7 M20 | 13 | - | Issue Loss of X ray after Pedal Tap x D When the foot switch is pressed and released rapidly without generating X rays due to a short duration of contact the system may enter a state where X ray generation is not available When this occurs the system does not show any message to the user indicating that X ray generation is not possible x D A warm restart may resolve the situation in some cases A cold restart will always resolve the situation x D x D Issue Loss of X ray after Phase Fault x D Philips Azurion systems are designed to automatically recover from phase faults being an abnormal condition in which one or more phase voltages drop to near zero When a phase fault occurs the X ray generator software initiates an automatic recovery process to restore full system functionality during which imaging X ray functionality is not available for approximately seconds x D x D Due to a software issue during this recovery the communication with the system software may be lost resulting in the system not recognizing that the generator has recovered and is ready for operation In this situation the system shows the user message Generator is busy starting up no X ray possible If the phase fault is detected by the generator during acquisition the system also shows the message Run Aborted Tube Problem x D x D A cold restart of the system will always restore system functionality x D x D Issue Missing live X ray on the FlexVision monitor x D When viewing high resolution images i e resolution above x pixels visible on screen as Live K the FlexVision monitor may not display the X ray live window blank screen when one of the following specific sequence of actions is followed x D a Changing the live viewset up by maximizing another view on the FlexVision monitor or by explicitly removing the Live X ray application from the screen or hiding it only possible if X ray is disabled or by activating the screensaver or by starting comfort Themes followed by a warm restart of the system x D x D b Performing a warm restart to the system and then changing the FlexVision screen layout from applications to less than applications x D A warm restart will resolve the issue However the issue will re occur if the sequence of actions described in scenario a or b are again performed x D A cold restart will resolve this issue unless afterwards a warm restart is performed for any reason in combination with the sequence of actions described in scenario a or b x D x D x D Issue System keeps restarting x D x D When the Patient database in the Azurion system s Suite PC which stores metadata about patients studies procedures and series objects contains more than series objects system functionality becomes unavailable x D x D To address this the Azurion system initiates a recovery restart However because the number of series objects remains unchanged the issue persists causing the system to repeatedly perform recovery restarts No message is displayed to indicate this situation x D x D To resolve this issue a reinstallation of the Suite PC software is required x D x D Issue AMC triple drive x D Due to a firmware issue in the Advanced Motion Controller AMC Axis drive integrated electronic and mechanical component responsible for the precise movement and positioning of the Azurion s stand x D x D a Movements of the stand might be slow when this occurs the message Adjustment of frontal stand is required is displayed to the user When this occurs D scan functionality is not available The message Rotational scanning is unavailable Reselect the X ray protocol is displayed to the user To resolve this issue a calibration of the system by service is required x D x D b Motorized movements of the stand might become unavailable when this occurs the message Adjustment of frontal stand is required is shown to the user A cold restart of the system may temporarily resolve the issue in some cases x D Manual movements of the stand remain available Table movements are not impacted by this issue x D x D x D Issue C Partition of Suite PC running out of free space x D The C drive of the Azurion Suite PC may gradually lose available space due to the accumulation of printer spool files x D The Windows operating system creates a spool file for each print job and deletes it once the job is completed successfully x D If a print job fails these spool files are not removed and accumulate after each system re start when the Windows operating system retries the print job consuming disk space x D When the system re starts and the C drive is full motorized movements are disabled The following messages appear on the screen Some geometry movements are not available and The geometry is starting Do not change SID x D x D x D Issue System remains in continuous restart mode after the start up x D x D Due to a software issue the License Manager which verifies installed licenses and enables optional functionality in the Azurion system based on the available licenses expands the Windows registry database at every system start or restart x D Over time the registry may grow beyond the maximum size that the License Manager can load When this happens the License Manager cannot start and the system will enter a continuous reboot loop The Azurion system will not complete booting and the message X ray system is booting remains displayed on the screen x D x D x D Issue Misalignment of Marker Tool Overlay Applicable only x D to Azurion systems with the Marker Tool option x D x D The Marker Tool allows clinical users to draw freehand markers on a previously acquired X ray image to add guidance or indicate a region of interest Markers are displayed on all images within the series and on any images acquired after they are created x D When geometry movement exceeds predefined thresholds e g angulation or rotation differences greater than or a Beam IsoCenter to Patient deviation greater than mm markers are hidden x D x D When the geometry returns within the threshold markers reappear However they may not be restored to their exact original position and can deviate by several millimeters x D x D If the Marker Tool is used during a clinical procedure to indicate the correct placement of a device this deviation could lead to incorrect placement of the device x D x D Issue Longitudinal Position Error Applicable only to Azurion x D systems with Poly G frontal stand x D When requesting a movement of the frontal stand in longitudinal direction the Azurion system may detect a mismatch between the expected longitudinal set position and the measured actual position of the stand This mismatch is caused by a glitch in the longitudinal position value provided by the position sensing potentiometer x D x D When the mismatch is detected the motion power to the frontal stand is automatically cut off The message Some stand movements are not available is displayed to the user x D x D In this situation X ray imaging and table movements will continue to function x D x D This issue can be resolved with a cold restart of the system or by restarting the geometry x D | Philips will address the identified issues by implementing software updates in all affected systems. Philips expects to have software update 3.1.5 (FCO72200613), that resolves issues 1 - 7, released by Q1 2026 (subject to regulatory clearance). The Misalignment of Marker Tool Overlay (issue 8) and Longitudinal Position Error (issue 9) will not be resolved with software update 3.1.5. The solution for those issues is planned for Q4 2026 (subject to regulatory clearance) with software update 3.1.15 (FC072200684). Your local Philips representative will contact you to schedule visits to install the software updates once available. | 2025-12-08 14:14
Philips Medical Systems Nederland B.V. | AZURION 7 B12 | 13 | - | Issue Loss of X ray after Pedal Tap x D When the foot switch is pressed and released rapidly without generating X rays due to a short duration of contact the system may enter a state where X ray generation is not available When this occurs the system does not show any message to the user indicating that X ray generation is not possible x D A warm restart may resolve the situation in some cases A cold restart will always resolve the situation x D x D Issue Loss of X ray after Phase Fault x D Philips Azurion systems are designed to automatically recover from phase faults being an abnormal condition in which one or more phase voltages drop to near zero When a phase fault occurs the X ray generator software initiates an automatic recovery process to restore full system functionality during which imaging X ray functionality is not available for approximately seconds x D x D Due to a software issue during this recovery the communication with the system software may be lost resulting in the system not recognizing that the generator has recovered and is ready for operation In this situation the system shows the user message Generator is busy starting up no X ray possible If the phase fault is detected by the generator during acquisition the system also shows the message Run Aborted Tube Problem x D x D A cold restart of the system will always restore system functionality x D x D Issue Missing live X ray on the FlexVision monitor x D When viewing high resolution images i e resolution above x pixels visible on screen as Live K the FlexVision monitor may not display the X ray live window blank screen when one of the following specific sequence of actions is followed x D a Changing the live viewset up by maximizing another view on the FlexVision monitor or by explicitly removing the Live X ray application from the screen or hiding it only possible if X ray is disabled or by activating the screensaver or by starting comfort Themes followed by a warm restart of the system x D x D b Performing a warm restart to the system and then changing the FlexVision screen layout from applications to less than applications x D A warm restart will resolve the issue However the issue will re occur if the sequence of actions described in scenario a or b are again performed x D A cold restart will resolve this issue unless afterwards a warm restart is performed for any reason in combination with the sequence of actions described in scenario a or b x D x D x D Issue System keeps restarting x D x D When the Patient database in the Azurion system s Suite PC which stores metadata about patients studies procedures and series objects contains more than series objects system functionality becomes unavailable x D x D To address this the Azurion system initiates a recovery restart However because the number of series objects remains unchanged the issue persists causing the system to repeatedly perform recovery restarts No message is displayed to indicate this situation x D x D To resolve this issue a reinstallation of the Suite PC software is required x D x D Issue AMC triple drive x D Due to a firmware issue in the Advanced Motion Controller AMC Axis drive integrated electronic and mechanical component responsible for the precise movement and positioning of the Azurion s stand x D x D a Movements of the stand might be slow when this occurs the message Adjustment of frontal stand is required is displayed to the user When this occurs D scan functionality is not available The message Rotational scanning is unavailable Reselect the X ray protocol is displayed to the user To resolve this issue a calibration of the system by service is required x D x D b Motorized movements of the stand might become unavailable when this occurs the message Adjustment of frontal stand is required is shown to the user A cold restart of the system may temporarily resolve the issue in some cases x D Manual movements of the stand remain available Table movements are not impacted by this issue x D x D x D Issue C Partition of Suite PC running out of free space x D The C drive of the Azurion Suite PC may gradually lose available space due to the accumulation of printer spool files x D The Windows operating system creates a spool file for each print job and deletes it once the job is completed successfully x D If a print job fails these spool files are not removed and accumulate after each system re start when the Windows operating system retries the print job consuming disk space x D When the system re starts and the C drive is full motorized movements are disabled The following messages appear on the screen Some geometry movements are not available and The geometry is starting Do not change SID x D x D x D Issue System remains in continuous restart mode after the start up x D x D Due to a software issue the License Manager which verifies installed licenses and enables optional functionality in the Azurion system based on the available licenses expands the Windows registry database at every system start or restart x D Over time the registry may grow beyond the maximum size that the License Manager can load When this happens the License Manager cannot start and the system will enter a continuous reboot loop The Azurion system will not complete booting and the message X ray system is booting remains displayed on the screen x D x D x D Issue Misalignment of Marker Tool Overlay Applicable only x D to Azurion systems with the Marker Tool option x D x D The Marker Tool allows clinical users to draw freehand markers on a previously acquired X ray image to add guidance or indicate a region of interest Markers are displayed on all images within the series and on any images acquired after they are created x D When geometry movement exceeds predefined thresholds e g angulation or rotation differences greater than or a Beam IsoCenter to Patient deviation greater than mm markers are hidden x D x D When the geometry returns within the threshold markers reappear However they may not be restored to their exact original position and can deviate by several millimeters x D x D If the Marker Tool is used during a clinical procedure to indicate the correct placement of a device this deviation could lead to incorrect placement of the device x D x D Issue Longitudinal Position Error Applicable only to Azurion x D systems with Poly G frontal stand x D When requesting a movement of the frontal stand in longitudinal direction the Azurion system may detect a mismatch between the expected longitudinal set position and the measured actual position of the stand This mismatch is caused by a glitch in the longitudinal position value provided by the position sensing potentiometer x D x D When the mismatch is detected the motion power to the frontal stand is automatically cut off The message Some stand movements are not available is displayed to the user x D x D In this situation X ray imaging and table movements will continue to function x D x D This issue can be resolved with a cold restart of the system or by restarting the geometry x D | Philips will address the identified issues by implementing software updates in all affected systems. Philips expects to have software update 3.1.5 (FCO72200613), that resolves issues 1 - 7, released by Q1 2026 (subject to regulatory clearance). The Misalignment of Marker Tool Overlay (issue 8) and Longitudinal Position Error (issue 9) will not be resolved with software update 3.1.5. The solution for those issues is planned for Q4 2026 (subject to regulatory clearance) with software update 3.1.15 (FC072200684). Your local Philips representative will contact you to schedule visits to install the software updates once available. | 2025-12-08 14:14
Philips Medical Systems Nederland B.V. | AZURION 7 B20 | 13 | - | Issue Loss of X ray after Pedal Tap x D When the foot switch is pressed and released rapidly without generating X rays due to a short duration of contact the system may enter a state where X ray generation is not available When this occurs the system does not show any message to the user indicating that X ray generation is not possible x D A warm restart may resolve the situation in some cases A cold restart will always resolve the situation x D x D Issue Loss of X ray after Phase Fault x D Philips Azurion systems are designed to automatically recover from phase faults being an abnormal condition in which one or more phase voltages drop to near zero When a phase fault occurs the X ray generator software initiates an automatic recovery process to restore full system functionality during which imaging X ray functionality is not available for approximately seconds x D x D Due to a software issue during this recovery the communication with the system software may be lost resulting in the system not recognizing that the generator has recovered and is ready for operation In this situation the system shows the user message Generator is busy starting up no X ray possible If the phase fault is detected by the generator during acquisition the system also shows the message Run Aborted Tube Problem x D x D A cold restart of the system will always restore system functionality x D x D Issue Missing live X ray on the FlexVision monitor x D When viewing high resolution images i e resolution above x pixels visible on screen as Live K the FlexVision monitor may not display the X ray live window blank screen when one of the following specific sequence of actions is followed x D a Changing the live viewset up by maximizing another view on the FlexVision monitor or by explicitly removing the Live X ray application from the screen or hiding it only possible if X ray is disabled or by activating the screensaver or by starting comfort Themes followed by a warm restart of the system x D x D b Performing a warm restart to the system and then changing the FlexVision screen layout from applications to less than applications x D A warm restart will resolve the issue However the issue will re occur if the sequence of actions described in scenario a or b are again performed x D A cold restart will resolve this issue unless afterwards a warm restart is performed for any reason in combination with the sequence of actions described in scenario a or b x D x D x D Issue System keeps restarting x D x D When the Patient database in the Azurion system s Suite PC which stores metadata about patients studies procedures and series objects contains more than series objects system functionality becomes unavailable x D x D To address this the Azurion system initiates a recovery restart However because the number of series objects remains unchanged the issue persists causing the system to repeatedly perform recovery restarts No message is displayed to indicate this situation x D x D To resolve this issue a reinstallation of the Suite PC software is required x D x D Issue AMC triple drive x D Due to a firmware issue in the Advanced Motion Controller AMC Axis drive integrated electronic and mechanical component responsible for the precise movement and positioning of the Azurion s stand x D x D a Movements of the stand might be slow when this occurs the message Adjustment of frontal stand is required is displayed to the user When this occurs D scan functionality is not available The message Rotational scanning is unavailable Reselect the X ray protocol is displayed to the user To resolve this issue a calibration of the system by service is required x D x D b Motorized movements of the stand might become unavailable when this occurs the message Adjustment of frontal stand is required is shown to the user A cold restart of the system may temporarily resolve the issue in some cases x D Manual movements of the stand remain available Table movements are not impacted by this issue x D x D x D Issue C Partition of Suite PC running out of free space x D The C drive of the Azurion Suite PC may gradually lose available space due to the accumulation of printer spool files x D The Windows operating system creates a spool file for each print job and deletes it once the job is completed successfully x D If a print job fails these spool files are not removed and accumulate after each system re start when the Windows operating system retries the print job consuming disk space x D When the system re starts and the C drive is full motorized movements are disabled The following messages appear on the screen Some geometry movements are not available and The geometry is starting Do not change SID x D x D x D Issue System remains in continuous restart mode after the start up x D x D Due to a software issue the License Manager which verifies installed licenses and enables optional functionality in the Azurion system based on the available licenses expands the Windows registry database at every system start or restart x D Over time the registry may grow beyond the maximum size that the License Manager can load When this happens the License Manager cannot start and the system will enter a continuous reboot loop The Azurion system will not complete booting and the message X ray system is booting remains displayed on the screen x D x D x D Issue Misalignment of Marker Tool Overlay Applicable only x D to Azurion systems with the Marker Tool option x D x D The Marker Tool allows clinical users to draw freehand markers on a previously acquired X ray image to add guidance or indicate a region of interest Markers are displayed on all images within the series and on any images acquired after they are created x D When geometry movement exceeds predefined thresholds e g angulation or rotation differences greater than or a Beam IsoCenter to Patient deviation greater than mm markers are hidden x D x D When the geometry returns within the threshold markers reappear However they may not be restored to their exact original position and can deviate by several millimeters x D x D If the Marker Tool is used during a clinical procedure to indicate the correct placement of a device this deviation could lead to incorrect placement of the device x D x D Issue Longitudinal Position Error Applicable only to Azurion x D systems with Poly G frontal stand x D When requesting a movement of the frontal stand in longitudinal direction the Azurion system may detect a mismatch between the expected longitudinal set position and the measured actual position of the stand This mismatch is caused by a glitch in the longitudinal position value provided by the position sensing potentiometer x D x D When the mismatch is detected the motion power to the frontal stand is automatically cut off The message Some stand movements are not available is displayed to the user x D x D In this situation X ray imaging and table movements will continue to function x D x D This issue can be resolved with a cold restart of the system or by restarting the geometry x D | Philips will address the identified issues by implementing software updates in all affected systems. Philips expects to have software update 3.1.5 (FCO72200613), that resolves issues 1 - 7, released by Q1 2026 (subject to regulatory clearance). The Misalignment of Marker Tool Overlay (issue 8) and Longitudinal Position Error (issue 9) will not be resolved with software update 3.1.5. The solution for those issues is planned for Q4 2026 (subject to regulatory clearance) with software update 3.1.15 (FC072200684). Your local Philips representative will contact you to schedule visits to install the software updates once available. | 2025-12-08 14:14
GS Elektromedizinische Geräte G. Stemple GmbH | Defib Unit Corpuls3 | 270 | - | The manufacturer has received reports in which the Display Module of the corpuls system has unexpectedly shut down and initiated an automatic reboot during patient care This results in a temporary interruption of certain device functions typically lasting between to seconds with an average duration of seconds x D x D This limited reboot is triggered by the device s internal software safety mechanisms as part of its fault management strategy The function is designed to preserve core functionality and maintain operational safety Importantly this does not represent a full system reboot once the Display Module restarts the device continues to operate as intended without the need for user intervention | As part of our post-market surveillance and risk control process, we are implementing a product correction by releasing an enhanced software version, REL 4.4.2, which addresses the root cause of the self-initiated safety reboot mechanism. | 2025-12-09 11:48
GS Elektromedizinische Geräte G. Stemple GmbH | Defib Corpuls 3 Touch SLIM | 270 | - | The manufacturer has received reports in which the Display Module of the corpuls system has unexpectedly shut down and initiated an automatic reboot during patient care This results in a temporary interruption of certain device functions typically lasting between to seconds with an average duration of seconds x D x D This limited reboot is triggered by the device s internal software safety mechanisms as part of its fault management strategy The function is designed to preserve core functionality and maintain operational safety Importantly this does not represent a full system reboot once the Display Module restarts the device continues to operate as intended without the need for user intervention | As part of our post-market surveillance and risk control process, we are implementing a product correction by releasing an enhanced software version, REL 4.4.2, which addresses the root cause of the self-initiated safety reboot mechanism. | 2025-12-09 11:48
Randox Laboratories Ltd | SEROLOGY ToRCH IgM POSITIVE CONTROL | 102 | - | Randox Laboratories can confirm that the analyte Herpes Simplex Virus HSV Type IgM is Negative in Serology ToRCH IgM positive control catalogue number SR lot SR | Transmission of Customer Notification: Send a copy of the notification to all affected customers and to those who need to be aware within the organisation. Complete and return the response form (12187-QA) to technical.services@randox.com within five working days. Please accept our apologies for any inconvenience caused. Thank you for your patience and understanding. If you have any questions or concerns, please contact Randox Technical Services. | 2025-12-09 12:35
Siemens Healthcare GmbH | Sensis Vibe Hemo | - | - | The issue was identified during regular field observation The root cause is within the Solid State Drive x D SSD controller of a certain SSD model which was used within the HP Flex Pro C PC | The affected SSD model will be replaced with a different model, which has not shown the Blue Screen error. | 2025-12-10 11:08
Siemens Healthcare GmbH | SENSIS Vibe Combo | - | - | The issue was identified during regular field observation The root cause is within the Solid State Drive x D SSD controller of a certain SSD model which was used within the HP Flex Pro C PC | The affected SSD model will be replaced with a different model, which has not shown the Blue Screen error. | 2025-12-10 11:08
Stago | Reagent | 32 | - | some strips of Reagent from Asserachrom HPIA lot had a coloring defect x D | We will proceed with the disposal of the kits(Destroy). | 2025-12-10 15:41
Medtronic Inc. | Vanta™/Sequentia™ LT Clinician Programmer Application | 15 | - | The Vanta A Clinician Programmer CP App v and v will display a System Error if the user selects the Start Usage button twice when interrogating a new implantable neurostimulator INS This error prompts the user to restart the CP app and contact Medtronic for assistance Upon subsequent interrogation of the INS the Vanta A CP App will display a Validation Error prompting the user to resolve the error in order to proceed with the INS interrogation If the Validation Error is not resolved the user is unable to complete the first attempt at interrogating the INS This issue does not occur with other Medtronic neurostimulators x D Based on the investigation it is found that the issue was caused when the user selects the Start Usage button twice when interrogating a new INS The Validation Error that displays prevents the user from completing first attempt at interrogating the INS if not resolved This issue presents at the point of use Updating the Vanta A CP App to or later addresses the issue as the telemetry dialog screen will transition directly into the next screen of the workflow which prevents the Start Usage button from being shown or selected a second time x D | Complete the enclosed Customer Acknowledgement Form | 2025-12-10 16:49
Zimmer Surgical Inc. | ZIMMER DERMATOME AN Dermatome AN Handpiece | 8 | - | Worldwide complaints confirmed that affected Air Dermatome devices may have a misaligned thickness control bar from manufacture resulting in cutting beyond the intended depth | Aligned with global FSCA ZFA-2025-00141. Saudi Arabia impacted (8 hospitals/8 units). Phased customer notification, quarantine of affected SNs, and phased return/repair to be initiated next week after EMEA FSN issuance. | 2025-12-11 02:39
IMIRIS imaging Inc. | Head Fixation Device | 1 | - | The torque screw component may crack or separate radially along the exterior housing x D Defect may occur during or immediately before clinical use x D Identified cause Design limitations in the current torque screw design x D A cracked damaged torque screw may lead to x D Inadequate head stabilization Compromised surgical precision Procedural delays Need for alternative fixation methods Potential patient injury | IMRIS will remove and replace the existing torque screw component with an improved design. Replacement program begins January 2026; IMRIS service team will schedule the correction. Customers must confirm receipt and return the acknowledgment form. Until replacement is completed, users must follow safety mitigations: -Ensure patient’s head is properly supported during pin penetration. -Discontinue use immediately if a crack is observed and notify IMRIS. -Inspect the torque screw before each use for wear, cracks, or damage. -Handle the component carefully during assembly, cleaning, and storage. -Stop using the part if dropped or impacted; notify IMRIS. -Follow instructions in the Operator Manual for proper assembly and use. -Share the recall notice with all staff handling the device and report issues. | 2025-12-11 12:45
HUMAN Gesellschaft fȕr Biochemica & Diagnostica mbH | Autocal | 245 | - | we noted that the concentration for CHOLESTEROL liquicolor in AUTOCAL REF LOT has increased over time x D These discrepancies occur when the above mentioned Cholesterol reagents are calibrated either with AUTOCAL LOT or with the kit standard included in the CHOLESTEROL liquicolor kits | As a measure the Cholesterol value in AUTOCAL LOT 0019 was revised as indicated below in the attached table also As a consequence, the Cholesterol values for corresponding kit standards were revised as well, as indicated below | 2025-12-11 19:30
HUMAN Gesellschaft fȕr Biochemica & Diagnostica mbH | Cholesterol Liquicolor | 245 | - | we noted that the concentration for CHOLESTEROL liquicolor in AUTOCAL REF LOT has increased over time x D These discrepancies occur when the above mentioned Cholesterol reagents are calibrated either with AUTOCAL LOT or with the kit standard included in the CHOLESTEROL liquicolor kits | As a measure the Cholesterol value in AUTOCAL LOT 0019 was revised as indicated below in the attached table also As a consequence, the Cholesterol values for corresponding kit standards were revised as well, as indicated below | 2025-12-11 19:30
Sysmex Corporation | Automated Tube Sorter | 9 | - | under specific conditions the lift s electromagnetic x D holding force may disengage unexpectedly allowing unintended downward movement of the lift x D This event may occur during the procedure described in the TS Instructions for Use IFU x D Chapter Removing a rack or foreign object from the carry in line | Actions to be taken by the customer: To prevent injury, never place your hands around the lift section at any time. In addition, please follow these instructions when handling rack jams or similar errors: ✓ Pay special attention not to insert your hands around the lift section during any intervention. ✓ If a rack, tube, or any object becomes lodged around the lift section, tum off the power. After confirming that the lift section has moved to the bottom dead point, carefully remove the rack and blood collection tubes. ✓ If the rack or blood collection tubes are jammed in the lift section and cannot be removed with one hand, contact Sysmex and request an FSR for assistance. Please ensure this notice is communicated to all relevant personnel and retained in a visible location until the software update is completed. Actions to be taken by Sysmex: Sysmex will implement a software update to improve lift position control to prevent unintended sudden drops of the lift section by ensuring that, after an error occurs, the lift automatically returns to the home position. Our service team will contact you to schedule the installation of software version 00-18 or later. | 2025-12-14 13:12
Medtronic Inc. | Vanta™ with AdaptiveStim™ Technology | 13 | - | Medtronic has determined that Vanta may be susceptible to not communicating with the clinician programmer and or the patient programmer system HH Handset and TM Communicator Although the INS can continue to deliver therapy to the patient within programmed parameters the inability to communicate can put the patient at risk In some instances if there is no communication with the INS during the implant procedure troubleshooting may delay the surgical procedure and may require a replacement INS In other instances if there is no communication with the INS post procedure it might not be possible to adjust therapy and may cause the patient to experience inadequate therapy i e return of underlying disease symptoms In rare instances if communication cannot be restored to the INS via a recommended mitigation an unanticipated surgical intervention may be needed to explant and replace the INS | • Please review your facility’s inventory of Vanta INS devices and take the following actions: • Identify Vanta INS devices manufactured with the old firmware by referencing the PIN number on the device packaging. PIN number 977006002H indicates the device has the old firmware. • Return unused affected product per the return instructions. • Complete the enclosed Customer Acknowledgement Form. When complete please return the form to Medtronic • This notice should be passed on to all those who need to be aware within your organization or to any organization where the potentially affected devices have been transferred. • Please maintain a copy of this notice in your records. | 2025-12-15 14:10
Philips Medical Systems Nederland B.V. | AZURION 7 M12 | 8 | - | Issue System remains in continuous restart mode after the start up x D Due to a software issue the License Manager which verifies installed licenses and enables optional functionality in the Azurion system based on the available licenses expands the Windows registry database at every system start or restart x D Over time the registry may grow beyond the maximum size that the License Manager can load When this happens the License Manager cannot start and the system will enter a continuous reboot loop The Azurion system will not complete booting and the message X ray system is booting remains displayed on the screen x D x D Issue Longitudinal Position Error Applicable only to Azurion systems with Poly G frontal stand x D When requesting a movement of the frontal stand in longitudinal direction the Azurion system may detect a mismatch between the expected longitudinal set position and the measured actual position of the stand This mismatch is caused by a glitch in the longitudinal position value provided by the position sensing potentiometer x D When the mismatch is detected the motion power to the frontal stand is automatically cut off The message Some stand movements are not available is displayed to the user x D x D In this situation X ray imaging and table movements will continue to function x D This issue can be resolved with a cold restart of the system or by restarting the geometry | Philips will address the identified issues by implementing software updates in all affected systems. Philips expects to have software update 3.1.5 (FCO72200635, FC072200671), that resolves issues 1, released by Q1 2026 (subject to regulatory clearance). The Longitudinal Position Error (issue 2) will not be resolved with software update 3.1.5. The solution for this issue is planned for Q4 2026 (subject to regulatory clearance) with software update 3.1.15 (FCO72200684). | 2025-12-15 15:31
Philips Medical Systems Nederland B.V. | AZURION 7 M20 | 8 | - | Issue System remains in continuous restart mode after the start up x D Due to a software issue the License Manager which verifies installed licenses and enables optional functionality in the Azurion system based on the available licenses expands the Windows registry database at every system start or restart x D Over time the registry may grow beyond the maximum size that the License Manager can load When this happens the License Manager cannot start and the system will enter a continuous reboot loop The Azurion system will not complete booting and the message X ray system is booting remains displayed on the screen x D x D Issue Longitudinal Position Error Applicable only to Azurion systems with Poly G frontal stand x D When requesting a movement of the frontal stand in longitudinal direction the Azurion system may detect a mismatch between the expected longitudinal set position and the measured actual position of the stand This mismatch is caused by a glitch in the longitudinal position value provided by the position sensing potentiometer x D When the mismatch is detected the motion power to the frontal stand is automatically cut off The message Some stand movements are not available is displayed to the user x D x D In this situation X ray imaging and table movements will continue to function x D This issue can be resolved with a cold restart of the system or by restarting the geometry | Philips will address the identified issues by implementing software updates in all affected systems. Philips expects to have software update 3.1.5 (FCO72200635, FC072200671), that resolves issues 1, released by Q1 2026 (subject to regulatory clearance). The Longitudinal Position Error (issue 2) will not be resolved with software update 3.1.5. The solution for this issue is planned for Q4 2026 (subject to regulatory clearance) with software update 3.1.15 (FCO72200684). | 2025-12-15 15:31
Philips Medical Systems Nederland B.V. | AZURION 7 B12 | 8 | - | Issue System remains in continuous restart mode after the start up x D Due to a software issue the License Manager which verifies installed licenses and enables optional functionality in the Azurion system based on the available licenses expands the Windows registry database at every system start or restart x D Over time the registry may grow beyond the maximum size that the License Manager can load When this happens the License Manager cannot start and the system will enter a continuous reboot loop The Azurion system will not complete booting and the message X ray system is booting remains displayed on the screen x D x D Issue Longitudinal Position Error Applicable only to Azurion systems with Poly G frontal stand x D When requesting a movement of the frontal stand in longitudinal direction the Azurion system may detect a mismatch between the expected longitudinal set position and the measured actual position of the stand This mismatch is caused by a glitch in the longitudinal position value provided by the position sensing potentiometer x D When the mismatch is detected the motion power to the frontal stand is automatically cut off The message Some stand movements are not available is displayed to the user x D x D In this situation X ray imaging and table movements will continue to function x D This issue can be resolved with a cold restart of the system or by restarting the geometry | Philips will address the identified issues by implementing software updates in all affected systems. Philips expects to have software update 3.1.5 (FCO72200635, FC072200671), that resolves issues 1, released by Q1 2026 (subject to regulatory clearance). The Longitudinal Position Error (issue 2) will not be resolved with software update 3.1.5. The solution for this issue is planned for Q4 2026 (subject to regulatory clearance) with software update 3.1.15 (FCO72200684). | 2025-12-15 15:31
Philips Medical Systems Nederland B.V. | AZURION 3 M12 | 86 | - | Issue System keeps restarting x D Issue AMC triple drive x D Issue C Partition of Suite PC running out of free space x D Issue System remains in continuous restart mode after the start up x D Issue Misalignment of Marker Tool Overlay Applicable only to Azurion systems with the Marker Tool option x D Issue Longitudinal Position Error Applicable only to Azurion systems with Poly G frontal stand x D | Philips will address the identified issues by implementing a software update (R2.1.15 or R2.2.15) in all affected systems (FCO72200660 to 2.2.15, FCO72200672 to 2.1.15). Philips expects to have this software update released by Q4 2026 (subject to regulatory clearance). | 2025-12-15 15:56
Philips Medical Systems Nederland B.V. | Azurion 3 M15, Azurion 5 M20,Azurion 7 B12,Azurion 7 B20, Azurion 7 M12, Azurion 7 M20 | 86 | - | Issue System keeps restarting x D Issue AMC triple drive x D Issue C Partition of Suite PC running out of free space x D Issue System remains in continuous restart mode after the start up x D Issue Misalignment of Marker Tool Overlay Applicable only to Azurion systems with the Marker Tool option x D Issue Longitudinal Position Error Applicable only to Azurion systems with Poly G frontal stand x D | Philips will address the identified issues by implementing a software update (R2.1.15 or R2.2.15) in all affected systems (FCO72200660 to 2.2.15, FCO72200672 to 2.1.15). Philips expects to have this software update released by Q4 2026 (subject to regulatory clearance). | 2025-12-15 15:56
Elekta Solutions AB | Leksell GammaPlan | 1 | - | Defining a new stereotactic reference in an examination for a frame treatment while failing to first change an already defined obsolete stereotactic reference representing a non current frame fitting to a pre plan reference may result in a treatment plan that irradiates the wrong patient location if the resulting inconsistent alignment between the stereotactic references is not detected by the user | Elekta will send an Important Field Safety Notice 100-01-102-019 to all affected customers 17 DEC-25. The notice informs users of the specific product and version numbers affected by the issue, and any work around that can be used to avoid the issue. Sites affected will be Customers Leksell GammaPlan® versions 11.1, 11.3 and 11.4 in use with Leksell Gamma Knife® Icon and Leksell Gamma Knife® Elekta Esprit. A copy of the Important Field Safety Notice is to be kept with the most current labelling and all personnel working with the product should be made aware of the content of the letter. The customers are instructed to complete and return the Important Field Safety Notice Acknowledgement form to their local Elekta representative as soon as possible but no later than 30 days from receipt. | 2025-12-15 16:04
LivaNova USA, Inc | VNS Therapy™ PerenniaFLEX™ | 1 | - | The root cause of the high impedance is currently unknown | No corrective action is required as no new cause or new harm has been established for patient or user. | 2025-12-16 11:17
Intuitive Surgical | da Vinci® Single Port System – Access Port Kits | 12 | - | Cracks may develop in the PETG sterile tray tub due to a supplier manufacturing defect Cracks may be difficult to visually detect A cracked tray tub compromises sterile barrier integrity x D Use of affected product may result in x D Loss of sterility Potential patient contamination Increased risk of infection No complaints adverse events or serious injuries have been reported to date x D | -Immediate stop-ship implemented for all affected products. -Customers instructed to: -Identify affected inventory -Quarantine affected units -Cease use immediately -Return affected products for credit -Complete and return the Customer Acknowledgement Form | 2025-12-17 15:41
Philips Medical Systems Nederland B.V. | Allura Xper FD20 | 5 | - | Certain Philips Allura Xper and Allura CV systems use a Laird cooling unit to control the flat detector s temperature This cooling unit is housed in a system cabinet in the technical room x D x D Philips has identified that a limited number of Allura Xper and Allura CV systems may lack a drip tray beneath the cooling unit This component is intended to prevent potential leaking coolant liquid from contacting electrical components If it is not installed there is a potential for coolant liquid to contact electrical components This could lead to electrical short circuits which may trigger system fuses and result in system shutdown Electrical short circuits could cause elevated temperatures in insulating materials which may result in smoke formation In severe cases deterioration of insulation could produce sparks and combined with flammable coolant liquid and high temperatures lead to fire x D | Philips will inspect the affected systems to confirm that the drip tray is installed and check if there are leakages. Your local Philips representative will contact you to schedule a visit for a Field Service Engineer to perform this activity. Philips expects to start this activity in February 2026. | 2025-12-17 16:48
Philips Medical Systems Nederland B.V. | Allura Xper FD10/10 | 5 | - | Certain Philips Allura Xper and Allura CV systems use a Laird cooling unit to control the flat detector s temperature This cooling unit is housed in a system cabinet in the technical room x D x D Philips has identified that a limited number of Allura Xper and Allura CV systems may lack a drip tray beneath the cooling unit This component is intended to prevent potential leaking coolant liquid from contacting electrical components If it is not installed there is a potential for coolant liquid to contact electrical components This could lead to electrical short circuits which may trigger system fuses and result in system shutdown Electrical short circuits could cause elevated temperatures in insulating materials which may result in smoke formation In severe cases deterioration of insulation could produce sparks and combined with flammable coolant liquid and high temperatures lead to fire x D | Philips will inspect the affected systems to confirm that the drip tray is installed and check if there are leakages. Your local Philips representative will contact you to schedule a visit for a Field Service Engineer to perform this activity. Philips expects to start this activity in February 2026. | 2025-12-17 16:48
Merit Medical Systems Inc. | BasixCOMPAK | 0 | - | In rare instances the handle may detach from the syringe during use due to an insufficient amount of adhesive on the handle The lack of adhesive is not detectable prior to use In a single instance the detachment occurred concurrently with an adverse event | 1. Holding the cap and barrel together during use may effectively prevent detachment. 2. Visual indicators such as pressure drops on the gauge can alert users to potential failures. Refer to the product Instructions for Use. 3. Simulated use testing showed that handle separation results in flow restoration. 4. Should the handle detach during use and the balloon not completely deflate, pull the handle out of the inflation syringe, disconnect the inflation syringe from the catheter, and use an alternate syringe to generate vacuum and aspirate residual contrast. | 2025-12-18 16:08
Philips Goldway (Shenzhen) Industrial Inc. | Efficia DFM100 | 805 | - | The Efficia DFM Defibrillator Monitor with software revision installed may fail to complete its boot up sequence This could occur either during manual boot up if the knob is turned to AED Manual Pacer or as part of a self test if the device is turned off and connected to any power source Instead of completing the boot up sequence the device may repeatedly display the splash screen and will not be available for clinical use This will be indicated to the user by displaying a solid red X in the Ready For Use RFU indicaton | Philips is providing this Urgent Field Safety Notice to inform affected customers. The software solution has been released (SW 2.00.46 or higher), and Philips will correct devices that have SW 2.00.33 installed via FCO86100240 | 2025-12-19 13:59
ICU Medical | Other | 311199 | - | ICU Medical has identified lot specific issues with the silicone seal on the Tego device The identified silicone seal defects include silicone seal doming which occurs because of a loose silicone seal which may bulge at the top surface or separate from the Tego body and may potentially result in fluid leaks and silicone seal tearing which can potentially result in a collapsed silicone seal which may lead to occluded fluid flow or fluid leaks | When using the device, all instructions, including warnings and cautions contained in the Instructions for Use must be followed with heightened awareness. Please complete the following actions below. 1. Check all inventory locations within your institution for the affected Tego products listed in Table 1 and discontinue use. Destroy all affected products following your institution’s process for destruction. If destroying is not immediately possible at your facility, then the product should be quarantined until disposal. 2. Share this notification with all potential users of the device, to ensure they are aware of this notification. If the devices are used at another location, please ensure this communication is delivered there. | 2025-12-21 17:48
Philips Medical Systems | Patient Information Center iX | 1 | - | During internal investigation Philips identified that the recent PIC iX software SW patch modifies Mobile Event Notification filter settings without providing any indication to the user when upgrading the system When patching a PIC iX X i e system that uses the default event filters for mobile notifications upgrading to overrides existing default filters and applies the new default filters for version x D The overridden filters can lead to the situation where alarms which were previously configured to distribute to mobile devices are no longer sent and no notification is provided by the system x D While all default event filters are overridden when patching to any customized configurations applied to high level filter categories subcategories or individual filters are maintained and migrated successfully x D | Philips representative will contact you to obtain PIC iX 4.5.1 SW patch for PIC iX systems running software version 4.5.0. | 2025-12-22 13:38
GE Healthcare | Centricity RISi | 55 | - | GE HealthCare has become aware of a cybersecurity vulnerability affecting Centricity RIS i configured with Instant Messenger When Instant Messenger is installed and running a malicious actor could exploit this vulnerability to perform unauthenticated Remote Code Execution in Instant Messenger which could potentially impact system availability and or lead to ability to manipulate data Instant Messenger is an optional module which is not installed by default The vulnerability is not present in RIS i systems without Instant Messenger or where Instant Messenger has been disabled x D x D There have been no injuries or unauthorized access to patient data reported to GE HealthCare as a result of this issue x D | Customers with RIS-i Versions 4.2.24 thru 7.0 SP0 will get a Customer Letter, Instant Messenger will be de-installed, and Instant Messenger setup will be removed. Customers with RIS-i Versions 7.0 SP1 thru SP6 will get a Customer Letter and software upgrade. | 2025-12-23 01:13
Boston Scientific Corporation | Hot Axios Stent and Electrocautery Enhanced Delivery System | 124 | - | Boston Scientific is initiating a removal of certain AXIOS Stent and Electrocautery Enhanced Delivery Systems mm mm mm mm and mm mm detailed in Attachment This action is being taken due to increased reports of stent deployment and expansion issues with these configurations This issue only occurs at the time of delivery of the stent and is expected to be noticed by the physician x D Successfully implanted stents are not affected by this issue | removal of certain AXIOS™ Stent and Electrocautery Enhanced Delivery Systems (6 mm × 8 mm, 8 mm × 8 mm, and 20 mm × 10 mm) | 2025-12-23 16:56
Sysmex Corporation | XN-20, XN-10 | - | - | Sysmex has been made aware that in the above mentioned software versions when measuring XN CAL in Blood Bank BB mode the WBC value was approximately higher than expected Furthermore if calibration was performed using that XN CAL subsequent measurements of residual leukocytes in actual blood products or samples would be erroneously underestimated by about Our internal investigation revealed that an incorrect calculation had been incorporated into these software versions x D The false low result only occurs under specific conditions as below x D Software Version x D BB mode is activated x D Calibration performed in BB mode with one of the affected aforementioned SW versions i e if calibration of BB mode has been performed with any other software version this issue does not affect your device x D | Actions to be taken by the customer: Please check the version of the BB mode currently in use. If it is version 22.20-00 or version 22.21-00, you may continue using it without any issues. If you are using an affected version, please contact your Sysmex representative. Please confirm with your Sysmex representative whether calibration of the device has been performed in BB mode using one of the affected aforementioned software versions. If calibration was performed in BB mode, please either refrain from reporting WBC count from BB mode until the Sysmex representative completes the corrective action or continue while taking into account that the true residual white blood cell count may be approximately 15% higher than the reported value. Sysmex advises you to consult your facility's physician and/or pathologist to determine any clinical implications (including retrospective review and/or re-testing). Please ensure this notice is communicated to all relevant personnel and retained in a visible location until the software update is completed. Actions to be taken by Sysmex: The Sysmex representative will visit your site to perform immediate action, either by adjusting the relevant calibration factor to 15% higher or by changing the version to 22.20-00 or 22.21-00. Sysmex will develop and will provide you the permanent corrective action as soon as possible. Communication of this Field Safety Notice: Distribute this FSN to all responsible persons within your organization and return the Acknowledgement of Receipt (AoR) with your signature to your authorized local Sysmex representative. | 2025-12-23 17:58
The National Medical Products Co Ltd | Damap Prep Pad Povidone-Iodine swabstick | 202400 | - | The swabstick color is light instead of dark brown | Recall | 2025-12-24 11:57
Codman and Shurtleff, Inc. | CODMAN® MICROSENSOR® Basic Kit | 66 | - | Integra LifeSciences is voluntarily issuing this Field Safety Notice for the recall of CODMAN MICROSENSOR Basic kit and CERELINK ICP sensor Basic kit listed in Table x D During an investigation Integra LifeSciences identified lots that were released with potential corrosion stains on the surface of the gauge Tuohy needle that is included within the sensor kits ref and x D To date there has been one adverse event of infection reported worldwide x D Integra LifeSciences will initiate a voluntary recall of the impacted lots This recall applies only to products within the specific lot numbers listed in Appendix of this communication | Actions to be taken by Distributors: 1. Please review and understand the information provided in this letter. 2. Determine if the product you have is subject to the recall: a. Identify the impacted reference and lot number. b. See Appendix 2 below for a sample of product label for where to locate the reference and lot number. The lot numbers are 7 digits long (only numbers) c. Open the excel file (appendix 3), use the find function Ctrl+F or use the dropdown arrow on the top of the column and see if your lot number(s) is (are) on the list. 3. If you have affected product(s) in your warehouse: a. Quarantine them immediately. b. Check the box “I do have affected unit(s)” in the enclosed reply form. c. Record on the form the total quantity of affected unit(s) and lot number(s) that you have. 4. If you do not have affected product(s) in your warehouse, check the box, “I do not have affected unit(s)”. 5. Please check your customer traceability records for shipments of affected products. 6. If you have shipped impacted products to your customers, please complete below: a. Create a customer reply form with your contact details. b. Forward a copy of the Field Safety Notice to any of your customers that have purchased the affected products and lot numbers. c. Collect completed response forms and affected product(s) from your customers and indicate the total quantities and lot(s) in the distributor reply form. | 2025-12-24 12:03
GE Medical Systems, LLC | Omni Legend | 1 | - | GE HealthCare has become aware of a potential intermittent issue on certain Omni Legend systems that can result in a streaking artifact in the PET clinical scan images In some cases the artifact can mimic pathology and has the potential to lead to misdiagnosis if undetected Note Standalone CT image acquisition is not impacted by this issue x D x D There have been no injuries reported as a result of this issue x D | SW kit & FMI instructions | 2025-12-24 13:37
optimed Medizinische Instrumente GmbH | Tentos4F | 424 | - | Risk due to loosening of the polyimide internal catheter near the coronal mark in the stent system Increased occurrence of the same fault pattern in which the tip of the system breaks off | Recall of the affected lots that are currently in the market. | 2025-12-24 15:09
Saudi Medical Solution Company | SMSCOZYM | 1 | - | No SOP controlling trail batches and improper human practice | Recalled all quantities from the customers and destroyed. | 2025-12-25 14:37
Cardinal Health 200, LLC | Sentinel Seal Chest Drainage Unit Aqua-Seal Chest Drainage Unit Aqua-Seal Chest Drainage Unit Thora-Seal Basic Chest Drainage Unit Sentinel Seal Chest Drainage Unit Thora-Seal Basic Chest Drainage Unit Thora-Seal Chest Drainage Unit Two Chamber System, Collection and Underwater Seal Chambers Thora-Seal Chest Drainage Unit Thora-Seal Replacement Collection Chamber | 79 | - | The purpose of this letter is to communicate to customers that the IFUs instructions for use for the above listed Cardinal Health Chest Drainage Units accessories CDUs have been updated to clarify the intended target patient population as adults years and older | 1. CONTINUE utilizing the CDUs on adult patients (not pediatric patients or infants). 2. REVIEW the updated IFU for the product codes listed above. https://www.mycardinalmsds.com/ 3. POST a copy of this notification in your storeroom and clinical areas. 4. NOTIFY any customers to whom you may have distributed/forwarded affected product (or to whom you intend to distribute/forward product) about this medical device product advisory and share a copy of this notice. 5. RETURN the enclosed acknowledgment form by following the directions on the form. | 2025-12-28 11:25
Cardinal Health 200, LLC | Altitude Chest Drainage Unit Dry Suction, Easy-Change Connector | 79 | - | The purpose of this letter is to communicate to customers that the IFUs instructions for use for the above listed Cardinal Health Chest Drainage Units accessories CDUs have been updated to clarify the intended target patient population as adults years and older | 1. CONTINUE utilizing the CDUs on adult patients (not pediatric patients or infants). 2. REVIEW the updated IFU for the product codes listed above. https://www.mycardinalmsds.com/ 3. POST a copy of this notification in your storeroom and clinical areas. 4. NOTIFY any customers to whom you may have distributed/forwarded affected product (or to whom you intend to distribute/forward product) about this medical device product advisory and share a copy of this notice. 5. RETURN the enclosed acknowledgment form by following the directions on the form. | 2025-12-28 11:25
NOxBOX Ltd | NOxBOXi® | - | - | Rapid succession of conflicting user inputs pressing two or more on screen buttons quickly x D This behavior can trigger a system restart x D The issue can occur on all currently released software versions of the NOxBOXi device x D Root cause is a software user interface vulnerability not a hardware failure | notify all affected customers. Maintain awareness of the notice for an appropriate period. Report device-related incident to SFDA | 2025-12-28 16:38
Maquet Cardiopulmonary GmbH | Bubble Sensor | - | - | Internal investigations have identified an issue with the durability of the connecting cable See Fig position near the connection to the Bubble Sensor Excessive bending of the cable can lead to poor contact which may trigger the errors Ven bubble sensor defective or Ven bubble sensor disconnected on the connected medical device ROTAFLOW II or CARDIOHELP i These errors can occur temporarily when the cable is moved or permanently if the connection is fully compromised | According to our post-market surveillance documentation, you may have products affected by this action. Please examine your inventory immediately to determine, if you have any affected product in your inventory. • Optional for CARDIOHELP-i accessories only: Instead of replacement, the affected accessory may be returned for credit. Upon return of the affected product, please contact your local Getinge representative for credit. • Please always report any adverse events potentially related to the affected products to your Getinge representative. • Duly fill out the enclosed Letter of Acknowledgement and return it to your local Getinge representative as soon as possible, latest by 2026-01-24 by mentioning FSCA - 1427918 as reference in the subject line of your mail. | 2025-12-30 15:32
Phadia AB | EliA ANA positive control 250 | - | - | Description of the problem x D An issue relating to EliA ANA Positive Control EliA ANA Positive Control and EliA x D ANA Positive Control has been discovered The investigation confirmed that the acceptance x D ranges for EliA La on the control certificates required revision following initial publication x D Consequently additional results may be found above outside the limit stated in the control certificates x D This does not have an effect on the correctness of the test results generated for patient samples since x D the EliA La wells are not affected and this will not have an impact on previously obtained quality x D control results x D No adverse events have been reported No other EliA test ranges on the referenced control certificates x D are impacted by this action x D x D Kindly review the attached | Actions to be taken by the customer/user: • Please use the below updated acceptance ranges for EliA La when running EliA ANA Positive Control 250, EliA ANA Positive Control 2500/5000 or EliA ANA Positive Control 200 Actions to be taken by the manufacturer: • Corrective Actions have been initiated to prevent this from recurring. Actions – Commercial organization • Scrap any existing stock of affected products and order replacement products free of charge. • Send the Commercial organization initial report to Global QA as confirmation that the information has been received. • Send the commercial organization final report to Global QA when communication with customers, and authorities if needed, has been performed. Transmission of this Notification: Please ensure that this notice is shared with anyone who needs to be made aware within your organization, or to any organization on which this notification potentially has an impact. We apologize for any inconvenience this event may cause. If you have any questions, please contact us. Kindly review the attached | 2025-12-31 15:41